## Slide 1: DT gas system
DT gas system items for 2021 :
DT gas analysis rack refurbishment
  Status of the new sketch (Louis-Philippe) and procurement of all components
  Inclusion of the gas chromatograph at the bottom
  Milestone for installation in the UGX. It was previously scheduled in Q4 2020.
Main flowmeters (SGX5) recalibration :
  Recalibration for closed loop and check of the already done recalibration in opened loop. It was scheduled in January 2021 in previous mails exchange.

## Slide 2: DT gas system
DT flowcells recalibration
  The observed loss over few years is 20%, ie. drop from 50 L/h (design value) to around 40 L/h
  Recalibration of few individual flowcells didn’t show any problem
  Status and conclusion from last tests
Exhaust anomaly
  when DT gas stopped, anomaly noticed in 2020
  include investigations in the plans in 2021