## Slide 1
LAViSta
First tests of the LAPP’s support system
LAPP
Laurent Pacquet
Guillaume Deleglise
Jean-Philippe Baud
Andrea Jérémie
Symme
Ronan Le Breton
Adrien Badel
Jacques Lottin
14/09/2010

## Slide 2
mid-lower magnet
1355mm
piezo
actuator
base
200mm
200mm
elastomeric joint
parallelism tuning
53mm
The active support
Notes: lets remind reminde
function ball joint to avoid shearing on piezoelectric material
pièces mâle / femelle
2 supports

## Slide 3
elastomeric strips for guidance
piezoelectric actuator below its micrometric screw
lower electrode of the capacitive sensor
fine adjustments for capacitive sensor (tilt and distance)
V-support for the magnet
carrying handles
micrometric screws in contact with the actuators
Active system principle

## Slide 4
Control strategy
geophone
ground
Base
Stabilized platform
geophone
piezo
Capacitive sensor
For feedback control, at least one sensor is required: (upper geophone)
The capacitive sensors give additional information for better control of the actuators elongation
If the lower geophone and the capacitive sensor is used, then the upper geophone is redundant

## Slide 5
Measuring range / voltage | 15 µm / 10 V
Noise at 100 Hz | 1.75 pm/√Hz
Resolution at 3 kHz | 0.075 nm
Linearity | ± 0.05%
Size (each electrode) | 15 x 8 x 4
Maximal displacement | 8 µm
Voltage range | -20 V -> 150 V
Resolution | 0.08 nm
Stiffness | 100 N/µm
Max load | 800 N
Size | 10 x 10 x 18
Piezoelectric actuator Cedrat PPA10M
Capacitive sensor PI D-015.00 - 2 sensitive electrodes
Intermediate
experimental setup
70mm
50mm
Notes: This is an intermediate experimental setup used to characterize piezo-actuators and sensors

## Slide 6
-> data (in blue) well fit with the factory’s calibration data (black plot)
Experimental setup – measures -

## Slide 7
Driving voltage = 20 mVpp
Measured displacement = 0.7 nm
Measurement noise = 0.1 nm
Critical point :
to reach a resolution of 0.1 nm, a clean driving voltage in the order of 1mVpp has to be generated
Roughly egal to the sensor noise
Measurements with the Cedrat actuator
Notes: &a

## Slide 8
Driving voltage = 1Vpp
Measured displacement = 0.25 nm
Measurement noise = 0.1 nm
Critical point :
Maximal range: about 40nm
Measurements with the home-made actuator

## Slide 9
 | Cedrat PPA10M | Ceramic patch
actuator |  | 
size | 10x10x18 | 10x10x0.3 (ceramic)
sensitivity measured | 40 nm/V | 0.25 nm/V
400mV excitation | 16 nm | 0.1 nm
150V excitation | 6 µm | 37.5 nm
Possibility to assemble both actuators to have a larger displacement dynamic.
-> control is more complex to coordinate (two driving voltage)
Comparison of both actuators
Nano-stabilization
Micro - positioning

## Slide 10
4 accelerometric sensors
2 electrodes of the capacitive sensors
Measurements on the instrumented support
 | PCB 352C34 | Bruel & Kjaer
sensitivity [mV/ms-2] | 10.02 | 10
frequency range [Hz] | 0.5 – 10 000 | 0.3 - 6000
Accelerometric sensor used
Experimental conditions :
acquisition rate : 2 kHz
- actuators fed with the same signal, which is a sequence from sinus 1Hz to 100Hz by step of 1Hz ( 0-20V)
accelerometrics signals pass an anti-aliasing filter and are amplified by 40dB, then high-pass filtered at 10Hz by software (poor coherence below)
no filter on capacitive data
- total length recording : about 5 min

## Slide 11
Frequency responses of the instrumented support
Signals are coherent
Capacitive sensor has the quickest response

## Slide 12
-> Sensibility depends on the operating point
(45nm/V if 10V, 55nm/V with 40V)
-> Implementation of a model of the actuator in the feedback control to counterbalance this hysteresis may be of interest.
Hysteresis effect

## Slide 13: Hysteresis compensation in piezoelectric actuator
Hysteresis effect in piezoelectric actuator :
Non linear effect
Asymmetric hysteresis
loop
Sensibility evolves with operating points (30%)
For a given voltage, the elongation depends on the history (ex: [-2.5; -4] at 60V)
×

## Slide 14: Hysteresis model
Need for a hysteresis model ?
To obtain a complete model for accurate simulation of the controller
To linearize controller using an inverse hysteresis operator

## Slide 15: Hysteresis model
How to modelize hysteresis effect?
Generalized Maxwell Slip (GMS)
model of hysteresis (symmetric)
Preisach model (assymmetric)
Other model ?
Representation of the Maxwell-slip model

## Slide 16: Preisach model
A weighted combination of elementary relay hysteresis elements : n × (n+1)/2
n: number of step in
the major loop
Used in simulation

## Slide 17: Lightweight hysteresis inverse model
A lightweight model [A. Badel]
Asymmetric loops
Only 6 parameters
Easily implementable for the real time inverse model

## Slide 18: Lightweight hysteresis inverse model

## Slide 19: Lightweight hysteresis inverse model
Real model of actuator

## Slide 20: Lightweight hysteresis inverse model
Real model of actuator
Nonlinear controller

## Slide 21
Conclusion
Measurement and actuation at the sub-nanometer scale is feasible (at least the measurement of the actuator elongation)
Compromise between actuator range and sensibility is critical. Two stages actuator may be a solution.
Actuator Hysteresis may be a problem (especially for “high” frequencies – still under study)
Real time hysteresis compensation should be feasible