## Slide 1: Results from the MBPL Measurements in NA64
Melvin Liebsch, Jens Kaeske and Thomas Zickler

## Slide 2: Magnetic Measurement Request
Two bending dipole magnets in the north area
Magnets are equipped with a vacuum chamber with unknown properties
Permeability of vacuum chamber is to be characterized
Domain inside the chamber is to be mapped
Requested accuracy in the % range

## Slide 3: The Vacuum Chamber
Stainless seel type: 304L
Two welding seams along the chambers length
= 1.07 +/- 0.01
= 1.07 +/- 0.01
Permeability was measured with a FOERSTER Magnetoscop
Thanks to Mariano for the permeability measurements!

## Slide 4: Field Simulations
The impact of the vacuum chamber was estimated by field simulations
The magnet and the vacuum chamber were modelled in Opera
The vacuum chamber mainly results in a reduction of the gap height Field enhancement
Chart: Field enhancement/attenuation due to vacuum chamber (304) at y = 0 mm
| -150 | -140 | -130 | -120 | -110 | -100 | -90 | -80 | -70 | -60 | -50 | -40 | -30 | -20 | -10 | 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150
I = 830 A | -0.21232 | -0.088077 | -0.022807 | 0.016857 | 0.043286 | 0.061401 | 0.07394 | 0.082475 | 0.08842 | 0.092488 | 0.095247 | 0.097097 | 0.098314 | 0.099071 | 0.099487 | 0.099624 | 0.099488 | 0.099072 | 0.098313 | 0.097101 | 0.095252 | 0.092492 | 0.088429 | 0.082514 | 0.073949 | 0.061369 | 0.043294 | 0.016897 | -0.023083 | -0.088081 | -0.21205
I = 500 A | 0.010638 | 0.1419 | 0.21311 | 0.25714 | 0.28651 | 0.30638 | 0.31982 | 0.32867 | 0.33464 | 0.33857 | 0.34113 | 0.34279 | 0.34384 | 0.34447 | 0.34481 | 0.34492 | 0.34481 | 0.34447 | 0.34383 | 0.34279 | 0.34114 | 0.33857 | 0.33464 | 0.32871 | 0.31982 | 0.30633 | 0.2865 | 0.25717 | 0.21279 | 0.14188 | 0.010888
Chart: Absolute field at I = 830 A and y = 0 mm
| -150 | -140 | -130 | -120 | -110 | -100 | -90 | -80 | -70 | -60 | -50 | -40 | -30 | -20 | -10 | 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150
no vacuum chamber | 16906.6889 | 16988.4664 | 17055.5586 | 17110.1358 | 17154.5463 | 17190.3518 | 17219.0146 | 17241.5966 | 17259.5992 | 17273.7526 | 17284.7385 | 17293.0926 | 17299.2248 | 17303.4025 | 17305.8405 | 17306.6709 | 17305.8443 | 17303.4024 | 17299.2159 | 17293.102 | 17284.7448 | 17273.7517 | 17259.6049 | 17241.6496 | 17219.0101 | 17190.2878 | 17154.5327 | 17110.146 | 17055.4004 | 16988.451 | 16906.6861
with vacuum chamber + welding (non-linear) | 16870.7933 | 16973.5034 | 17051.6687 | 17113.0201 | 17161.9718 | 17200.9069 | 17231.7463 | 17255.8166 | 17274.8602 | 17289.7288 | 17301.2017 | 17309.8837 | 17316.2323 | 17320.5452 | 17323.0575 | 17323.9124 | 17323.0615 | 17320.5452 | 17316.2232 | 17309.8938 | 17301.2089 | 17289.7285 | 17274.8674 | 17255.8764 | 17231.7434 | 17200.8374 | 17161.9596 | 17113.0371 | 17051.4634 | 16973.4874 | 16870.8355
Thanks Thomas for these simulations!

## Slide 5: Magnetic Field Maps
Constraints:
  3D mapper is occupied to set up the FAIR measurements
  Translating Fluxmeter does not fit in the vacuum chamber
  Time: Measurements needed to be performed in March 2021
Solution:
The Hall probe array
Thanks to Jens for his work on the development of this system!

## Slide 6: The System in a Nutshell
Guiding rails 2 mm sag over 5 m
Sled
Rods pulling the sled
Retroreflector
7 Hall probes
Asensor HE244
Field can be mapped on 5 vertical positions

## Slide 7: The System in a Nutshell
Support
Motor
Leica Laser Tracker

## Slide 8: The System in a Nutshell

## Slide 9: Sources of Noise
Keithley introduced noise at 8 kHz
Aliasing Filter with 2 kHz corner frequency

## Slide 10: Sources of Noise
The motor was is a large noise contributor
Signal cable was shielded to motor ground

## Slide 11: Sources of Noise
Signal noise of 20 uVpp at sampling frequency of Hz
Integration period = 10 ms
This corresponds to 1 Gauss in field

## Slide 12: Calibration
Hall probes were calibrated between +/- 1.4 T in 311
The linearity error in this range is below 5 units of 10 000

## Slide 13: Verification in 311
System was used to measure the „Panzer“ in 311
A comparison to the SSW integral measurement was made
15 mTm
Why do the Hall probes see more field?
We are missing some of the field!

## Slide 14: Verification in 311
System was used to measure the „Panzer“ in 311
A comparison to the SSW integral measurement was made
A Large contribution to the error can be related to the smaller measurement range
wire
Hall Array
Hall Array sees more field at the bottom positions!
This is due to the sag of 2 mm

## Slide 15: Measurements in 887

## Slide 16: Temperature Dependency
The calibration was performed in the controlled environment in 311 T = 23 °C
At 887 the temperature is much lower T = 18.5 °C
We have to recalibrate the sensitivity!
Typical value given in the datasheet

## Slide 17: Magnet Cycles
Chart: Magnet Current [A]
 | 0 | 850 | 850 | 0 | 0 | 100 | 100 | 850 | 850 | 0 | 0 | 300 | 300 | 850 | 850 | 0 | 0 | 500 | 500 | 850 | 850 | 0 | 0 | 800 | 800 | 850 | 850 | 0 | 0 | 810 | 810 | 850 | 850 | 0 | 0 | 820 | 820 | 850 | 850 | 0 | 0 | 830 | 830 | 850
830 A
100 A

## Slide 18: Magnet Cycles
Chart: Magnet Current [A]
 | 0 | 850 | 850 | 0 | 0 | 100 | 100 | 850 | 850 | 0 | 0 | 300 | 300 | 850 | 850 | 0 | 0 | 500 | 500 | 850 | 850 | 0 | 0 | 800 | 800 | 850 | 850 | 0 | 0 | 810 | 810 | 850 | 850 | 0 | 0 | 820 | 820 | 850 | 850 | 0 | 0 | 830 | 830 | 850
830 A
100 A
1
2
3
4
5
6
7

## Slide 19: Transient Effects
Probe in the magnet center
Ramping to requested currents
How long you need to wait until decay?
The yoke is bulk Eddy currents with large time constants
We are at the moment preparing for another measurement of this type separating ramp up and down

## Slide 20: Field Integrals
B(I)
x coordinate wrt to chamber center. Magnet is misaligned wrt chamber!
magnetic center [mm] | des. mech. center [mm]
-4.6479 | -4.702
Coincidence within 50 um!

## Slide 21: Field Profiles Field Quality

## Slide 22: Measurements done for 2 magnets at 7 current levels
300 A
500 A
800 A

## Slide 23: Conclusion
Measurement request satisfied
Total costs of new components < 10 000 CHF
Possible Improvements:
Instrumental amplifiers to get single ended signals for all Hall probes  parallel acquisition with DAQ
Controlling the Leica from FFMM
Mechanism to pull back the cable
Full automation
7x faster

## Slide 24: Acknowledgements
The whole project was finished within 3 months.
Mainly because of the help of the 311 team!
Special thanks to:
Dmitry for the design
Nini for machining
Fatia for electronics and cabling
Richard for manufacturing the Hall probe array
Guy for building the Leica trigger box