## Slide 1: Results of the laser-polished sample B-5.5 test (preliminary)
D. Tikhonov, Authors
EASITrain – European Advanced Superconductivity Innovation and Training. This Marie Sklodowska-Curie Action (MSCA) Innovative Training Networks (ITN) has received funding from the European Union’s H2020 Framework Programme under Grant Agreement no. 764879
ARIES WP meeting
Presenter
Dmitry Tikhonov
Early stage researcher
Helmholtz-Zentrum Berlin
dmitry.tikhonov@helmholtz-berlin.de
+49-30-8062-12922

## Slide 2: Rs vs B for 400 MHz
Higher surface resistance
Results for pre-laser polishing and after for 2.5 K, ~400 MHz:

## Slide 3: Rs vs B for ~400 MHz
Bmax ~27 mT
Rs vs B is more “smooth”, one spot is better
Possible problems: problematic spot could be at higher field region
Maximum film is about same as previously tested (27 mT). Probably, problematic spot was not reached by laser.
Results for pre-laser polishing and after for 4.5 K, ~400 MHz:

## Slide 4: Max field, 416 MHz
Pulse shape (transmitted)
At max field 27 mT, 416 MHz: Temperature significantly increased (but not complete quench)
Only part of the film quenched (not full film because temperature stays <9.3K)
Quench field depends on temperature:
Chart: Temperature, K
| 3 | 4 | 5 | 5.5 | 6 | 6.5 | 6.75 | 7 | 7.25 | 7.5 | 7.75 | 8 | 8.25 | 8.5 | 8.75 | 9 | 9.25 | 9.5 | 9.75
 | 30.3 | 29 | 27.2 | 25.9 | 24.25 | 21.85 | 20.35 | 18.6 | 16.7 | 14.7 | 12.2

## Slide 5: Rs vs B for ~850 MHz
Bmax ~22 mT
Maximum film is about 22 mT.
Results for pre-laser polishing and after for ~850 MHz:

## Slide 6: Rs vs B results for 2 modes
Q1 vs Q2

## Slide 7: Rs vs T, ~400MHz
Rs vs T: Most of the ‘additional’ resistance come from temperature-depend part
Residual resistance is approximately same:
Rs vs T results for pre-laser polishing and after for ~416 MHz:
 | Rres
B-5.4 | ~67 nOhm
B-5.5 | ~76 nOhm
Residual resistance
‘eye’ estimation (416 MHz):

## Slide 8: Rs vs T, ~400MHz
Rs vs T: Most of the ‘additional’ resistance come from temperature-depend part
2.5
4.5
Rs vs T results for pre-laser polishing and after for ~416 MHz:
~35 nOhm
~85 nOhm

## Slide 9: Rs vs T, ~400MHz
Rs vs B results for pre-laser polishing and after for ~416 MHz:
Rs with subtracted ~BCS temperature dependent part at 4.5 K (approximation):
Probably increased BCS came from impurities and defects due to wrong polishing procedure
With subtracted BCS contribution from both tests the field dependent part is more or less comparable except higher field ‘jump’

## Slide 10: Jumps in surface resistance: edges

## Slide 11: Jumps in surface resistance: edges
B-5 sample
B-4 sample
Conclusions:
For the reliable data, next samples should be treated same way (flash etching, etc) to remove sharp pieces of Nb from edges

## Slide 12: F vs t results for pre-laser polishing and after

## Slide 13: F vs t results for pre-laser polishing and after
Impurities/Defects?

## Slide 14: Post-test sample scan: comparison
Pre LP:
Post LP:

## Slide 15: Post-test sample scan: comparison
Pre LP:
Post LP:
Some defects removed or partially removed

## Slide 16: Post-test sample scan: Overview
Sample scanned with the laser microscope:
Defects detection
Roughness measurements
Full size JPEG picture (34 MB) and other data are on the CERNbox folder (note, that some features = dust)

## Slide 17: Post-test sample scan: Surface defects & roughness
Some observed features:
1
2
3
4
Major defect near the edge
Minor craters in high field region
Laser traces (overpower?) near the edge
A “group of craters”
This pattern = effect of stitching

## Slide 18: defect: number 1
Defect 1 found near the edge:
Optical microscope view:
Laser microscope view (5x):

## Slide 19: defect: number 1
Defect 1 found near the edge:
Laser microscope view (5x):
Laser microscope view (20x):

## Slide 20: defect: number 1
Defect 1 view from optics in 3D:
The defect was not found so far on pre-laser coated sample
Shape of the defect: the laser could give overpowered pulse
The film is completely penetrated by the “shot”, substrate is exposed to RF
Cu?

## Slide 21: defect: number 1
Defect 1 view in 3D:

## Slide 22: defect: number 1
Defect 1 view in 3D:
The film deepness can be estimated about 5 mkm

## Slide 23: defect: number 2
Defect 2 found in the high RF field region:
Optical microscope view:
Laser microscope view (5x):

## Slide 24: defect: number 2
Laser microscope view (5x):
Laser microscope view (20x):
Defect 2 found in the high RF field region:

## Slide 25: defect: number 2
Defect 2 found in the high RF field region:
Defect again looks like a crater (series of craters), deepness up to 1 mkm
Roughness from substrate shape (due to mechanical production) is about 1 mkm
Rings on Cu from mech. turning

## Slide 26: defect: number 3 & 4
Laser microscope view (5x):
Craters again?

## Slide 27: defect: number 3 & 4
Laser microscope view (5x):
Craters & overpowered region

## Slide 28: Conclusions
Conclusions
Laser polishing practically can work
Minor improvements at Rs vs B behavior. The problem with the rough edges is still present, and % of contribution of it to the final result is not clear.
Film (probably) got some contaminations and defects during the process (conclusion from RBCS and FvsT data).
The process needs to be adjusted and surface analyzed more carefully.

## Slide 29
B-2 | CERN (HiPIMS Nb 7.3 micr.) | Rf test: 02.2021 CERN | ??? Nb Film INFN | ??? Rf test: 04.2021 HZB | ? | ?
B-3 | HZB (Siegen 1st SIS) | HOLD (for future test) | ? | ? | ? | ?
B-4 | HZB (Siegen 2nd SIS) | Nb Film strip. INFN | 3ed SIS USiegen | Rf test: 02.2021 HZB | ? | ?
B-5 | HZB (RTU laser) | Rf test: 03.2021 HZB | Nb Film strip. INFN | SIS coat. STFC | Rf test: 04.2021 HZB | ?
Now 25.01.2021
Next plan
Future plan

## Slide 31: Conclusions
Suggested Next steps
If possible (not at HZB): Any analyses regarding molecular components of the film and/or SEM microscope scan
Suggestions: Stripping of the top layer of the film and retesting to see if the contaminations are removed?