## Slide 1: Magnetic Measurements on MCBXFBP2
Matthias Bonora
Lucio Fiscarelli
With the help of colleagues from TM and CIEMAT
Topical meeting on MCBXFB/A, 3rd March 2021

## Slide 2: Introduction
Reference radius 50 mm
Vertical field from the inner coils:
- main field B1 (normal)
- first allowed b3
Horizontal field from the outer coils:
- main field A1 (skew)
- first allowed a3
Rotating-coil in the helium bath of vertical cryostat clusterD in SM18
  5 segments (500 mm each, 2.5 m total)
2

## Slide 3: Powering cycles

## Slide 4: Transfer function
 | Unit | Inner Dipole | Outer Dipole
Nominal current | A | 1625 | 1474
Integrated field (model) | Tm | 2.62 | 2.59
Integrated field (measured) | Tm | 2.56 (-2.3% wrt. model) | 2.60 (+0.4% wrt. model)
Integrated field (previous measurement) | Tm | 2.64 (+0.8% wrt. model) | 2.62 (+1.2% wrt. model)

## Slide 5: Multipoles vs current

## Slide 6: Multipoles vs current

## Slide 7: Multipoles vs current

## Slide 8: Multipoles vs current

## Slide 9: Multipoles vs current

## Slide 10: Multipoles vs current

## Slide 11: Multipoles vs current

## Slide 12: Multipoles vs current

## Slide 13: Field-quality table
INNER at 1625A Outer at 0A |  |  | INNER at 0A OUTER at 1474A |  |  | INNER at 1200A OUTER at 1090A |  |  | INNER at 1200A OUTER at -1090A |  | 
n | bn | an | n | bn | an | n | bn | an | n | bn | an
2 | 1.06 | 3.65 | 2 | -0.06 | -0.05 | 2 | 1.01 | 1.18 | 2 | 0.07 | -1.64
3 | -10.15 | 0.32 | 3 | -0.06 | -4.97 | 3 | 4.63 | 2.26 | 3 | 4.74 | 3.35
4 | 0.49 | 2.55 | 4 | -0.41 | 0.28 | 4 | -0.13 | 1.43 | 4 | 0.39 | -1.22
5 | -7.90 | -0.32 | 5 | -0.07 | -1.03 | 5 | -4.85 | -1.09 | 5 | -4.74 | -0.44
6 | 0.19 | 0.89 | 6 | 0.67 | -0.23 | 6 | 0.75 | 0.70 | 6 | -0.24 | -1.00
7 | -3.10 | 0.36 | 7 | 0.10 | 2.66 | 7 | -2.83 | 2.10 | 7 | -3.00 | 1.95
8 | -0.25 | -0.07 | 8 | -0.15 | 0.02 | 8 | -0.09 | 0.10 | 8 | 0.13 | -0.07
9 | 0.34 | 0.47 | 9 | 0.01 | -0.50 | 9 | -0.24 | -0.25 | 9 | -0.26 | -0.48
10 | -0.22 | -0.95 | 10 | -0.04 | 0.01 | 10 | -0.25 | -0.64 | 10 | -0.20 | 0.63
11 | 2.26 | -0.15 | 11 | -0.02 | -0.05 | 11 | 1.57 | -0.19 | 11 | 1.60 | 0.11
12 | 0.19 | 0.88 | 12 | 0.02 | 0.02 | 12 | 0.25 | 0.64 | 12 | 0.15 | -0.61
13 | -2.21 | -0.02 | 13 | -0.03 | 0.06 | 13 | -1.31 | 0.17 | 13 | -1.25 | 0.04
14 | -0.05 | 0.02 | 14 | -0.00 | 0.08 | 14 | 0.00 | -0.01 | 14 | -0.02 | 0.15
15 | 0.66 | -0.17 | 15 | 0.02 | 0.03 | 15 | 0.16 | -0.06 | 15 | 0.17 | 0.11
Units of 10-4 at the reference radius of 50 mm

## Slide 14: Comparison with previous measurement
Units of 10-4 at the reference radius of 50 mm
INNER at 1625 A OUTER at 0A |  |  | previous measurement |  |  | difference |  | 
n | bn | an | n | bn | an | n | bn | an
2 | 1.06 | 3.65 | 2 | 1.75 | 2.81 | 2 | 0.69 | -0.84
3 | -10.15 | 0.32 | 3 | -6.91 | -0.07 | 3 | 3.24 | -0.39
4 | 0.49 | 2.55 | 4 | 1.78 | 0.81 | 4 | 1.29 | -1.74
5 | -7.9 | -0.32 | 5 | -7.45 | 0.67 | 5 | 0.45 | 0.99
6 | 0.19 | 0.89 | 6 | 1.33 | -0.18 | 6 | 1.14 | -1.07
7 | -3.1 | 0.36 | 7 | -3.77 | 0.76 | 7 | -0.67 | 0.4
8 | -0.25 | -0.07 | 8 | 0.11 | -0.08 | 8 | 0.36 | -0.01
9 | 0.34 | 0.47 | 9 | 0.00 | -0.06 | 9 | -0.34 | -0.53
10 | -0.22 | -0.95 | 10 | -0.8 | 0.38 | 10 | -0.58 | 1.33
11 | 2.26 | -0.15 | 11 | 1.45 | -0.64 | 11 | -0.81 | -0.49
12 | 0.19 | 0.88 | 12 | 0.89 | -0.24 | 12 | 0.7 | -1.12
13 | -2.21 | -0.02 | 13 | -1.47 | 0.77 | 13 | 0.74 | 0.79
14 | -0.05 | 0.02 | 14 | -0.06 | -0.13 | 14 | -0.01 | -0.15
15 | 0.66 | -0.17 | 15 | 0.12 | 0.06 | 15 | -0.54 | 0.23
INNER at 0 A OUTER at 1474 A |  |  | previous measurement |  |  | difference |  | 
n | bn | an | n | bn | an | n | bn | an
2 | -0.06 | -0.05 | 2 | 0.78 | 0.55 | 2 | 0.84 | 0.6
3 | -0.06 | -4.97 | 3 | -0.22 | -7.56 | 3 | -0.16 | -2.59
4 | -0.41 | 0.28 | 4 | -0.05 | -0.16 | 4 | 0.36 | -0.44
5 | -0.07 | -1.03 | 5 | 0.29 | 2.00 | 5 | 0.36 | 3.03
6 | 0.67 | -0.23 | 6 | 0.09 | 0.7 | 6 | -0.58 | 0.93
7 | 0.1 | 2.66 | 7 | -0.61 | -2.64 | 7 | -0.71 | -5.3
8 | -0.15 | 0.02 | 8 | 0.07 | -0.15 | 8 | 0.22 | -0.17
9 | 0.01 | -0.5 | 9 | 0.17 | 0.54 | 9 | 0.16 | 1.04
10 | -0.04 | 0.01 | 10 | -0.05 | -0.02 | 10 | -0.01 | -0.03
11 | -0.02 | -0.05 | 11 | -0.01 | 0.02 | 11 | 0.01 | 0.07
12 | 0.02 | 0.02 | 12 | 0.02 | 0.03 | 12 | 0 | 0.01
13 | -0.03 | 0.06 | 13 | -0.06 | -0.02 | 13 | -0.03 | -0.08
14 | 0 | 0.08 | 14 | 0 | 0 | 14 | 0 | -0.08
15 | 0.02 | 0.03 | 15 | -0.04 | 0.01 | 15 | -0.06 | -0.02

## Slide 15: Conclusions
Transfer function at nominal level
  Inner plane 2.558 Tm
  No saturation on inner plane with individual powering
  Outer plane 2.600 Tm
  Saturation on outer plane -1.3% with individual powering
Field quality at nominal level
  Inner plane
    b3 of -10.2 units with individual powering
    b5 of -7.9 units with individual powering
  Outer plane
    a3 of -5.0 units with individual powering