## Slide 1: MCBXFBP2b Mechanical measurements
Mechanical Measurement Lab, EN/MME
M. Guinchard / S. Mugnier
3rd March 2021

## Slide 2
Outline
      Instrumentation overview
      Strain measurement during collaring
      Thermal cycle
      Powering tests
      Bath instability at 4.5 K
      Conclusion

## Slide 3
1-Instrumentation overview

## Slide 4: Instrumentation overview
Azimuthal strain measurements on the inner coil :
12 instrumented collars in total :
    6 instrumented collars equipped for temperature and magnetic field compensation
    6 instrumented collars equipped only for temperature compensation (with external compensators)
Azimuthal strain measurements on the outer coil were performed only during initial collaring and removed for final assembly

## Slide 5: Instrumentation overview
Equipment | Reference
Strain gauges | HBM 1-XC11-1.5/350
Bridge power supply | 2.5 V @ 4.8 kHz
Bridge Configuration | Half bridge – Poisson compensation
Quantity | 6 collars / both sides
a/b
c/d
Strain gauges a & c are on the same side of the collar, b & d are on the other side

## Slide 6: Instrumentation overview
Equipment | Reference
Strain gauges | HBM 1-LC11-1.5/350
Bridge power supply | 2.5 V @ 4.8 kHz
Bridge Configuration | Quarter bridge + compensators
Quantity | 6 collars / both sides
a/b
c/d
Strain gauges a & c are on the same side of the collar, b & d are on the other side

## Slide 7
Instrumentation overview
MX840 for SG reading
MX1615 for SG reading
Settings
Data
Mechlab Server

## Slide 8
2-Strain Measurement during collaring

## Slide 9
Stress measurement overview
MCBXFBP2a target value for inner collaring : 120 MPa at RT
MCBXFBP2b target value for inner collaring : 140 MPa at RT
 Avg. Measured after collaring : 135 MPa
 Avg. Measured before cooldown : 115 MPa
MCBXFBP2b target value for outer with a shim at 0,3 mm: +20 MPa more than MCBXFBP2a at RT
MCBXFBP2b target value at cold : 60 MPa at 4.5 K
 Avg. Measured at cold : 45 Mpa

## Slide 10
Last Inner collaring at 293 K in Building 927

## Slide 11
Last Inner collaring at 293 K in Building 927
MCBXFBP2b
MCBXFBP2a
Overall stiffness : 900 µm/m/mm
Overall stiffness : 650 µm/m/mm

## Slide 12
Outer comparison under press with shimming at 0.3 mm
The outers collars were removed for the final collaring phase before the cold tests.

## Slide 13
At 0.3 mm of shimming: The target value was + 20 MPa (about 100 µm/m) more than for MCBXFBP2a
MCBXFBP2b
MCBXFBP2a
155 MPa
180 MPa
Shimming at 0,3 mm
Shimming at 0,3 mm
Outer comparison under press with shimming at 0.3 mm

## Slide 14
3-Thermal Cycle

## Slide 15
Thermal cycle
Raw values without K factor corrections.
Cooldown
Cooldown
Cooldown
Warm-up
Warm-up
Warm-up

## Slide 16
4-Powering tests

## Slide 17
Inner dipole at ultimate current: 1690A
16 February 2021 at 02:00PM

## Slide 18
Outer dipole at ultimate current: 1580A
16 February 2021 at 05:30PM

## Slide 19
Training inner and outer dipole at nominal current, in Quadrant 1 (I Inner=1579A and I Outer=1430A)
15 February 2021 at 06:00PM

## Slide 20
5-Bath instability at 4,5 K

## Slide 21
Bath instability at 4.5 K (Boiling)
Change from 1.9 K to 4.5 K: Instability begins around 6:00 p.m. at a bath temperature of 2.2 K (Raw values without K factor corrections)
2,2K
2,2K
2,2K

## Slide 22
Conclusion
MCBXFBP2b was successfully monitored during the collaring process, thermal cycles and powering phases thanks to the 48 electrical strain gauges installed.
MCBXFBP2b was loaded with an additional stress of 20 MPa comparing to MCBXFBP2a.
A stress relaxation of 20 MPa was measured in 3 weeks.
Thermal cycles and powering phases do not affect the mechanical stresses of the structures.
During powering phases, a linear response of the inner coil strains was observed versus I2 Inner. For the outer coil powering and both coils powering: additional investigation needs to be done to understand coupling effects: Data can be provided.
Some instability on the gauges was observed only at 4,5 K in the cryostat (circuit flowing from top to bottom inside the cryostat)

## Slide 23
Thank you for your attention

## Slide 24
Warm up (raw values)
23 February 2021 to 26 February 2021

## Slide 25
Ramp Inner at 1580A then Outer at 1384A
16 February 2021 at 10:05AM

## Slide 26
Training inner and outer dipole at nominal current, in Quadrant 2 (I Inner=-1529A and I Outer=1395A)
18 February 2021 at 12:18PM

## Slide 27
Quench free cycle – 12 ramps
19 February 2021 at 04:18PM to 06:06PM