## Slide 1
E. Benedetto
Synchrotron injection energy
HITRIplus WP7 meeting
7/5/2021
see also slides @ Joint SEEIIST/NIMMS Meeting 13/11/2020

## Slide 2: Motivation
What is the best injection energy?
(Why Japanese inject at 4-5 MeV/u
and Europeans at 7 MeV/u?)

## Slide 3: Injection energy in the synchrotron
  Space charge
    small for C-ions (…is it still the case if 20x intensity?)
  Geometrical emittance
    beam dimensions, magnet aperture
  Revolution frequency
    # particles that can fill 1 turn
  Stripping foil
    efficiency C+4  C+6 is 95% at 4 MeV/u
    Multiple Coulomb scattering
  Linac design
    length, RF power, … cost

## Slide 4: Scaling
  Space charge: ~bg2
  Geometrical emittance: ~bg
  Revolution frequency: ~b
  g ~1
    Dn ~0.13 for C-ions @ 4 MeV and increases w. energy
…~not a problem for C-ions!
    Beam size ~ sqrt(b)
Energy MeV/u | Beta
4 | 0.09
5 | 0.10
7 | 0.12

## Slide 5: Space Charge
  Dn ~0.13 (*) for C-ions @ 4 MeV and increases w. energy
…not a problem for C-ions!
  Dn ~0.42(*) @ 7 MeV protons (Dn=0.35(*) @ 10 MeV/u)
     Dn=0.31 @ 7 MeV/u if 3um (more realistic!)
(*) with 1um ring emittance, worst case
Dn < 0.25-0.3 “safe”
Dn ~ 0.5-0.6 “feasible but…”

## Slide 6: Stripping efficiency C+4  C+6
Amra, Mariusz
95% @ 4 MeV/u

## Slide 7: Stripping efficiency O+6
Amra, Mariusz
~85% @ 4 MeV/u

## Slide 8: Multiple Coulomb Scattering @ foil
E. Benedetto – NIMMS/SEEIIST Meeting 13/11/2020
5% increase, if 0.6 pi mm mrad
(x2 if emi ~0.3 pi mm mrad)
Amra, Mariusz

## Slide 9: Bfield in SC magnets
For C-ions:
  B@inj = 0.4 T @ 7 MeV/u
  For protons:
  B@inj = 0.2 T @ 7 MeV/u TOO LOW!
  Bmin= ~0.3 T corresponds to 16 MeV protons
  if protons @ 12 MeV  B=0.26 T,
  if protons @10 MeV 0.24 T

## Slide 10: Conclusions (for now…)
Lower energy for C-ions
  7 MeV/u looks OK
  5 MeV/u
    space charge not an issue
    Bfield in SC magnet getting low
    MCS at stripping foil increases
    beam size (and aperture) increases
Higher energy for protons?
  16 MeV ideal, but can compromise
    space charge OK
    Bfield = 0.3 T OK

## Slide 11: Source
~500 euA of C+4 @inj point(!)
  assuming 100% linac transmission!
  with 30 turns x 50% efficiency MT, final emi 5um instead of 1um should be ok
emi = 0.7 pi um rms, normalized
AISHA, Lanzhau:
800euA
0.2-0.3 pi mm mrad
CNAO Supernanogan
250 euA
0.52 pi mm mrad
make sure all the emittances are rms (vs. full) and currents refers to C+4