## Slide 2: Modern SEM materials and methods at CERN
Christophe Vuitton Benjamin Moser

## Slide 3: Modern SEM materials and methods at CERN
Christophe Vuitton Benjamin Moser
Presentation to Aries

## Slide 4: Modern SEM materials and methods at CERN
Christophe Vuitton Benjamin Moser
Presentation to Aries

## Slide 5: Overview
Functionality of a SEMGRID
Production of SEMGRID platines
Preparation of wires for SEMGRIDS
Electrical connections for SEMGRIDS
DAQ
Sticking wires

## Slide 6: How does a SEMGRID work
When the beam passes secondary electrons are ejected from the wires
This results in current flowing in the wire
Spatial resolution is given by the wire spacing
Time resolution by the DAQ system

## Slide 7: How does a SEMGRID work
When the beam passes secondary electrons are ejected from the wires
The current flowing back onto the wires is measured
Resolution is given by wire distance

## Slide 8: By design, SEM grids measure the beam transverse profiles. The profile data can be used to infer:
Transverse emittance
Position
Intensity
Energy and energy spread

## Slide 9: Components of a SEMGRID HEAD
Support structure
BIAS
Wire support
Wires
Connector
“cabling”

## Slide 10: Components of a SEMGRID HEAD
Support structure
BIAS
Wire support
Wires
Connector
“cabling”

## Slide 11: Wire support
A PCB with pads to solder the wire grid onto
Materials used need to conform to CERN vacuum requirements
Overall PCB design is given by beam aperture, vessel for the instrument and spacing of the grid
Using Resistors for test signal to confirm the state the wires without beam

## Slide 12: Wire support
Alumina (Al2O3) ceramic substrate 96-97% Biggest Panel possible with our providers so far are 160x240mm
PCB tracks single or double sided
VIAS diameter 0.3 mm
Tracks thickness 35 μm (1oz) copper with ENIG finish
SMD Resistor 1 MΩ packs
Possible to order as a “standard product”
Performed vacuum acceptance tests at CERN and so far had no issues with them, all cleared for installation

## Slide 13: Examples
Grid with 0.4 mm spacing

## Slide 14: Examples
Grid with 1 mm spacing

## Slide 15: Examples
Grid with smaller spacing in the centre of the beam

## Slide 16: Examples
With gold solder, can be baked out to 200 C

## Slide 17: Components of a SEMGRID HEAD
Support structure
BIAS
Wire support
Wires
Connector
“cabling”

## Slide 18: Wires
Material | Diameter | Tension of the wire | Smallest wire distance
Tungsten with gold coating | 40 μm | 1.5 N | 0.4 mm
Carbon ex-brai | 34 μm | 0.2 N | 0.4 mm

## Slide 19: Preparation of the wire grid
Guides on either side, which need to be aligned by hand
Laying the wires by hand
Used for Carbon wires as they are more stiff and tend to break when bent too much

## Slide 20: Preparation of the wire grid
Winding machine
Grid pattern can be machined onto guiding rails
With a program the linear motor moves while the winding machine winds the grid
Creates 4 grids in one go
Used for tungsten wires

## Slide 21: Preparation of the wire grid
Winding machine
Grid pattern can be machined onto guiding rails
With a program the linear motor moves while the winding machine winds the grid
Creates 4 grids in one go
Used for tungsten wires

## Slide 22: Preparation of the wire grid
Winding machine
Grid pattern can be machined onto guiding rails
With a program the linear motor moves while the winding machine winds the grid
Creates 4 grids in one go
Used for tungsten wires

## Slide 23: Components of a SEMGRID HEAD
Support structure
BIAS
Wire support
Wires
Connector
“cabling”

## Slide 24: Cabling
Single kapton isolated wires wired from the connector to the grid
Issues of cable/pin swapping
Can become a bit messy

## Slide 25: Cabling
FLEX-PCB for cabling
Easy to design
Easier to replace in case of issues
Easy to produce

## Slide 26: Cabling
Single and double layer
APICAL NP polyimide film as base
Copper finish of 30 μm with a 5μm base
Using empty polyimide films to protect against shorts with the vacuum vessel or each other
Performed vacuum acceptance tests at CERN and so far had no issues with them, all cleared for installation

## Slide 27: Components of a SEMGRID HEAD
Support structure
BIAS
Wire support
Wires
Connector
“cabling”

## Slide 28: Connector
Flex PCB requires a different kind of connection than single wires on both the PCB and the feedthrough side
Connectors for both side have been developed and used at CERN

## Slide 29: Connector feedthrough side

## Slide 30: Connector feedthrough side
Macor® piece connects onto D-Sub 50
Cut crimp contact
Macor® counterpiece
Spring loaded contact
FLEX-PCB

## Slide 31: Connector feedthrough side
Macor® piece connects onto D-Sub 50
Cut crimp contact
Macor® counterplate
Spring loaded contact
FLEX-PCB

## Slide 32: Connector feedthrough side
37 pin round connector

## Slide 33: Connector feedthrough side
37 pin round connector

## Slide 34: Connector PCB side
Flex PCB folds and pads make direct contact with pads on PCB

## Slide 35: Connector PCB side
Shell made out of aluminium
RF finger strip mounted on a support piece (Al)
Fully assembled contact strip for one side

## Slide 36: Connector PCB side
Clamped onto the PCB with the Flex

## Slide 37: DAQ
Machine | Current range | Gain range | ADC type | Comments
Linear amplifier (LINAC) | 1 μA -> 10 mA | 1/10 -> x1000 | Samples: 4 μs | A linear line
Integrator amplifier (circulating) | 2.5 pC -> 250 nC | X1->x100’000 | Samples: 4 μs | Bunch beam
TbT Booster (linear) | 1 μA -> 100 μA | x1 -> x100 | Samples: 240 ns | Turn by turn measurement
TbT PS (linear) | 10 nA -> 1 mA | 1/100 -> x1000 | Samples: 240 ns | Turn by turn measurement
SPS VME acquisition 2 mode | Linear: 50 nA -> 20 uA Integrator: 1nC -> 400 nC | Linear: 1, 20 ,400 | Samples: 20ms | Use integrator for injection to SPS Use linear for Experiment extraction

## Slide 38: Sticking wires
Issue of gold coated tungsten wires “sticking” to each other observed in LINAC 4 after being used in beam
Due to beam hitting the wire the wires start heating up
Thermal expansion of the wire causes them to lose part of the tension
We suspect tungsten wires become magnetic due to the beam and attract each together
Wires bond together, temperature at bonding point must be higher than 1100°C as that is the melting point of gold
Once the heat is gone they are stuck together

## Slide 39: Solution ?
Create an additional pocket on the PCB after the solder
Apply some force which will act on the wire when it is heated up and loses tension to re-tension
Design exists needs to be tested with beam
PCB
Solder point
Tungsten wire

## Slide 40: Special Thanks to
Patrick Bouvier, Michel Duraffourg, Guilhem Elias, Roman Riac, Federico Roncarolo, Jean Tassan-Viol, Raymond Veness,

## Slide 41: Any questions? Comments? Similar experiences? Additional ideas?