## Slide 1: Beam instrumentation based on Electro-Optical detection
T. Lefevre for BI group
EP-ESE & SY-BI collaboration meeting - 14 June 2021

## Slide 2: Outline
Electro-optical detection
Commercial products and R&D work
On-going work in SY-BI
Future needs and possible synergies with EP-ESE

## Slide 3: Electro-optical detection in Beam instrumentation
Aim : Convert Coulomb field of e-bunch or beam induced voltage signals into an optical intensity variation
How: Use the polarization/phase change of a laser beam which passes through a birefringent crystal itself polarized by the electric field of charged particle beams
Pockel/Kerr effect
particle beam
Coulomb field
probing laser pulse
EO crystal
Relative phase shift between polarizations
increases with the beam electric field

## Slide 4: Electro-optical detection in Beam instrumentation
Commercial E-O modulators used in many applications in laser technology and telecommunications
Phase/intensity modulator
without/with polariser
Intensity modulator
using Mach-Zehnder interferometer
IQ Intensity modulator
using Mach-Zehnder interferometers

## Slide 5: Electro-optical detection in Beam instrumentation
Commercial E-O modulators used in many applications in laser technology and telecommunications
LiNbO3, LiTaO3, GaAs, BGO, ..
Large bandwidth from DC to 50GHz

## Slide 6: Electro-optical detection in Beam instrumentation
R&D on E-O techniques in accelerators for short bunch length monitoring (sub picosecond time response)
ZnTe
GaP
Thinner crystal - need higher power (short pulse) laser

## Slide 7: Electro-optical detection in Beam instrumentation
R&D on E-O techniques in accelerators for short bunch length monitoring (sub picosecond time response)
Encoding
Time to Frequency
Decoding
Frequency from Position
Time
Position
Spectral
Decoding
Wilke et.al., PRL 88 (2002) 124801

## Slide 8: E-O systems in SY-BI
Spectral decoding EO system at CLEAR for short bunch length measurement
Er laser 780nm
150fs – 12ps
Spectrometer with grating and intensified gated CCD camera
ZnTe Crystal in between crossed polarisers
System being simplified for integration in AWAKE and upgraded for measuring 1ps long bunches
S. Mazzoni and E. Senes

## Slide 9: E-O systems in SY-BI
EO development for high bandwidth (20ps resolution) intra-bunch beam position monitor at LHC
Button with custom brazed electrode
  In-air fiber based EO waveguide
  encoding manufactured by UK industry
DC Laser at 780nm
Collab. with RHUL
CST simulations

## Slide 10: E-O systems in SY-BI
EO BPM measurement for DC beams in SPS
x
z
EOpt
Transmission
direction
Transmission
direction
Photodetector
DC Bias
Data processing system
Modulation
signal
Feedback signal
External field
to measure
LiNbO3
LiTaO3
Laser at 1550nm
A. Cristiano and M. Krupa

## Slide 11: E-O systems in SY-BI
All systems aim at fiber based solutions
Possibly generalising time stretch techniques (spectral decoding) to improve the time resolution

## Slide 12: EO detection system for LHC
25ns
LHC Beam signal
40GHz EO modulator / PU
Prism + 40 photo-detectors (600MHz)
High-speed spectrometer
dt = 25ps
1ns
25ns
2ns
laser

## Slide 13: EO detection system for LHC
25ns
LHC Beam signal
40GHz EO modulator / PU
Prism + 40 photo-detectors (600MHz)
High-speed spectrometer
dt = 25ps
1ns
25ns
2ns
40MHz, 80mW
1560nm Laser system
Dispersive / fiber bragg grating
stretching
Laser temporal manipulation
Pulse duration 50 fs 2.5 ns
Peak power 16 MW 200 mW (30%loss)
laser

## Slide 14: EO detection system for LHC
25ns
LHC Beam signal
40GHz EO modulator / PU
Prism + 40 photo-detectors (600MHz)
High-speed spectrometer
dt = 25ps
1ns
25ns
2ns
40MHz, 80mW
1560nm Laser system
Dispersive / fiber bragg grating
stretching
Laser temporal manipulation
Pulse duration 50 fs 2.5 ns
Peak power 16 MW 200 mW (30%loss)
40 ADCs
(16bits, 40MSa/s)
Large FPGA
digesting a data flow of 25.6Gbit/s
Signal
Conditioning
laser

## Slide 15: EO detection system for LHC
25ns
LHC Beam signal
40GHz EO modulator / PU
Prism + 40 photo-detectors (600MHz)
High-speed spectrometer
dt = 25ps
1ns
25ns
2ns
40MHz, 80mW
1560nm Laser system
Dispersive / fiber bragg grating
stretching
Laser temporal manipulation
Pulse duration 50 fs 2.5 ns
Peak power 16 MW 200 mW (30%loss)
40 ADCs
(16bits, 40MSa/s)
Large FPGA
digesting a data flow of 25.6Gbit/s
Signal
Conditioning
laser
Equivalent Bandwidth of 40GHz
Time resolution of 25ps
Sensitivity would depend on EO modulator/photo-detectors noise/DR
Measuring all bunches (40MHz) - Ok
Acquiring larger number of turns – depending on DAQ

## Slide 16: EO detection system for short bunches

## Slide 17: Future…a fully integrated..
High rep rate chirped Fiber laser…. Stretching
…Custom fiber coupled EO modulator… Stretching
…Dispersive Grating…
… linear array of Si/InGaAs photosensor…
….ADC/FPGA

## Slide 18: Thanks for your attention

## Slide 19
100GSa/s -16bits (E-O) Oscilloscope (250kCHF)
100 ADCs
(16bits, 1GSa/s)
read-out system
FPGA digesting the data flow 100x16x109 = 1.6Tbit/s
Signal to measure
50GHz EO modulator
40mW CW laser with periodic chirping
1ns 1ns 1ns
100MHz, 80mW
1560nm Laser system
Prism + linear photo-detector array
High-speed spectrometer
100 channels
Output signal < 0.3 mW (0.3mA) per detector
10ps time resolution
1ns
Chirping
Dispersive
stretching
Laser temporal manipulation
x 10
Splitting and
Frequency multiplication
Longitudinal
shaping
50 fs 5 ps 1.5 ns 1 ns Laser pulse duration
16 MW 140 kW 600 mW 400 mW Laser peak power
Signal
Conditioning
50kCHF
20kCHF
20kCHF
40kCHF
120kCHF

## Slide 20
Laser for E-O detection at LHC
Chart: Optical Spectrum
Central wavelength: 1560nm +/- 20nm
Spectral range(typ. <3 dB): ~ 20nm;
Output average power: 130mW (linear amplifier)
Repetition rate: 100MHz +/- 1MHz
Pulse width: prechirped to 1ns using 2 fiber bragg gratings
50kCHF with synchronisation
ΔFrev of few kHz

## Slide 21
Time stretching for E-O detection at. LHC
Dispersive
stretching
Laser temporal manipulation
Study time streching methods
  Dispersion in long fibers
  Fiber bragg grating
  Combination of both
Characterise and optimise : spectral flatness vs peak power
  Work can start re-using the CLEAR EO laser system

## Slide 22
Linear photodetector atLHC
Would also need to launch price enquiry : Need to work on specifications first

## Slide 23
Linear photodetector for HT on LHC
Would also need to launch price enquiry : Need to work on specifications first

## Slide 24
DAQ for HT at LHC
40 ADCs
(16bits, 40MSa/s)
FPGA
digesting a data flow of 25.6Gbit/s
Signal
Conditioning
What exists on the Market ?
96 transceivers
memory bandwidth of up to 512 GBps
12k
48 transceivers
memory bandwidth of up to 128 Gbps
4k
Arria 10 GX 660
ADS42JB69IRGCT
16 Bit ADC
2 Inputs
220k/pieces