## Slide 1: Development of a ChDR BPM for the AWAKE experiment - Status
Eugenio Senes for the BI AWAKE & Cherenkov teams
… with many thanks to CLEAR OP for the support !

## Slide 2
Outlook on the tests
Test of ø6 and ø15 mm buttons
Horn antennas vs waveguide transitions
Waveguide transition tests & status
Outline
Lessons learned
Transition design status
Conclusions

## Slide 3: April 2021 test campaign at CLEAR
The initial plan Detection in the bunker
PRO: short waveguides
CON: inaccessible

## Slide 4: April 2021 test campaign at CLEAR
The initial plan Detection in the bunker
PRO: short waveguides
CON: inaccessible
Final solution Detection in the klystron gallery
PRO: accessible, troubleshootable
CON: >10 dB attenuation
Too difficult to access and fix ….
Needed a lot of fixing …

## Slide 5: Detection scheme: horns
Detection with horn antennas
OSCILLOSCOPE
28 dB
500 MHz
Schottky diode zero-bias
Tektronix
6 GHz
ø6mm button
Horn antenna
f0 29.986 GHz BW 300 MHz
BEAM
Waveguide network
15m 10 dB
BUNKER
PRO: - Simple(r)
CON: - Hard to align
- Requires RF shielding

## Slide 6: Test description
Both ø6 and ø15 mm radiators were tested
Mostly charge 80-140 pC
Different bunch lengths
Very different attenuations from the two buttons
  0.3 dB difference @30GHz in WG
  Buttons, transitions …
  Diodes …
A ’sweet spot’ in the configuration exists … however charge drift is also a problem
Response after charge
compensation
σ = 3.8 ps
143 ± 2 pC
~3 dB total

## Slide 7: Response to long and short bunches
SHORT BUNCH
RESPONSE
σ = 1.1 ps
81 ± 3 pC
Raw signal
Raw signal
15 mm buttons.
Expectation for Gaussian bunches
- 32% signal for long bunch
LONG BUNCH
RESPONSE
σ = 4.8 ps
89 ± 3 pC

## Slide 8: Detection scheme: transitions
OSCILLOSCOPE
28 dB
500 MHz
Schottky diode zero-bias
Tektronix
6 GHz
ø6mm button
Transition ø6 mm – WR28
f0 29.986 GHz BW 300 MHz
BEAM
Waveguide network
15m
10 dB
BUNKER
Detection with transitions
PRO: - Intrinsically shielded
- Couples larger signal
CON: - Custom part
- Early prototype
Alumina-loaded circular waveguide ø6 mm to WR 28 transition

## Slide 9: Test description
Transition available only for radiators ø6 mm
Mostly charge 80-120 pC
Response after charge
compensation
σ = 3.8 ps
121 ± 15 pC

## Slide 10: Horn antenna vs transition
HORN ANTENNA
429±65 pC
WAVEGUIDE TRANSITION
121±15 pC
To get similar
output voltage

## Slide 11: Lessons learned
Different channels might present rather large attenuation differences
  Buttons asymmetry
  Transitions asymmetry
  Waveguide length difference
  Unknown assembly tolerances
Reduced dependence on bunch length
  Expected 32% difference in power
  Bunch shape might play a role
Dramatic signal increase with transitions
  Compared to horn antennas, same signal with factor ~4 charge difference
  Early transition prototype, still reflecting ~40% of the power

## Slide 12: Transition designs
Large simulation work done by Collette
CST Studio Suite 2020
Level
0
No transition
Ptra = 53%
1
Single dielectric
Ptra = 64%
2
Multiple dielectrics
Ptra = 99%
Just showing worst case, details in a separate talk of Collette

## Slide 13: Conclusions
A design with waveguide transition is desirable
  Alumina buttons ø6 mm are the preferred choice
  Button + alumina is vacuum tight
Important effort in transition R&D
  Modular design: the transition slides onto the alumina, outside the vacuum
  Large efforts to understand the impact of tolerances
Studies pending
  Possible additional test campaign at CLEAR for stability studies
Brazed collar

## Slide 14: ChDR BPM project summary

## Slide 15: ChDR BPM project summary
BPM412346 (ChDR #1)
  Beampipe resizing done
  Support in manufacturing (end of June)
  Body installation once tested with buttons
  Buttons final design pending (end of June)
BPM412352 (ChDR #2)
  Body installed
  Buttons final design pending (end of June)

## Slide 16: ChDR BPM project summary
BPM412351
  Will detect both p+ and e-
  Discussions pending for signal splitting

## Slide 17: 40 GHz buttons option
In the past the use of high frequency buttons was discussed
CST studio
M. Wendt
Mitigate the risk on ChDR BPM project
Compare the two technologies
This is an alternative, additional R&D project, and not a fallback solution
CST simulations showed satisfactory performance with respect to ChDR BPM
Electrons 100 pC, 4 ps-σ

## Slide 18: 40 GHz buttons option
Valid quote for the buttons for 2 more weeks
Body to be designed
This is an alternative, additional R&D project, and not a fallback solution
ChDR BPM #1
HF BPM
Proposal:
Procure min 5 buttons and 1 body
Install in the other free drift close to fire door
Cost in the order of 50kCHF
Minimum 12 months before installation

## Slide 19: Thanks for your attention !