## Slide 1: Advances in Magnetic Measurements – Avatars and Twins
Stephan Russenschuck, 07.2021

## Slide 2: The Measurement Laboratory 311
Air-condition system (1.4 + 0.4 MCHF)
Avoid stratification (laser applications)
Air speed < 0.4 m/sec (laminar)
+- 2 oC in large hall, +- 0.4 oC in “measuring room”
Power converter cluster,
motorized switches,
PLC interlocked distribution
feedboxes

## Slide 3: Environment: We Have Come a Long Way
ISR-Tunnel Segment: The most stable environment at CERN

## Slide 4: The Measurement Paradigm
Measurements for design
  Material properties and design choices
  Software and model validation
  Dynamic effects
Measurements for magnet production
  Prototype qualification
  Quality assurance (error tracking); acceptance testing
  Magnet to magnet reproducibility (check tolerances)
  Shimming
  Requalification after repair
Measurements for accelerator operation
  Field description (feed forward compensation)
  Field quality (resonances)
  Online monitoring (pre-cycling)

## Slide 5: Magnetic Measurement Techniques
Rotating coil systems
3D Hall mapper
Stretched-wire
Stationary induction-coil
magnetometers

## Slide 6: The Avatar and Twin (1)
Test and magnetic
measurement
Raw signal
(integrated
voltages)
Processed
signal (harmonics,
flux density, VI)
Physical object (magnet)
Capture noise
Read-out noise
Intrins., approx., gross errors
Num. magnet model (design)
Materia/geom. parameters
Modeling errors
?
Pre-processing
(meshing)
Matrix solving
Post-processing
Discretization errors
Numerical errors
Approximation errors

## Slide 7: Beam Physicists, Magnet Designers, and Measurement Engineers
Everybody believes in measurements, but the measurement engineer himself.
Nobody believes in field simulations, but the field computation expert himself.
Establish a C3 coherence between requirements for the three parties

## Slide 8: BEM-FEM Coupling and Dynamic Effect Calculation in ROXIE
10-4 predictive properties

## Slide 9: 2D Field Harmonics
2D Field map Good field region

## Slide 10: Rotating Coil Measurements

## Slide 11: Reparametrization to Arc Length
Low drift, low-noise amplifier = resolution of 10 nVs
One cycle = 1 – 10 s
1024 trigger points for discrete Fourier transform

## Slide 12: Rotating Coil Systems
380 mm
8 mm

## Slide 13
Sensors and transducers for magnetic measurements: rotating coils
Horizontal stations room temperature
Optical targets installed at both coil ends  measurement of the rotation axis

## Slide 14
Using the FAME system and the FFMM framework with acquisition rates of up to 8 Hz
Measurements of Dynamic Effects in SC Magnets
Superconducting magnetization decay and snapback on non-allowed multipole a3 (at 1.8 K)
Flux jumps at 4.5 K

## Slide 15: The Avatar and Twin (2)
Test and magnetic
measurement
Raw signal
(integrated
voltages)
Processed
signal (harmonics,
flux density, VI)
Physical object (magnet)
Capture noise
Read-out noise
Intrins., approx., gross errors
Num. magnet model (design)
Modeling errors
RMS
Pre-processing
(meshing)
Matrix solving
Post-processing
Discretization errors
Numerical errors
Approximation errors
Num. model (as built)
Materia/geom. parameters
Parameter identification (non-linear optimization)

## Slide 16: Tracing manufacturing errors

## Slide 17: Uncertainty Mitigation Measures (Systematic Errors and Gross Errors)
Calibration: to establish the transfer function of a given instrument by comparison with a reference instrument of known accuracy
  cross-checks with instruments of comparable uncertainty to gain confidence or eliminate gross errors
  in-situ calibration: use the magnet to be tested as the reference via a previous measurement with a reference instrument
Traceability: an unbroken chain of comparisons, each with a stated uncertainty, from a measurement to a primary standard.
  Fundamental concept to certify a measurement; maintaining systematic records, databases, documented procedures

## Slide 18
MM requests
Kanban Board
NORMA
MTF
EDMS
MM Equipment
Calibration data
MM report
FSU job request
Operations management
Database “management”
Asset management
Project management and
Business intelligence
Released data
FFMM config.
Raw data
Processed data
Material
Kn
Traveller
NC magnet
SC magnet
Bench preparation
Client
Client
The Problem Setting

## Slide 19: Mitigation Measures (Systematic and Gross Errors)
Flexible C++ Framework for Magnetic Measurements
Web-services
(XML (Soap))
Text-file
SQL

## Slide 20: Calibration: CERN Dipole Field Reference
Old CERN ISR bending dipole, cycled up to 1 T in a rigorously reproducible way
Yearly mapping to establish field profile, averages and reproducibility
We map By, but since the field is not perfectly uniform we have also Bx, Bz  the resulting (small) error is propagated to all instruments calibrated in this reference !
high precision NMR probe used to map the field

## Slide 21
Sensor and PCB coil calibration

## Slide 22
3D-Helmholtz-Coil
Biot-Savart's law for computing the magnetic flux density
Effect of construction errors on field < 10-6
α
ca
cf
Parallelism
Flatness

## Slide 23: Magnetic Measurements with Hall Sensors in Cryomagnet
PCB with 5 coils and 5 Hall probes mounted in magnet aperture

## Slide 24: Cold-Calibration of Hall Sensors
Calibration to compensate the transversal field inhomogeneity, fn(B) = F
10 % !

## Slide 25: Cross-Calibration by Integrating the Induced Quench Voltage
Integrator drift: 0.5% over 60 s on average

## Slide 26: Stair-Step Cycles (On Central Hall Sensor)
Induction and damping of interstand coupling currents in high TC superconducting magnet

## Slide 27: Riemann Lebesque Lemma
Always plot your results in logarithmic scale
Simulations
Measurements
The Fourier coefficients tend to zero as n goes to infinity

## Slide 28: Going Big (FAIR SuperFRS multiplet)
Diameter 380 mm

## Slide 29: Geometric Oversampling

## Slide 30: Combining Radius Scaling and Feed-Down (Analytical Continuation)
c
c

## Slide 31: Matrix Conditioning
N Number of multipoles in small coil
I Number of measurement positions
For 15 multipoles reconstructed on r0
rc/r0
k

## Slide 32: Results: 8 Positions, 8 Harmonics on the Small Coil

## Slide 33: The Avatar and Twin (3)
Physical state
Regularity conditions of the
magnetic field
Magnetic flux
density in W
Observation function
Harmonics on large radius
Test and magnetic
measurement
Raw signal
(integrated
voltages)
Processed
signal (harmonics,
flux density, VI)
Physical object (magnet)
Capture noise
Read-out noise
Intrins., approx., gross errors
Num. magnet model (design)
Modeling errors
Field inversion
Overdetermined
(Mapping)

## Slide 34: Going small

## Slide 35: Stretched-Wire Systems

## Slide 36: Stretched Wire Measurements
Oscillating wire
Vibrating wire
Map of Second Wire Resonance
Stretched wire

## Slide 37: Solenoid Center and Axis (Vibrating)
Force distribution requires vibration at the second resonance too add up consistently

## Slide 38: Linear Regression in the Line Search Procedure
Resolution in the mm range
Uncertainty 20 mm

## Slide 39: Multipole Measurements (Oscillating)
Measure the oscillation amplitudes on K rulings of a cylindrical domain.
Develop into a Fourier series.
Compare coefficients

## Slide 40
Local Field Distribution

## Slide 41
Pseudo-Multipoles (Fourier Bessel Series)

## Slide 42: The Leading Term is NOT the Measured One

## Slide 43
Search Coils Must be Isoperimetric

## Slide 44: The Avatar and Twin (4)
Physical state
Boundary Sources (FEM / BEM), Snapshots)
Magnetic flux
density in W
Observation function
(Voltages, flux increments) including uncertainties
Test and magnetic
measurement
Raw signal
(integrated
voltages)
Processed
signal (harmonics,
flux density, VI)
Physical object (magnet)
Capture noise
Read-out noise
Intrins., approx., gross errors
Num. magnet model (design)
Properties of the Equation system | Method | Application
Diagonal | Orthogonal proj. | Harmonic fields
Square | Collocation | Boundary data
Overdetermined | Least-squares | Mapping
Underdetermined | Bayesian inference | Sensor fusion
Modeling errors
Field inversion
Avatar

## Slide 45
The Avatar of the Faser Magnet

## Slide 46: Grid Map versus Boundary Data

## Slide 47
Mid-Plane Field

## Slide 48: The Moving Fluxmeter
For FAIR
For NA64

## Slide 49: Avatar (Generalized Field description)

## Slide 50
Solenoidal Field Transducer

## Slide 51: Combined Function Magnets (CERN PS)
4 pole face windings
affecting B1 to B5
Figure-of-eight loop
to control B2
Strongly coupled
excitation circuits
No 10-4 predictive model
Remanent field  0.2%
FMR field
marker
Fluxmeter
for gradient
measurements

## Slide 52
  CNGS
  CNGS
  CNGS
PS Beam Stability Test at Injection
Specific powering cycles (CNGS) lead to reproducible radial positioning errors
from relative field changes on the order of 10-5

## Slide 53: The Avatar and Twin (5)
Physical state
Boundary Sources (FEM / BEM), Snapshots)
Magnetic flux
density in W
Observation function
(Voltages, flux increments) including uncertainties
Test and magnetic
measurement
Raw signal
(integrated
voltages)
Processed
signal (harmonics,
flux density, VI)
Model Order Reduction
Physical object (magnet)
Capture noise
Read-out noise
Intrins., approx., gross errors
Num. magnet model (design)
Modeling errors
Mode updating by proper orthogonal decomposition
Design, constr, calibration of field transducer
Model-based system engineering
Value-shop model and asset management
Avatar
Twin
Model assurance criterion

## Slide 54
Twin (data-driven modelling)
Stefano Sorti
simulated
reconstructed