## Slide 1
State of the art of Positron Emission Tomography:
technology and current challenges
Dr. Maurizio Conti Director, PET Physics and Reconstruction
Siemens Medical Solution USA, Inc, Knoxville, TN, USA
Presented at CERN, Geneva, September 6, 2021

## Slide 2: Introduction to myself
Knoxville, TN

## Slide 3
What is Tennessee famous for?

## Slide 4: Table of contents
TOF PET methods and applications
Long Axial FOV scanner and applications
Evolution of PET and state of the art of PET scanners today
Positron Emission Tomography (PET)
The role of AI in PET
Future technology development and needs for fundamental research

## Slide 5: Table of contents
TOF PET methods and applications
Long Axial FOV scanner and applications
Evolution of PET and state of the art of PET scanners today
Positron Emission Tomography (PET)
The role of AI in PET
Future technology development and needs for fundamental research

## Slide 6
P N + e+ + n + energy
A radioisotope b-decays by emitting a positron, a neutrino. Positron energy is typically 0-3 MeV, half-life typically 1-120 minutes.
The positron and an electron annihilate producing two 180º opposed 511 keV gamma rays.
The PET detector system detects the annihilation by detecting the occurrence of two gamma rays within a few nanoseconds.
The positron annihilation occurred somewhere along the line connecting the two detectors.
Positron Emission Tomography

## Slide 7
Radiopharmaceutical | # of Times Scanned | % of Total
Fluorodeoxyglucose (F-18) | 346442 | 90.68%
Rubidium chloride (Rb-82) | 10211 | 2.67%
Fluorethylcholine (F-18) | 5087 | 1.33%
Germanium (Ge-68) | 1790 | 0.47%
Gallium Solution (Ga-68) | 1468 | 0.38%
Fluoro-L-dopa (F-18) | 1264 | 0.33%
DOTA-TATE (Ga-68) | 1241 | 0.32%
Ammonia (N-13) | 1233 | 0.32%
Sodium Fluoride (F-18) | 1111 | 0.29%
Oxygen (O-15) | 938 | 0.25%
Solution (F-18) | 823 | 0.22%
Acetate (C-11) | 806 | 0.21%
Choline (C-11) | 787 | 0.21%
DOTA-TOC (Ga-68) | 726 | 0.19%
Fluorethyltyrosin (F-18) | 723 | 0.19%
Methionine (C-11) | 606 | 0.16%
Solid Solution (Ge-68) | 598 | 0.16%
Crossed Rods (Ge-68) | 550 | 0.14%
PSMA (Ga-68) | 442 | 0.12%
18F-FDG
(fluorodeoxyglucose)
Use of PET, 2013
Main application: oncology
Measuring Glucose Metabolism
Others: brain studies, cardiac, etc.
Tracers used in PET

## Slide 8
Detector block
PMT

X1 = 0.95 / (0.95 + 0.05) = 0.95
X2 = 0.55 / (0.55 + 0.45) = 0.55
A
B
0
1
PET detector

## Slide 9
PET events are distributed across a range of energy, not only in the 511 keV range. An energy window is employed to qualify good events and reject scattered events.
ENERGY
WINDOW
Counts
Energy (keV)
0
100
200
300
400
700
600
500
SCATTER
511 keV PHOTONS
Lower Level Discriminator (LLD) 435 keV
Upper Level Discriminator (ULD) 650 keV
Energy and time qualification of a detected photon
Positron (e+)
511 keV gamma rays
Is
EventTime A-B less then
4.06 ns
Coincident Event
(belongs to the same positron annihilation)
Yes
Event Time A
Event Time B

## Slide 10
Sinogram
Radial offset
Angular Offset
angle
Radial position
Data organization (sinogram)

## Slide 11
Algorithms for reconstruction:
analytical (back projection FBP, etc) and iterative reconstruction (MLEM, OSEM, MAP, RAMLA, etc)
How to make an image (the principle)
But the data need “corrections”…

## Slide 12
Normalization
A
B
C
D
Normalization array
Each detector has a different response: A calibration or normalization is needed
A
B
W
detector
detector
(e
–µA
)(e
–µB
) = e
–µ(A+B)
= e
–µW
Photons are scattered and absorbed:
A correct quantification of the original source needs attenuation correction
Data correction (1)
Attenuation correction

## Slide 13
True Coincidence
Random Coincidence
Scatter Coincidence
Data correction (2)
Randoms correction
Scatter correction

## Slide 14
Emission Sinogram
m-map (CT or MR based)
Attenuation correction
Scatter correction
Normalization
Scatter simulation
Randoms
How to implement data corrections (1)

## Slide 15
=
_
A x [Nx(E-R) – S]
A x S
recon
recon
X
=
A x [N x (E-R)]
A
E-R
X
=
N
N x (E-R)
recon
Non-attenuation corrected image
Attenuation corrected image
Attenuation and scatter corrected image
How to implement data corrections (2)

## Slide 16: Table of contents
TOF PET methods and applications
Long Axial FOV scanner and applications
Evolution of PET and state of the art of PET scanners today
Positron Emission Tomography (PET)
The role of AI in PET
Future technology development and needs for fundamental research

## Slide 17
Time-of-Flight PET (TOF) – two photons in coincidence

## Slide 18
TOF reconstruction = sensitivity amplification or noise reduction
D is patient “diameter”, Dt is time resolution of TOF PET scanner
*estimate based on uniform cylindrical object and analytical reconstruction:
T.F. Budinger, “Time-of-flight positron emission tomography: status relative to conventional PET”, J. Nucl. Medicine, vol. 24, pp. 73-78, 1983.
M. Conti, "State of the Art and Challenges of Time-of-Flight PET", Physica Medica, no. 25, pp. 1-11, 2009.
L. Eriksson and M. Conti, “Randoms and TOF Gain revisited”, Phys. Med. Biol. 60 1613-1623 (2015).
Dt (ps) | Dx (cm) | GNEC (D=40cm) | GSNR (D=40cm)
200 | 3 | 13.4 | 5.2
300 | 4.5 | 8.9 | 3.0
400 | 6 | 6.7 | 2.6
500 | 7.5 | 5.3 | 2.3
600 | 9 | 4.4 | 2.1
Time-of-Flight PET advantages: noise reduction

## Slide 19
same time,
better image
3 minute/bed
550 ps
non-TOF
TOF
Time-of-Flight PET advantages: noise reduction and improved contrast recovery

## Slide 20
  Multiple lesions (<2cm) analyzed in 30 patients
  Body Mass Index used as a measure of the size
Time-of-Flight PET advantages: noise reduction and patient size

## Slide 21
50 cm diameter water cylinder with cylindrical inserts.
Variable contrast to background.
m-map was assumed uniform, and normalization
was incorrect.
TOF
non-TOF
Iter 1
Iter 2
Iter 3
Iter 4
Iter 5
13 mm
25 mm
51 mm
51 mm
101 mm
Inaccurate m-map and normalization
Time-of-Flight PET advantages: more robust with inconsistent data

## Slide 22: Time-of-Flight PET and the low counts challenge: improved reproducibility
Reliability of quantification at low counts:
variation between realizations
Geant4 simulation
Contrast recovery coefficient (CRC) vs. noise
17 mm sphere and 600 ps time resolution
Plotting 10 realizations and mean value
40x106 counts
1x106 counts
0.05x106 counts
Bgrnd
Signal
V. Westerwoudt, M. Conti and L. Eriksson, “Advantages of improved time resolution for TOF PET at very low statistics”, TNS 61, 126-133(2014)
noise

## Slide 23: Time-of-Flight PET and the low counts challenge: improved reproducibility
Line = mean value over 10 realizations
Color band = +/- 1 standard deviation
At very low counts, the uncertainty of the quantification gets considerably smaller with better time resolution
Typical FDG patient statistics today (30-40 x 106 counts)
Contrast recovery
10 mm sphere
1/sqrt(counts)
V. Westerwoudt, M. Conti and L. Eriksson, “Advantages of improved time resolution for TOF PET at very low statistics”, TNS 61, 126-133(2014)

## Slide 24
    Use imaging to plan therapy, monitor therapy, assess response to therapy.
    Personalize the medical treatment.
Theranostics
mAb imaging
Theranostics
a targeting agent which specifically binds to a molecular target on the surface of a cell or tumor
an imaging agent and a therapeutic drug/radioisotope pair
a linker which serves to connect the two entities
More…
Small compound imaging
New TOF PET technology as enabler for new applications: beyond “diagnostic-only” imaging
Y90
Hadron therapy

## Slide 25: Time-of-Flight PET and the low counts challenge: Y90 imaging
90Y is used as a therapeutic radioisotope in microspheres for targeted radiation therapy of the liver. Very weak PET signal from e+e- pair creation (BR=30x10-6)
PET scan characteristics:
long (30 minutes)
low number of counts [105-106]
high random fraction from LSO background (>80%)
high flux of Bremsstrahlung singles in 90Y

## Slide 26: Time-of-Flight PET and the low counts challenge: Y90 imaging
Courtesy of T. Carlier, University Hospital, Nantes, France
non TOF
TOF
Extremely low counts:
Yttrium-90 imaging
Patient after Y-90 microspheres injection in the liver
(50000 counts, 95% random fraction)

## Slide 27
PET scan characteristics:
long (30 minutes)
low number of counts [105-106], while typical FDG PET is [107-108]
with high random fraction (>80%)
    PET imaging of irradiation with proton or 12C induces β+ emitter (15O, 11C, etc)
    Hadron and proton therapy are used for local treatment of tumors: liver, brain, eyes, prostate.
    Very weak PET signal from nuclei activation inside the tissue.
Courtesy of K. Parodi, Heidelberger Ionenstrahl-Therapiezentrum (HIT)
    Off-beam PET (about 5-10 minutes post-therapy)
Prediction
Measurement
Time-of-Flight PET and the low counts challenge: hadron therapy

## Slide 28
Courtesy of Heidelberger Ionenstrahl-Therapiezentrum (HIT)
OSEM, 21 subsets,
3 iterations
OSEM+PSF+TOF, 14 subsets, 2 iterations
Reconstruction methods need to be tailored/optimized for challenging applications
image size 200x200, 5 mm Gaussian filter
Time-of-Flight PET and the low counts challenge: hadron therapy

## Slide 29
    Matched imaging–therapeutic agents:
    68Ga-DOTA-PSMA 177Lu-DOTA-PSMA
177Lu
    Matched imaging–therapeutic radioisotopes: 124I and 131I, 86Y and 90Y
Theranostics: small compound imaging
Imaging
Therapy

## Slide 30
Monoclonal antibodies (mAb) can target a specific tumor cell surface marker.
They can be used to disease-selective imaging and disease-selective therapy.
Intact mAbs have a long biological time in humans (days), and they are the format of choice for therapy (pharmaceutical or radiotherapy).
The corresponding imaging radioisotopes must have a long half-life
Theranostics: mAbs and immunoPET
Diabody
55 kDa
Peak 1-2 h
Minibody
80 kDa
Peak 4-6 h
Intact antibody
150 kDa
Peak 50-100 h
Engineering Antibodies

## Slide 31
Number of therapeutic monoclonal antibodies entering clinical study during 1980 to 2006
Theranostics: mAbs and immunoPET

## Slide 32
    non-pure beta decay -> accurate prompt gammas corrections
    Long positron range -> range deblurring
    Fewer counts -> reconstruction optimization
Challenges of theranostics for PET physics: radioisotopes
Isotope | Half-life | Branching (b+) in % | Emax (MeV) | Emean (MeV) | Rmax (mm) | Rmean (mm)
11C | 20.4 min | 99.8 | 0.960 | 0.386 | 4.2 | 1.2
13N | 10.0 min | 99.8 | 1.199 | 0.492 | 5.5 | 1.8
15O | 2 min | 99.9 | 1.732 | 0.735 | 8.4 | 3.0
18F | 110 min | 96.9 | 0.634 | 0.250 | 2.4 | 0.6
64Cu | 12.7 hours | 17.5 | 0.653 | 0.278 | 2.5 | 0.7
89Zr | 78.4 hours | 22.7 | 0.902 | 0.396 | 3.8 | 1.3
Isotope | Half-life | Branching (b+) in % | b+ Emax (MeV) | Branching (g) in % | g E (PG) (MeV)
68Ga | 67.8 min | 87.7, 1.2 | 1.899, 0.821 | 3.2 | 1.077
76Br | 16.2 hours | 25.8, 6.3 | 3.382, 0.871 | 74.0, 15.9 | 0.559, 0.657
82Rb | 1.3 min | 81.8, 13.1 | 3.378, 2.601 | 15.1 | 0.777
86Y | 14.7 hours | 11.9, 5.6 | 1.221, 1.545 | 82.5, 32.6 | 1.077. 0.627
124I | 100 hours | 11.7, 10.7 | 1.535, 2.138 | 62.9, 11.2 | 0.602, 1.691
Prompt gamma emitters
Pure beta emitters

## Slide 33
F18 decay scheme
“simplified” Y86 decay scheme
Challenges: prompt gamma corrections

## Slide 34: Table of contents
TOF PET methods and applications
Long Axial FOV scanner and applications
Evolution of PET and state of the art of PET scanners today
Positron Emission Tomography (PET)
The role of AI in PET
Future technology development and needs for fundamental research

## Slide 35: The evolution of Siemens scanners
2000
2010
2017
2020
scintillator(LSO)
TOF (SW/HW)
Photosensor (SiPM)
Volume & data handling

## Slide 36: Biograph Vision – system overview
Biograph Vision
SiPM-based LSO detector with 3.2 x 3.2 x 20 mm crystals
82 cm transaxial diameter and 26 cm axial length
19 detector module assemblies with 16 detectors per assembly
60,800 total crystals. 760 per ring and 80 rings
Water-cooled gantry with 3 zones of cooling
Large 78 cm bore
64 or 128-slice CT
FlowMotion continuous bed motion

## Slide 37: A new approach to detector design with Biograph Vision™
3.2 mm LSO crystal element
2x4 mini -blocks
100% coverage1 of the LSO array
5x5 crystal mini-block
Coupled to a 4x4 array of 4.0 x 4.0 mm SiPMs.
2x8 detectors axially =26.3 cm axial FOV

## Slide 38: “SiPM TOF-PET Detector”, Sanghee Cho et al. (IEEE MIC 2017)

## Slide 39: Biograph Vision performance (IEEE MIC 2018)

## Slide 40: Biograph Vision performance

## Slide 41: Clear visualization of the 2.4 mm cylinders of the Mini-Derenzo Phantom with Biograph Vision
Biograph mCT
Biograph Vision
Diameter (mm)
1.2
1.6
2.4
3.2
4.0
4.8

## Slide 42: Table of contents
TOF PET methods and applications
Long Axial FOV scanner and applications
Evolution of PET and state of the art of PET scanners today
Positron Emission Tomography (PET)
The role of AI in PET
Future technology development and needs for fundamental research

## Slide 43: Biograph Vision Quadra
Multi-organ coverage
High axial coverage: 106 cm axial length
High sensitivity
High time resolution: ̴ 225 ps time resolution
High detection efficiency: 243200 LSO crystals
3.2 x 3.2 x 20 mm crystals
High spatial resolution (it doesn’t hurt )

## Slide 44: The evolution of Siemens scanners

## Slide 45
A TOF image is equivalent to a non-TOF image obtained with a larger number of counts: TOF reconstruction acts as a “virtual counts or sensitivity amplifier” for a PET scanner. The sensitivity gain is proportional to the size of the object that is being scanned (D) and the system time resolution :
TOF sensitivity gain or TOF effective sensitivity
NEMA sensitivity (70cm line)
Two “modes” available:
High Sensitivity and high resolution (HS, 18°)
Ultra High Sensitivity (UHS, 52°) [2nd software release)

## Slide 46: Applications of long axial FOV scanners (LAFOV) or Total Body scanners
The Present: today’s clinical applications
Faster scans:
  Faster workflow
  Patient comfort, particularly for sick patients
  Motion artifact reduction
Lower dose:
  Low dose PET
  Repeated scans to monitor therapy
  Pediatric PET
The Future: long axial FOV PET in research and new clinical applications
References:
S.R. Cherry, T. Jones, J.S. Karp, et al., “Total-Body PET: Maximizing Sensitivity to Create New Opportunities for Clinical Research and Patient Care”, J Nucl Med 2018; 59:3–12
S.Surti, A.Pantel, J.Karp, “Total Body PET: Why, How, What for?”, Transactions on Radiation and Plasma Medical Sciences, vol. 4, no. 3, pp. 283-292, 2020.
S. Vandenberghe, P. Moskal and J.S. Karp, “State of the art in total body PET”, EJNMMI Physics (2020) 7:35, pp: 1-33.

## Slide 47: The Future: LAFOV as enabler for new applications
PET/CT screening
Fast breath-hold PET/CT
Low dose immunoPET (Zr89 and others)
Pediatric imaging
High sensitivity
Interaction inter organs (brain-guts, brain-heart, brain-spine)
PET angiography and PET perfusion
Multi-organ coverage
Accurate whole-body parametric imaging
DV
SUV
Ki
Multi organ tracer kinetics and input functions

## Slide 48: The Future: LAFOV as enabler for new applications (1)
Fast breath-hold PET/CT
Breath-hold scans
The amplification of sensitivity compared to standard scanners like Vision, allows for standard dose but very short scans (how many seconds? TBD as part of the research) – this will provide:
  Patient comfort: sick patients cannot hold breath much!
  Full matching between CT and PET (sometimes because of the longer PET scan there is mismatch, which is problematic)
  Particularly important for lung cancer

## Slide 49: The Future: LAFOV as enabler for new applications (2)
PET/CT screening
Screening of patients at risk
Early detection of disease like lung cancer: this and other deadly diseases in which patients show symptoms when the cancer is in a very advanced state.
Population at risk can be defined for those diseases, and if dose can be reduced to “minimal dose” (TBD within the research) large scale screening of those populations can be done, and lives can be saved with early detection

## Slide 50: The Future: LAFOV as enabler for new applications (3)
Pediatric applications
For this application probably we will investigate simultaneous reduction of dose and scan time.
A x10 gain in sensitivity allows to reduce both (1/2 of the time and 1/5 of the dose, for example).
Dose reduction is needed for children
Short scans will reduce motion during the scan: children often need to be sedated to keep them still!
Pediatric imaging

## Slide 51: The Future: LAFOV as enabler for new applications (4)
Low dose immunoPET (Zr89 and others)
ImmunoPET and theranostics
PET imaging (immunoPET) can be associated with therapy (immunotherapy) to predict efficacy of the drug or the radioactive therapy agent, using the same molecule or antibody as a carrier for both (Theranostics).
Given the long biological half-life of monoclonal antibodies, long life radioisotopes like 89Zr are used, and the low positron abundance (23% for 89Zr) together with the slow decay, results in very low counts in the scan, and need for high sensitive PET scanners.

## Slide 52: The Future: LAFOV as enabler for new applications (5)
PET angiography and PET perfusion
Whole-body PET angiography and multiple organs PET perfusion
Fast and full body scans are feasible
Whole-body vascular disease can be imaged dynamically
Imaging of perfusion in multiple organs, simultaneously.

## Slide 53: The Future: LAFOV as enabler for new applications (6)
Interaction inter organs (brain-guts, brain-heart, brain-spine)
Interaction between organs
Brain-skeleton, research to develop new opioid drugs that control pain and avoid addition
Brain-guts, research associated with Parkinson’s
Brain-heart, investigation of amyloid deposits associated with Alzheimer’s
Brain-heart, studying cognitive impairment of people with heart attack

## Slide 54: The Future: LAFOV as enabler for new applications (7)
Accurate whole-body parametric imaging
kinetic/parametric parameters imaging tends to be “noisy” because of needing to fit very short frames. High sensitivity allows for more robust parametric images.
Accuracy in quantification of kinetic parameters is greatly improved
Image quality is improved as noise is reduced
Accurate whole-body parametric imaging
DV
SUV
Ki

## Slide 55: The Future: LAFOV as enabler for new applications (8)
Multi organ tracer kinetics and input functions
Multi organ tracer kinetics and input functions
Fast, low dose, multi organ tracer kinetics for drug development
We can study uptake kinetics in all organs in the same time, and also study the uptake in organs that are not target (collateral damage to healthy organs/tissue)
“local” image derived input functions can be computed and compared, input function dispersion can be modeled

## Slide 56: Table of contents
TOF PET methods and applications
Long Axial FOV scanner and applications
Evolution of PET and state of the art of PET scanners today
Positron Emission Tomography (PET)
The role of AI in PET
Future technology development and needs for fundamental research

## Slide 57
Credit: Vijay Shah
Publications on AI in Molecular Imaging

## Slide 58: Artificial Intelligence in PET
Detector and system
Improve & speed-up calibrations and normalization
Improve detector performance
Reconstruction
Corrections (attenuation, scatter, normalization, motion)
Aid & speed-up reconstruction
Low count reconstruction
Direct reconstruction
Post-reconstruction
radiomics
Image processing (denoising, registration)
Support to diagnostics and research

## Slide 59: Use of Deep Learning: denoising or low counts application
IEEE 2017 (Stanford)
IEEE 2018 (Singapore)
The idea: train the network with some high statistics training set
Applications: mAb, 124I, 90Y, low dose imaging, low dose PET screening

## Slide 60: Use of Deep Learning: alternative or faster reconstruction
TRPMS 2021 (King’s College London)
The idea: “direct” DL reconstruction from raw data, or mix of trained network and “conventional” methods
Applications: all applications, possibly faster or better reconstruction
TRPMS 2021 (Siemens)

## Slide 61: Use of Deep Learning: improved data correction
TRPMS 2021 (Seoul)
The idea: train networks to computer faster or better attenuation, scatter, motion correction
Applications: all applications, but in particular attenuation
EJNMMI 2018 (Madison)

## Slide 62: Use of Deep Learning: detector performance improvement
PMB 2020 (Munich, CERN)
The idea: train networks with controlled experiments to be able to extract extra information from signals
Applications: improve corrections for depth of interaction, inter crystal scatter, time of flight
PMB 2018 (UC Davis)

## Slide 63: Table of contents
TOF PET methods and applications
Long Axial FOV scanner and applications
Evolution of PET and state of the art of PET scanners today
Positron Emission Tomography (PET)
The role of AI in PET
Future technology development and needs for fundamental research

## Slide 64: Direct TOF back projection from list mode, or histoimage (simulation)
“cardiac defect“ simulation: 4-cm diameter ring, 1-cm thick , with on 1-cm defect
500 ps
200 ps
10 ps
What happens if we improve TOF resolution?

## Slide 65
No normalization, attenuation, scatter correction
Data Courtesy of University of Tennessee
Direct TOF back projection from list mode (aka histoimage), MIP image
An example of “no reconstruction”
mCT (550 ps resolution)
Vision (215 ps resolution)

## Slide 66: Do we still need reconstruction physicists?
Probably yes…
Consistency of space-time information is enforced by iterative reconstruction, together with noise reduction
TOF backprojection or histoimage (200ps)
OSEM TOF reconstruction (200 ps, 4i5s)

## Slide 67: How to improve time resolution?
new scintillators
new materials with fast emission
new detector geometries (monolithic detectors, side readout, plates, etc)
Cherenkov light
Alternative detector systems (MCP etc)
Deep learning to improve performance
…

## Slide 68
New and improved detectors and architectures
Better time and spatial resolution
Mathematics and reconstruction for challenging new applications (extremely low counts as in cell tracking, motion correction and dynamic studies)
Technology for multidisciplinary teams (physician, biologist, physicist, pharmacologist…
My conclusive remarks and questions:
The great technological leap of the past few years provided powerful tools for fundamental research
The future belongs to biological research, medical research, pharmaceutical research
But, is there still space for groundbreaking technological innovation?
Yes! Oui! Si! Ja! Da! Tak ! …

## Slide 69: Thanks to:
  Maurizio Conti
  Director, PET physics and reconstruction
  Phone: +1 865 548-2443
  maurizioconti@siemens-healthineers.com
Siemens Healthineers
Molecular Imaging
Research PET
810 Innovation Drive Knoxville, TN 37919, USA
John Prior (CHUV, Lausanne, Switzerland)
Axel Rominger (Inselspital, Bern University Hospital, Switzerland)
Joel Karp (University of Pennsylvania, USA)
Thomas Carlier (CHU Nantes, France)
Katia Parodi (Ludwig-Maximilians-Universität München, Germany)
Jorge Cabello, Bernard Bendriem, Mike Casey, Lars Eriksson, Curtis Howe, Carl von Gall, Deepak Bharkhada, Paul Schleyer (Siemens, Knoxville, USA)