## Slide 1: Alignment network error propagation and alignment FCC week 2022, June 1st 2022
Georg Gassner

## Slide 2: Content
Alignment steps
Error Propagation in tunnel alignment networks
Issues with error propagation for tunnel measurements
Examples from SLAC
Alignment examples from FCC-ee optics tuning and alignment mini-workshop
Summary

## Slide 3: Alignment steps
Establish reference network
LCLS-I
LCLS-II
FACET II
LCLS
Fiducialize components / rafts
Align components relative locally to reference network
Smoothing to maintain alignment afterwards

## Slide 4: Error Propagation - Vertical
Level measures height difference normal to gravity
dh15=dh12+dh23+dh34+dh45
n*
Assumption: Random errors, normal distribution.
dh12
dh23
dh34
dh45
H1
H3
H2
H5
H4

## Slide 5: Error Propagation - Horizontal
Simplified (not how tunnels are really measured):
Total station measures angle between previous and next point (plus distance)
X2=X1+dx12=X1-cos(hz13-o10)*d
=
Assumption: Random errors, normal distribution.
X0
X2
X1
X4
X3
hz02
hz13
hz24
d
d
d
d
o10

## Slide 6: Issues with error propagation for tunnel measurements
Assumption was that we only have random errors, a large portion of errors are not random, simulation results are usually optimistic.
None random factors:
Mainly environmental factors like refraction (e.g. the tunnel wall is usually at a different temperature than the air. Especially in ring where inside wall is closer to line of site.)
Instrument errors (calibration at certain conditions, differ from field).
Guide to the expression of uncertainty in Measurements (GUM)

## Slide 7: Simulations for 100km ring – conventional system
Wall monuments
floor monuments
50m
Simulation
Part of errors not random

## Slide 8: SLAC - Reference Network – Linac Laser Alignment System (LLAS)
270 girders of 12m and 3m length
Every girder has one fresnel zone plate target

## Slide 9: Linac Alignment System - Girder
Height / Roll
Adjustment jacks
Horizontal
Adjustment Jack
Telescope mount
Fiducials
Fresnel Lens Target

## Slide 10: LCLSII – Cryomodule installation - Reconfigure LLAS
LCLS-II
Spherical
Lenses

## Slide 11: Conventional Alignment Network S0-S30 – Layout
First kilometer no longer has a light pipe and in addition bypass lines are not tied to girders -> Monument based network
Opposing wall monuments to mitigate effect of refraction
3km
12m
3.3m
Cryomodule
Wall Monument
Floor Monument

## Slide 12: Alignment Network S0-S30 - Measurements
311 Laser Tracker (Leica AT401/402) setups with 5082 measurements
577 Height differences (Leica DNA03)
200 Linac Light Pipe readings
(3km - 6 weeks with 1 crew (2-person)
->100km –> 200 weeks 1 crew + above ground network)
Least Squares Adjustment with GEONET which allows the Linac Light Pipe readings (=±0.5mm) to be included as deviations from a straight line

## Slide 13: Alignment Network S0-S30 – Results (1)
13 µrad
Comparison of coordinate results between adjustment with LLAS and without LLAS.
Location of
radiation chicane
no direct line of sight

## Slide 14: SLAC LCLS – 1km new construction
LCLS

## Slide 15: Network Measurements – Portable Wire Measurement System
In house development
Position 1
Position 2 …
Laser System
Scan of a wire with the laser system
Laser Diode
Laser Beam
Wire
Window (10mm slit)
Photo Cell
Laser system

## Slide 16: LCLS-I Portable Wire Results
Position [m]
Sigma [µm]
Network measurements alone
Wire measurements included
Position [m]
Coordinate change [µm]
Impact of wire on standard deviations
Impact of wire on coordinates
Network of 393 points with the following observation schema: triplets from 69 tracker setups, 465 height differences and 57 offset measurements to two overlapping wires (240m long and 370m long).
A-priori standard deviations:
Laser Tracker: sD=50 µm, sHz=7 µm/m, sV=10 µm/m
Level: sDh=70 µm
Offset measurement: sDo=30 µm

## Slide 17: Alignment tolerances
Fiducialization (4m) (Laser Tracker / Arm)
  Usually within 50µm
Field alignment tolerances (20m) (Laser Tracker):
  Local alignment relative to the network 100µm
Network
  Mid range (200m) 0.3mm (Laser Tracker / Level / Wire)

## Slide 18: Alignment Network – Building stability
The floor in the undulator hall showed very little movement over the last 10 years, the only significant movement was in the FEE and NEH where another floor was installed.

## Slide 19: Alignment Network – Building stability
1mm
Deformation Scale:
Feb. 08 – July 08
July 08 – Feb. 09
Feb. 09 – Jan. 10
Jan. 10 – Sep. 12
Sep. 12 – Oct. 16
Undulator Hall (160 m)
E-Dump
LTU
The walls in the undulator hall are still moving inward, the rate has slowed down.

## Slide 20: FCC-ee alignment mini-workshop
All labs reported that their respective rings / accelerators are constantly moving with especially large motion during new construction and new equipment installation requiring re-alignment.
Several labs apply smoothing techniques to re-align tunnel.
  Conventional methods are very time consuming.
  No other trade should be working in same location.
Number of surveyors is limited, it took SLAC almost a year to hire one. FCC would need 50 (25 teams, data from Mark Jones) with a 4-month duration.
Fiducialization of components / rafts was highlighted and the importance of temperature control and automation of setups is significant.

## Slide 21: Summary
Alignment network:
  Vertical is not an issue (measurements relative to gravity)
  Horizontal (normal to beam) uncertainty grows faster and might need additional measurement system.
Conventional alignment methods will either take a long time or a very large workforce (very repetitive job).
Fiducialization should be automated as much as possible.
Tunnels are not stable, especially right after construction. Re-establishing the alignment network based on tolerances and smoothing of the ring is required.