## Slide 2
Status of the technology development for the FCC-ee vacuum system
C. Garion, CERN/TE/VSC
Content:
  Introduction to the FCC-ee vacuum system in the arcs
  UHV connections with Shape Memory Alloy couplers
  Additive manufacturing with gas dynamic cold spray coating
  Prototyping phase
  Summary

## Slide 3
The FCC-ee Vacuum system has to cope with beam parameters from low-energy (45.6 GeV) high current (1390 mA) version to high-energy (182.5 GeV) low current (5.4 mA) configuration.
Dipole
Dipole
Quad
Shielded bellows
Flange connection
~12 m
FCC-ee vacuum system in the arcs
Layout:
2 rings of ~ 100 km.
Cell length: ~55.9 m.
Beam Energy E(GeV) | Beam Current I(mA) | Photon Flux F’(ph/s/m) | Dynamic Gas Load Q’(mbar·l/s/m)
45.6 | 1390 | 7.17·1017 | 2.90·10-8
80 | 147 | 1.38·1017 | 5.58·10-9
120 | 29 | 4.13·1016 | 1.67·10-9
182.5 | 5.4 | 1.18·1016 | 4.78·10-10
Relevant FCC-ee parameters for the vacuum system design
For the vacuum system, the synchrotron radiation leads to:
High local heat deposition:
  The limit is given by 50 MW/beam synchrotron radiation losses (~650 W/m).
High outgassing:
  Pressure: low 10-9 mbar range.

## Slide 4
Geometry: Tube with two winglets
2 mm thick, 70 mm ID
Material: Copper
  Good thermal conductivity and low electrical resistivity
  Shielding for the X-Ray synchrotron radiation fan and minimizing the irradiation of machine and tunnel components
Surface treatment: NEG coating
  Distributed pumping speed
  Low SEY
  Quick vacuum conditioning
Lumped SR photon absorber:
Distanced by about 5.8 m
Lumped pump: no need for a systematic installation in vicinity of the absorbers 1 or 2 per cell
FCC-ee vacuum chambers
The whole vacuum system shall be designed with a cost-effective and sustainable approach.
119
70
Vacuum chamber prototype cross-section
Illustration of vacuum chambers with absorbers and pumps
Present design as presented in the CDR:

## Slide 5
Copper extrusion
Welding extrusion/flange
Cold sprayed copper for BPM integration
SMA connector
New technologies in development for HEP and of interest for FCC-ee
Technical solutions shall be defined to minimize the cost of the system. The production should be based on series industrial processes and with a minimum of interfaces.
Some technologies are being developed and assessed for the main ring vacuum chambers:
  Shape Memory Alloy UHV connectors for:
    Interconnection.
    BPM pick-ups.
  Gas dynamic cold spray for:
    Additive manufacturing of copper.
    Permanent radiation hard bake out system.
  Weld of extruded chamber/flange.
Beyond the vacuum challenges related to the synchrotron radiation and dynamic vacuum, the vacuum system shall have an affordable cost.

## Slide 6
Concept of SMA connectors for UHV applications
Standard vacuum coupling systems
SMA ring coupler
Shape memory alloy couplers
Heating/mounting
Cooling/dismounting
  Image source: web
Soft/hard materials + mechanical force
Soft/hard materials + temperature control
Martensitic enlarged cold-shape
Austenitic contracted hot-shape
Thin cylindrical gasket

## Slide 7
Concept of SMA connectors for UHV applications
Heating/mounting
Cooling/dismounting
Martensitic enlarged cold-shape
Austenitic contracted hot-shape
Twinned (low stress) or detwinned (high stress) martensite
Heating
Cooling
AS
Af
MS
Mf
As: Austenite start temperature
Af: Austenite finish temperature
Ms: Martensite start temperature
Mf: Martensite finish temperature
Transformation Temperatures (TTs) depend on:
Chemical composition
Internal stress/strain field (dislocation arrays/precipitates)
Thermo-mechanical cycling
Reversible phase transformation: displacive transformation without diffusion process
Martensite: stable at low temperature and high stress
Austenite: stable at high temperature and low stress
Austenite

## Slide 8
Operation principle
1: T<As (40 °C)
2: T>Af (>100 °C)
pmax
Mounting at room temperature
Tightening by heating above 100 °C
Leak Rate < 10-10 mbar·l·s-1 at room temperature
Dismounting by cooling down to -40°C
SMA ring (martensitic)
Tube
Mounting
Clamping
3: T = RT
Operation
4: T≈Mf (-40 °C)
Dismounting
pop
1
2
3
4
As: Austenite start temperature (Ring contraction starts)
Af: Austenite finish temperature (Ring contraction ends)
Ms: Martensite start temperature ( Ring expansion starts)
Mf: Martensite finish temperature (Ring expansion ends)

## Slide 9
Advantages and potential applications
    A compact, leak tight and easily mountable\dismountable connection system, compatible with accelerator environment:
      Magnetic permeability < 1.002
      Thermal outgassing < 10-13 mbar.l-1.s-1.cm-2
      Radiation hard (up to 4 MGy at least)
    Possibility of remote controlling\activation
    Possibility to connect dissimilar materials
    Possibility to use in high demanding areas (e.g. collimator areas, machine/detector interface)
Declamping with a robot
LHC dump windows
LHC collimators
Vacuum module for the MDI area

## Slide 10
f134
f30
SMA-based UHV prototype joints validated
Compatibility with various pipe geometries (DN16, DN25, DN50, DN100) and metals (steel, aluminum, copper, etc.)
Bimetallic joints (St/Ti,St/Alu)
Zero - longitudinal gap connection
Examples of different validated SMA connectors

## Slide 11
145
122
Oval-shaped connectors for FCC-ee chamber interconnections
Rigid oval chamber
SMA connector
Preliminary elliptic profile for SMA connector compatibility based on existing DN100 QCF.
Challenges:
Non Uniform contact pressure
Training of Oval SMA rings
Simulations of contact pressure
Preliminary analytical calculations and FE simulations
Plastic strain in the flange
Replacement of bolted flanges by SMA connectors
Typical expected space in the interconnection

## Slide 12
Training setup for oval-shaped connectors
Next steps:
  Future tests
    Recovery stress test (contact pressure measurement)
    Repeatability of training
  Optimisation of the ring/flange and training
b
a
Design, Simulations and Manufacturing
Trained ring after ovalisation

## Slide 13
Circular connectors for FCC-ee chamber/BPM button
SMA connectors could be advantageously used to integrate the BPM pickup to the vacuum chamber :
More compact than DN16 CF flange.
Easier to assemble.
Transition with copper (no brazing)  cheaper.
Design update of the FCC-ee BPM block on the vacuum chamber, incorporating the proposed equipment (BPM design given CERN/SY/BI for illustration)
Preliminary considerations

## Slide 14
The gas dynamic cold spray coating is based on the projection of solid powder at high velocity.
Compressed gas
Powder feeder
Substrate
De Laval Nozzle
Principle of Cold Spray
A compressed gas is heated before entering in a DeLaval Nozzle.
The gas is accelerated in the nozzle.
Powder is injected in the gas stream and accelerated as well.
The powder reaches velocity up to 1200 m/s.
The powder is plastically deformed and the coating is building up.

## Slide 15
Cold Spray Advantages and Limitations
Advantages:
No powder melting.
  No phase change.
  No grain growth.
  Low heating of the substrate.
No significant impact on the oxide content w.r.t. initial material.
Powder mixture possible.
Compressive residual stress (fatigue life increase).
Nozzle geometry can be tuned for a given jet size.
Thick coating.
High deposition rate.
Limitations:
One constituent has to be ductile.
Accessibility to the surface to be coated.
Possible applications:
Additive manufacturing
Coating (surface or local coating, metallization of polymer)
Metallization of
New materials (composite material)

## Slide 16
Additive manufacturing for:
Manufacturing of components of potentially dissimilar materials
Joining of dissimilar materials
New feature
Repair
Local reinforcement of thin walled structure
New Manufacturing Process
J. Villafuerte, ADVANCED MATERIALS & PROCESSES, 2014
FCC-ee quadrupole chamber with BPM
Impact Innovations website

## Slide 17
[1] Etude de la relation entre porosité et étanchéité à l'ultra-vide de dépôts à base d'aluminium obtenus par projection dynamique par gaz froid ("cold spray"), Sébastien Weiller: http://www.theses.fr/2021UPSLM004
Suitability of cold spray additive manufacturing for UHV applications?
Leak tightness of cold sprayed additive materials:
A previous study on aluminium coating has shown [1]:
Helium leak tight coatings are achievable.
The presence and the role of microporosities formed during the process.
The influence of powder morphology on the porosity formation.
Influence of porosity on leak tightness for different powders and process parameters [1]
Porosity formation in cold sprayed aluminium [1]
Preliminary tests on copper are in preparation:
Copper-to-Copper Cold-Spray samples manufactured
Test bench for thermal outgassing fully refurbished
Tooling for leak tightness test available
Samples for leak tightness and thermal outgassing tests
Test bench for thermal outgassing measurement by accumulation of baked samples

## Slide 18
A technology successfully applied in Accelerators
Local coating:
Electrode:
Measurement of electrons for the FCC-hh beam screen prototype experiment at ANKA.
Clearing electrode (optimisation of ceramic layer required for non-baked system).
Copper and aluminium cold spray coating on ceramic insulated copper coated stainless steel sheet
Heat transfer:
Copper trips on FCC-hh beam screen prototype.

## Slide 19
Copper plate with plasma sprayed ceramic and 0.2 mm thick cold sprayed titanium heating track
Permanent radiation hard heating element for bake out (in or outside vacuum)
Measurement of the temperature field
Possible Applications in Accelerators
Local coating:
Substrate
Ceramic
Ceramic
Cold sprayed resistive track
Insulating layer
Protective layer

## Slide 20
Copper Chamber
Possible Applications in Accelerators
Local coating:
225
231
Expected temperature field along the vacuum chamber
Connectors
Titanium heating track
Expected temperature profile along the vacuum chamber
Permanent radiation hard heating element for bake out (in or outside vacuum)

## Slide 21
Al2O3+13TiO2
Study of the ceramic layer
Samples with different ceramics, ~ 0.5 mm thick, plasma sprayed on copper plates, have been manufactured:
Al2O3 99%
Al2O3 – 13 TiO2
Cr2O3 – 4 SiO2 – 3 TiO2
ZrO2 – 8Y2O3
The different ceramic layer revealed different adhesion quality of the Ti track.

## Slide 22
Significant difference is observed between the different ceramics.
Al2O3-13TiO2 exhibits the lowest residual stress, in the 10 MPa range.
Residual stresses
Residual stresses have been assessed by hole drilling method, on the ceramic and on the track.
Functional tests have been successfully carried out on the Al2O3-13TiO2 plate, including 50 thermal cycles to 250 °C.
Comparative strain measurements
Strain measurements for Al2O3-13TiO2 ceramic layer

## Slide 23
A first prototype of a copper tube with a permanent radiation tolerant bake out system has been produced:
OFE copper tube, 84 mm * 2 mm, 500 mm long
Al2O3-TiO2 ceramic layer
Track in titanium, ~ 110 mm thick, 8 mm width
Distance between the tracks: ~30 mm
Prototype – Proof of concept
Interfaces for the electrical connections
Expected temperature field and profile along the 80 mm vacuum chamber prototype

## Slide 24
Thermal tests:
Low thermal conductance supports
DC current power supply
Electrical connections clamped on the tube
Temperature field
Prototype – Proof of concept
Longitudinal temperature profile
Successful heating to more than 250 °C.
Good temperature homogeneity: +/- 10 °C.

## Slide 25
Thickness irregularities have been observed in the corners of the track (different speed of the gun)  cold spots (thicker) and hot spots (thinner).
Better management of the robot displacement required.
Tests of different track U turn paths.
Measurements of the electrical resistance is not in agreement with expected value  measurement of resistivity of cold sprayed material ongoing.
Electrical connections to be defined for a safe operation.
Prototype – Proof of concept
Thickness irregularities
Hot/cold spots induced by thickness irregularities
Some further developments are required:

## Slide 26
Prototyping phase
Chart: Power density on the FCC-ee absorber
300 mm
170 mm
Water channel:
Typical expected temperature profile
Heat deposition of ~4kW per absorber
700 mm
3.8 W power transmitted in the air (over 700 mm)
Different technologies are being assessed, on small samples, independently and in parallel:
SMA connectors
Copper cold spray additive manufacturing
Radiation hard permanent bake out system
Friction stir welding
In parallel:
Design of lump absorbers: thermal mechanical study, water cooling, assembly to the vacuum chamber.

## Slide 27
Prototyping phase
In parallel:
Alternative designs and manufacturing processes will be explored to take into account, in particular, the design and integration of the synchrotron radiation absorbers.
Discussions about the vacuum system integration, in particular in the magnets, just started. Significant changes, i.e. interbeam distance, length of interconnections with shielded bellows, might be required.
In a second step, the technologies will be applied for the manufacturing of a representative short vacuum chamber prototype:
~ 2 m
Based on copper extrusion and lump absorber
Early 2023

## Slide 28
Conclusion
In a desire to propose cost-effective technical solutions for the FCC-ee vacuum system, the CERN vacuum group has undertaken a series of developments in UHV technologies:
SMA connectors is a mature technology for UHV applications. Implementation study to FCC-ee has been initiated as well as the verification of its suitability to respond to the FCC-ee particularities.
Radiation hard permanent bake out system is under development. Good progress has been done so far: proof of concept test has been successfully carried out.
Additive manufacturing of copper by cold spray is in a first exploratory phase.
Further activities are initiated and will be reinforced in the near future:
Lumped absorber: thermal mechanical design, integration study and prototyping.
Interconnection: Some conceptual designs available. Technical design to be done.
Vacuum chamber prototype manufacturing.
The development of the FCC-ee vacuum system and the related technologies started but requires significant efforts.

## Slide 29
Thanks for your attention.
Thanks to the vacuum team, in particular Christian Duclos, Roberto Kersevan, Hendrik Kos, Marco Morrone, Fabrizio Niccoli and last but not least Samuel Rorison.

## Slide 30
Principle of Cold Spray
1D isentropic flow equations:
De Laval Nozzle
P0, v0, T0
P1, v1, T1,
A1
g=5/3 (1.67) for monoatomic perfect gas
g=7/5 (1.4) for diatomic perfect gas and
g=1,33 for polyatomic perfect gas
Rs is the specific gas constant given by R/Mmolar
A*
Nitrogen is commonly used. Helium is used to reach higher velocity.
Notes: P0 ~30 – 50 bars
T0 : 300 C

## Slide 31
Copper on copper
Assadi et al., Bonding mechanism in cold gas spraying, Acta Materialia, 51, 4379-4394, 2003
T. Schmidt et al., From Particle Acceleration to Impact and Bonding in Cold Spraying, Journal of Thermal Spray Technology, 18, 5-6, 794-808, 2009
Bounding mechanism
Rebound zone
Material jet
Typical surface around the critical velocity
A1050 powder on 2024-T3 substrate
Q. Blochet, Influence of substrate surface roughness on cold-sprayed coating-substrate bond strength in aluminum-based systems, PhD Thesis, Mines ParisTech, 2015
Notes: Adiabatic shear instability

## Slide 32
After Jeandin et al., Coating properties in Modern cold spray, Ed. J. Villafuerte, Springer, 2015
Electrical conductivity
Bounding strength
Coating strength
 Material properties are affected by the cold spray process but they can significantly be recovered by dedicated post treatment.
Some properties: example of copper
T. Stoltenhoff et al., Microstructures and key properties of cold-prayed and thermally sprayed copper coatings, Surface & coatings Technology, 200, 2006
With helium
With nitrogen
C.H. Boyle, Mechanical performance of integrally bounded copper coatings for the long term disposal of used nuclear fuel, Nuclear Engineering and design, 293, 2015
F. Gärtner, Mechanical properties of cold-sprayed and thermally sprayed copper coatings, Surface & Coatings Technology 200, 2006
N2
He

## Slide 33
Application to FCC
Colaminated copper- stainless steel internal screen
Cooling channel
Pumping holes
Copper strip for heat transfer
Requirements:
Copper material (thermal conductivity)
Stainless steel substrate
Discountinuous (longitudinally  reduced Lorentz forces during a magnet quench)
Continuous as close as possible to the cooling channel (better cooling and temperature control)
Done after beam screen assembly (welds)
No spray or coating contamination inside the beam screen
Industrial process
FCC-hh beam screen

## Slide 34
Application to FCC beam screen
Cold sprayed copper on austenitic stainless steel.
Surface preparation by blasting (Al2O3)
Possible improvements:
Surface preparation: laser treatment
Nozzle geometry: correct width
Process parameters