## Slide 1
Dr. A. Faus-Golfe
VHEE with High-Gradient linacs

## Slide 2
OUTLINE
  The Challenges of Radiotherapy
    New RT approaches
      VHEE RT
      New delivery doses modalities: FLASH
  Facilities for VHEE R&D
  Technologies for VHEE: HG linacs
  Summary and Future perspectives

## Slide 3
CHALLENGES IN RADIOTHERAPY
New RT approaches
RT treatment of some radio resistant tumours, paediatric cancers and tumours close to a delicate structure (i.e. spinal cord) is currently limited
  One of the main challenges is to find approaches to increase the normal tissue resistance
  Standard RT is restricted to the few temporal and spatial schemes, dose rates, broad field sizes: mainly photons, 2 Gy/session, 1 session/day, 5 days/week, dose rates ~ 2 Gy/min, field sizes > cm2, homogeneous dose distributions
Possible strategies to spare normal tissue
  Different particle types: Very High Energy Electrons (VHEE)
  Different dose delivery methods: Grid Mini-beam or FLASH RT
Notes: Limitation of conventional radiotherapy. Toxicity to healthy tissue limits the dose.
- How to improve the treatment?
- Preserving healthy tissue
- Induce more effective tumor irradiation
- Particle / energy, ballistic approach

## Slide 4
VHEE THERAPY
State of the Art
With recent High-Gradient linac technology developments, Very High Energy Electrons (VHEE) in the range 100–250 MeV offer the promise to be a cost-effective option in anticancer RT and open up innovative treatment modalities (Grid mini-beam, FLASH,..)
Their ballistic and dosimetric properties can surpass those of photons, which are currently the most commonly used in RT
Their position compared to protons need to be evaluated, but they can be produced at a reduced cost
Notes: CLIC RF X-band cavity prototype (12 Ghz, 100 MV/m) Compact Linear Collider
Current radiotherapy methods:
  6 - 18 MeV photons
  5 - 20 MeV electrons
  50 - 300 MeV/u hadrons (protons, Carbon-12)
Beam simulations with Geant4
Several VHEE beam simulations has been conducted by Manchester group (A. Lagzda)
  Longitudinal and transverse dose profiles of 50 – 250 MeV VHEE beams (with CALIFES parameters)
  250 MeV electron beams can sufficiently penetrate tissue up to 20-25cm – ideal for deep-seated tumour irradiation in non-obese patients

## Slide 5
NEW DELIVERY DOSE MODALITIES
What is FLASH RT?
New way of dose delivery consisting in ultra short irradiation time or “beam on time” (<500 ms) and very high-dose (60-200 Gy/s)
No FLASH effect in vitro with clonogenic cell death as an endpoint (Nias et al. Br J Radiol 42: 553, 1969). The FLASH effect relates to normal, living tissue only.
The beam-on time, not the dose-rate appears to be the major determinant of the FLASH effect
Ultrahigh dose-rate pulses (e.g. 1013 Gy/s) will NOT provide a FLASH effect as long as the dose per pulse is low and the total length of radiation exposure exceeds 100-500 ms
100-500 ms

## Slide 6
VHEE – FLASH Open questions:
What High-Energy ?
Field sizes?
Inhomogenities?
Skin/Bone dose ?
Penumbra?
Neutrons ?
              Scanning small beams ?
              Focusing and Shaping?
              In vivo experiments for validate FLASH RT?
              Pulse structure impact in FLASH RT?
              ......
Mini-beams
FLASH RT
Notes: -Effect of inhomogeneities on beam dose (Manchester team)
-Protons: Shift in Bragg peak position (~diameter of embedded air cavity) – results in large dose deposition in healthy tissue!
-Photons: Dose build-up after the air cavity for photons
-VHEE: relatively insensitive to intervening low-density media (dose difference range ~0.3%) – precise and reliable dose delivery to cancerous tissue
FIGURES:
Dose maps of VHEE, photon and proton beams in a 30×30×30 cm+ water phantom with an air cavity placed at depth , = 10 cm and radius / = 1 cm
Experiments with > 2° converging beams planned at CERN’s CLEAR

## Slide 7
VHEE – FLASH Open questions:
What High-Energy ?
Field sizes?
Inhomogenities??
Skin/Bone dose ?
-Penumbra?
-Neutrons ?
              Scanning small beams ?
              Focusing and Shaping?
              In vivo experiments for validate FLASH RT?
              Pulse structure impact in FLASH RT?
              ......
Mini-beams
FLASH RT
To answer these questions, a lot of experiments are being performed and new facilities are been designed
Notes: -Effect of inhomogeneities on beam dose (Manchester team)
-Protons: Shift in Bragg peak position (~diameter of embedded air cavity) – results in large dose deposition in healthy tissue!
-Photons: Dose build-up after the air cavity for photons
-VHEE: relatively insensitive to intervening low-density media (dose difference range ~0.3%) – precise and reliable dose delivery to cancerous tissue
FIGURES:
Dose maps of VHEE, photon and proton beams in a 30×30×30 cm+ water phantom with an air cavity placed at depth , = 10 cm and radius / = 1 cm
Experiments with > 2° converging beams planned at CERN’s CLEAR

## Slide 8
CLEAR CERN
CLARA STFC CI
PTB Berlin
Operating
Facilities
FLASH IC
ORIATRON CHUV
ELBE HZDR
LOA Salle Noire IPP
FEERIX/ AERIAL
AWA Argonne
Notes: - FLASH R&D platform AERIAL-FEERIX IBA 10 MeV Rhodotron in Strasbourg
- IPP Institute Politechnique de Paris

## Slide 9
CLEAR CERN
CLARA STFC CI
PTB Berlin
Operating
Facilities
FLASH IC
ORIATRON CHUV
ELBE HZDR
LOA Salle Noire IPP
FEERIX/ AERIAL
AWA Argonne
Low-energy e- & FLASH
Cyclos & linacs
High-energy e- & FLASH
NC HG linacs
High-energy e- &FLASH
LPA
High-energy e- & FLASH
SC linacs
Notes: - FLASH R&D platform AERIAL-FEERIX IBA 10 MeV Rhodotron in Strasbourg
- IPP Institute Politechnique de Paris

## Slide 10
PITZ DESY
On design
Facilities
e-beam PHASER
IDRA LOA IPP
DEFT CERN-CHUV
FRIDA Sapienza-INFN
High-energy e-
LPA
High-energy e-
NC HG linacs
High-energy e-
SC linacs

## Slide 11
PITZ DESY
On design
Facilities
e-beam PHASER
IDRA LOA IPP
DEFT CERN-CHUV
FRIDA Sapienza-INFN
High-energy e- & FLASH
LPA
High-energy e- & FLASH
NC HG linacs
High-energy e- & FLASH
SC linacs
W. Wuensch’s talk
Notes: Focus in NC HG linac technology

## Slide 12
Sapienza & INFN on Flash-RT
Development of 5-7 MeV S-Band e-Linac for Flash-RT with industrial partnership (SIT), nowadays installed at IC Orsay
Biperiodic structure ~ 3 GHz
Accelerting mode
Magneting coupling
Dosimetric analysis - 7 MeV – Measurements vs Fluka
7 MeV
D30
PDD
5 October 2021
40th CLIC Project Meeting
FLASH IORT machine
Notes: Sapienza & INFN group have invested on the e-FLASH technology, and has already commercial patnership on it. The group is centered at the SBAI Departement and has a twofold interest on FLASH RT:
- A huge effort has been put on realization of the first FLASH IORT machine with SIT company. The next future target of this team is the development of compact VHEE demonstrator
- The second pillar of this effort has been the developmentof a new TPS software based on MC dose evaluation, able to treat electron, photon and protons and to run on GPU

## Slide 13
Sapienza & INFN on Flash-RT
Development of a C-Band Linac 12 MeV – SIT Company
32 biperiodics cells structure
5 October 2021
40th CLIC Project Meeting
Notes: The second step is the development of a C-band linac

## Slide 14
Sapienza & INFN on Flash-RT
Development of a C-Band Linac 12 MeV – SIT Company
32 biperiodics cells structure
5 October 2021
40th CLIC Project Meeting
Final full prototype expected end 2021

## Slide 15
Monitoring & Dosimetry
Radiobiology
Pre-clinics
KLYSTRON
60-100 MeV LINAC
9 m
Very High-Charge | Medium-Charge
Up to 600 nC per RF pulse | Up to 200 nC per RF pulse
IPAC Conference, 2021
Sapienza & INFN on Flash-RT
Proposal of a VHEE-LINAC- FLASH RT Research Laboratory (Concept)

## Slide 16
Monitoring & Dosimetry
Radiobiology
Pre-clinics
KLYSTRON
60-100 MeV LINAC
9 m
Very High-Charge | Medium-Charge
Up to 600 nC per RF pulse | Up to 200 nC per RF pulse
IPAC Conference, 2021
Sapienza & INFN on Flash-RT
Proposal of a VHEE-LINAC- FLASH RT Research Laboratory (Concept)
FRIDA Project INFN
FLASH Radiotherapy with hIgh Dose-rate particle beAms
A community of 80 researchers distributed in 7 INFN sections & Labs: CT, LNS, Milano, Pisa, Roma1, TIFPA, Torino.
Main Objectives:
    Explore the time scales at which the FLASH effect occurs
    Develop compact, high intensity sources and delivery solutions for EBRT with e and p
    Explore novel detection strategies both for dosimetry and beam monitoring applications
    Explore clinical potential of FLASH EBRT
Budget ~1 ME, Project Approved

## Slide 17
Normal Conducting RF
Super Conducting RF
Laser-Driven e-
Technologies
Short tunable pulse length Laser
110 GHz RF structure
>100 MeV/m is now achievable in labs
12 GHz RF structure
5.712 GHz RF structure
2.998 GHz RF structure
Rhodotrons
Longitudinal shape

## Slide 18
NOVEL ACCELERATOR TECHNOLOGIES FOR VHEE – FLASH Applications
VHEE RT requires beam energies between 50-200 MeV, an improved dose conformity and scale to higher doses rates, in the case of the FLASH-RT until 50 Gy/s are needed.
An international R&D global effort is being made by labs such as CERN, INFN, KEK, MIT, UCLA, LANL, SLAC and industry partners and focused towards the:
Material origin and purity, surface treatments, manufacturing technology
Consistency and reproducibility of test results.
Some promising R&D are:
The distributed coupling accelerator (SLAC)
The use of cryogenic copper
X-band p-mode Distributed Coupling RF structures
Breakdown probability measurement in a single cell cryogenic test
Quality factor for cryogenic copper accelerator structures

## Slide 19
NOVEL ACCELERATOR TECHNOLOGIES FOR VHEE – FLASH Applications II
Another approach for the next generation of compact, efficient and high-performance VHEE accelerator is the use of:
Higher frequencies millimetric waves (~100 GHz)
Higher repetition rates using THz sources
Comparison of X-band and mm-wave linac parameters
Built and tested structures at 110 GHz to 230 MeV (SLAC)
Relativistic electron sources with charges of 160 nC delivered in 1 second, with high repetition rate (ultimately 400Hz) utilizing field emission (100 fC per bunch, 40 ns pulse) (SLAC-NSF)
Commissioning device for relativistic electron source
Relativistic electron source
Laser-triggered semiconductor switch for pulse shaping (MIT)

## Slide 20
Detailed information in: https://indico.cern.ch/event/939012/

## Slide 21
SUMMARY and FUTURE PERSPECTIVES
Challenge in RT is to find possible strategies to spare normal tissue
  Different particle types: Very High Energy Electrons (VHEE)
  Different dose delivery methods: Grid mini-beam or FLASH RT
FLASH consists of ultra short irradiation time or “beam on time” (<500 ms) and very high-dose (60-200 Gy/s), that spare normal tissues from radiation-induced toxicities
  Understanding the FLASH-RT limits (min/max dose/time)
    Minimum beam on-time: ms => 𝜇m=> ns ? Bunching structure?
    Maximun dose > 200 Gy/s ?
  Understanding the FLASH-RT underlying mechanism:
    Biological studies
    Chemical and kinetic: role of the oxygen
  Extending to other species: p, 12C
Dosimetry for ultra-high dose rate is a key stone, development of dosimeters is mandatory for a clinical application (EURAMET - EMPIR 2018)
Which Accelerator is the most suitable to cope the necessary performances?