## Slide 1
Update of the 100cm×50cm RWELL detector
for DHCAL application
Daojin Hong
On behalf of the USTC MPGD Group
State Key Laboratory of Particle Detection and Electronics
University of Science and Technology of China
16th.Nov.2021
RD51 Collaboration Meeting and topical workshop on Wide Dynamic Range operation of MPGDs
15th -19th November 2021

## Slide 2: Options for CEPC PFA HCAL
CEPC HCAL
Iron
Digital
Analog
Scintillator
RPC
MPGD
MPGD：one of the candidates
Resistive WELL detector (RWELL):
Only a drift gap
Only one stage amplification, high gain
Resistive layer-DLC
Requirements of sensitive detector:
Compact
High detection efficiency
Scalable to large size
RWELL: A promising candidate for sensitive detector of DHCAL

## Slide 3: Design of 100cm×50cm RWELL detector
100cm×50cm RWELL
20 sectors
25cm×25cm
100cm×50cm resistive layer
100cm×50cm readout PCB was divided into 8 pieces with 25cm×25cm each
Gluing
Diameter：0.5mm
Pitch: 1mm
Thickness：0.5mm
no rim
Parameters of RWELL foil：
Divide RWELL foil into 20 sectors
Separate resistive layer from readout PCB
Grounding circuit was designed on the resistive layer
Large readout PCB was composed of 8 pieces
Design methods:

## Slide 4: Fabrication of 100cm×50cm RWELL detector
  We make a special PCB and use it as the gluing mask;
  We put glue on both resistive layer PCB and THGEM PCB;
  A vacuum platform was used for gluing;
  8 pieces of readout PCB are used, there are 25 pad(Pad size 5cm×5cm) on each pcb
Painting the glue
Seal the platform
Pumping and drying
Assembling
100cm×50cm RWELL detector:
25cm×25cm

## Slide 5: Performance of 2nd version RWELL detector
Gain vs HV:
Gas: Ar-5%iC4H10
Mean gain:~1500
Daojin Hong，Development of a large-area RWELL detector with DLC coating for CEPC-DHCAL Application，RD51 Collaboration Meeting and Topical Workshop on FE electronics for gas detectors, 14th-18th June 2021
Gain uniformity : RMS/Mean=35.46/256.4～13.8%@~1500 Gain
～25%

## Slide 6: Problem and improvement
Cleaning
Drying
Cu electrode was oxidized after several times cleaning
Cleaning & drying
Cu reduction(to ～5μm) & Micro-etching (～1μm)
Gold deposition: avoid oxidation

## Slide 7: Assembly
～ 40 spacers in the detector
The long metal strip on the readout PCB of 2nd version RWELL detector was replaced by five short metal strips
Acrylic cylinder spacer
improvement

## Slide 8: Performance
Gain vs HV:
8 keV X ray
RWELL
MCA
ORTEC
142 AH
ORTEC 671
PC
(pre-amplifier)
(Spectroscopy amplifier)
Test setup
Gas: Ar-5%iC4H10
～20%

## Slide 9: Gain uniformity & Rate capability
Rate capability：>100 kHz/cm2
Gain uniformity
Collimator dia:5.5mm, 8keV X-ray
（G0curr/G0eff ～ 2.15)
RMS/Mean: ～14.0%@ ～ 5175 Gain

## Slide 10: Detection efficiency to MIPs
Test setup
Detection efficiency: ～95.9%
Gas: Ar-5%iC4H10
signal
Trigger
PMT1
RWELL detector

## Slide 11: MICROROC based electronic system
64 Channels
3 threshold per channel
Dynamic range:1~500 fC
Readout board: integrated with MICROROC
625 channels
Readout pads: 1cm×1 cm
MICROROC Parameters
FEB
DIF
Pad size:1cm×1cm
MICROROC analog part

## Slide 12: Preliminary result
Low gain
S curve
Calibration
High gain
Spectrum measured with external ADC
Cosmic ray test is ongoing

## Slide 13
Gain uniformity of the 100 cm×50 cm RWELL detector detector is about 14%@～5200 gain, and rate capability is large than 100 kHz/cm2. Detection efficiency to MIPs is about 95.9%.
Readout PCB with MICROROC integrated has developed and started the test with RWELL detector.
Advanced test of the MICROROC based FEE with RWELL detector will be studied in the future.
Summary
Outlook
Thanks