Minutes, EP Magnet Working Group Meeting, 23/7/21

Minutes prepared by: Matthias Mentink

Attendees: Francois Garnier, Hans Danielsson, Matthias Mentink, Alexey Dudarev, Benoit Cure, Nicola Pacifico, Xavier Pons, Maciej Ostrega, Sylvain Ravat, Raphael Dumps, Piet Wertelaers, Giovanna Lehmann Miotto, Burkhard Schmidt, Luisa Fabiola Fleig, Stefan Russenschuck.

Associated Indico event link: https://indico.cern.ch/event/1058819/

- Overview of TE-MSC-TM activities (Stefan Russenschuck)

- Stefan presented an overview of TE-MSC-TM activities. In his presentation, Stefan presented the different magnetic field measurement systems available within his group

- He discussed digital twins, where experimental observations on a magnet system is compared to the simulated behaviour

- Stefan presented a geometric over-sampling technique where a smaller probe is used to measure at various positions within a larger bore, where the overall magnetic field quality can be reconstituted through post-processing, thus allowing to characterize a large bore with a rotating field probe without requiring the rotating field probe to match the size of the bore

- Burkhard asked Stefan about on-going and planned measurement projects and Stefan provided an overview of projects.

- Test results with the new 5 T solenoid (Tim Mulder)

- Tim presented the measurement results of a new superconducting 5 T solenoid with a free bore of 200 mm. This Nb-Ti-based solenoid was produced in collaboration with EP-DT colleagues

- In three training steps, the solenoid was trained to 5.5 T in the centre of the bore, after which it was decided to not push the magnet further. This allows operation at 5 T with margin.

- Tim explained that the solenoid was built to allow for general purpose characterization, including superconducting conductors, small coils, hall probes, etc.

- First measurements on the NA60+ demonstrator (Matthias Mentink)

- Matthias gave a short overview of the first measurements on the NA60+ demonstrator.

- Matthias explained that the demonstrator was experimentally characterized in terms of low-current behaviour (up to 20 A), inductance, and magnetic field, without water cooling

- All measurements are thus far consistent with expectations.

- A wide range of joint resistances, up to 144 µΩ, was observed, and it was agreed that the joints should be improved before proceeding to higher currents

- Following a short discussion between Hans, Francois, and Burkhard, a plan was formulated to improve the joints in the subsequent week.

- DT-DI Magnet Activities (Xavier Pons)

- Xavier presented a comprehensive overview of DT-DI magnet-related activities, including maintenance and operation of LHC and non-LHC experimental magnets, as well as magnet R&D projects

- For the LHC experimental magnets, Xavier discussed the present statuses of Alice, LHCb, ATLAS, and CMS.

- Subsequently, Xavier briefly discussed Aegis, Compass, M1, the Vertex magnets, and the Morpurgo dipole.

- For R&D activities, Xavier mentioned a new cryogenic test bench and a diffusion pump test stand in B180, the efforts to produce a cost-effective acquisition system needed for measuring LHCb and CMS strain-gauges, radiation-hard vacuum sensors, a future ATLAS Vacuum Control System.

- Burkhard inquired about the present status of the strain gauge development, and Xavier indicated that this project is ongoing.

- Xavier mentioned the possibility to use the B180 vacuum test bench to investigate radiation hard vacuum sensors and Matthias indicated that he agrees.

- Xavier discussed the possibility of switching over from National Instruments CompactRIO to other technology, given the possible future increase in cost in utilizing this technology. Xavier mentioned that covering the increased future cost of CompactRIO from the M&O budgets could become problematic. Burkhard mentioned that a future increase of licences is commonly observed in many CERN activities, and that it would not necessarily imply a switch-over to different technologies in all cases.