## Slide 1
2021 Commissioning:
RF campaign in EAR1 with C6D6
V. Alcayne, A. Casanovas, F. García, J. Lerendegui,
E. Musacchio and J.A. Pavón

## Slide 2
Goals
  Validate the RF from the simulations from 0.5 eV to 100 keV
  Determine the BIF in the usual capture position for 20/30 mm samples
  Measure the backgrounds in the usual capture position
  All above: comparison Ph4 vs. Ph2/3 if possible
Specific requirements for EAR1
  Lower neutron flux (compared to EAR2) → Longer measurements → Time constraint
  High energy resolution → Large energy range can be covered with same sample before resonances overlap
  Strategy → Minimum number of samples with sufficient time for high quality data
  Big samples
RF @ EAR1: Goals and planning

## Slide 3
RF @ EAR1: Experimental Setup
Standard capture setup of SiMon + 4xL6D6 (D,E,G,L) + Sample Exchanger mounted on 27/28-9
L6D6 D
L6D6 E
L6D6 G
L6D6 L

## Slide 4
RF @ EAR1: Experimental Setup
Standard capture setup Simon + 4xL6D6 + Sample Exchanger mounted on 27/28-9
Detectors-sample distance: 10 cm, 125º from beam direction
PMT at 1500~1600 V

## Slide 5
RF @ EAR1: Experimental Setup
Gafchromic foil inserted behind 65x2 mm natFe sample to check if alignment OK after change of collimator

## Slide 6
EAR1 RESOLUTION FUNCTION CAMPAIGN SUMMARY
RF @ EAR1: Summary
ISOTOPE | DIMENSIONS / MASS | REF. | Protons | Goal
197Au | 20 x 0.1 mm / 0.596 g | 101 | 1.05·1017 | RF @ 5eV- 1 keV, RF vs sample size, BIF
197Au | 30 x 0.1 mm / 1.308 g | 228 | 8.60·1016 | RF @ 5eV- 1 keV, RF vs sample size, BIF
197Au | 45x0.1 mm + 48x0.1 mm / 5.057 g | 22 | 1.76·1017 | RF @ 5eV- 1 keV, RF vs sample size, BIF, Activation
natFe | 65 x 2 mm / 54.55 g | 153 | 5.01·1017 | RF @ keV – 100 keV and beyond
natIr | 20 x 0.1 mm / 0.709 g | 369 | 9.01·1016 | RF < eV, flightpath vs En
natSi | 50.8 x 5 mm / 23.50 g | 629 | 2.97·1017 | RF @ keV – 100 keV and beyond
Empty |  |  | 1.48·1017 | Background
Lead* | 20 x 0.1 mm / 7.281 g | 184 | 1.51·1017 | Background, Ph2/Ph4 comparison
Carbon* | 20 x 6.6 mm / 0.596 g | 18 | 1.26·1017 | Background comparison Ph2/Ph4 comparison
Total RF samples: 1.26·1018 Backgrounds: 4.25·1017
*Peformed during 176Yb(n,γ) campaign
28/9
to
16/10
10/11
to
14/11

## Slide 7
RF @ EAR1: Experimental Setup
Sample Exchanger in time for RF campaign! (Thanks Pino, Sylvain, Michi!!)
20x0.1 mm natIr sample
45x0.1 mm Au sample
65x0.2 mm natFe sample
20x0.1 Au sample

## Slide 8
EAR1 RESOLUTION FUNCTION CAMPAIGN SUMMARY
RF @ EAR1: Summary
28/9
to
16/10
ISOTOPE | DIMENSIONS / MASS | REF. | Protons | Goal
197Au | 20 x 0.1 mm / 0.596 g | 101 | 1.05·1017 | RF @ 5eV- 1 keV, RF vs sample size, BIF
197Au | 30 x 0.1 mm / 1.308 g | 228 | 8.60·1016 | RF @ 5eV- 1 keV, RF vs sample size, BIF
197Au | 45x0.1 mm + 48x0.1 mm / 5.057 g | 22 | 1.76·1017 | RF @ 5eV- 1 keV, RF vs sample size, BIF, Activation
natFe | 65 x 2 mm / 54.55 g | 153 | 5.01·1017 | RF @ keV – 100 keV and beyond
natIr | 20 x 0.1 mm / 0.709 g | 369 | 9.01·1016 | RF < eV, flightpath vs En
natSi | 50.8 x 5 mm / 23.50 g | 629 | 2.97·1017 | RF @ keV – 100 keV and beyond
Empty |  |  | 1.48·1017 | Background
Lead* | 20 x 0.1 mm / 7.281 g | 184 | 1.51·1017 | Background, Ph2/Ph4 comparison
Carbon* | 20 x 6.6 mm / 0.596 g | 18 | 1.26·1017 | Background comparison Ph2/Ph4 comparison
Total RF samples: 1.26·1018 Backgrounds: 4.25·1017
10/11
to
14/11
*Peformed during 176Yb(n,γ) campaign

## Slide 9
RF @ EAR1: Summary
BIF and RF up to 1 keV
RF at 1 – 100 keV range and beyond
RF and flight path < 1eV
EAR1 RESOLUTION FUNCTION CAMPAIGN SUMMARY

## Slide 10
RF: First look at 197Au, 2021 vs 2018
Au-197, 20x0.1 mm (2021) vs Au-197, 20x0.125 mm (2018)
Calib. in ToF to En: L0=184.28 m (60 eV Au res.)

## Slide 11
RF: First look at 197Au, 2021 vs 2018
197Au, 20x0.1 mm (2021) vs 197Au, 20x0.125 mm (2018)
Calib. in ToF to En: L0=184.28 m (60 eV Au res.)

## Slide 12
RF @ EAR1: natFe
natFe, 65x0.2 mm (2021), 92% Fe-56, Fe-54 (5.85%), Fe-57 (2.1%)
Calib. in ToF to En: L0=184.28 m
5x1017 protons

## Slide 13
RF @ EAR1: natFe
natFe, 65x0.2 mm (2021), 92% Fe-56, Fe-54 (5.85%), Fe-57 (2.1%)
Calib. in ToF to En: L0=184.28 m
5x1017 protons

## Slide 14
RF @ EAR1: natFe
natFe, 65x0.2 mm (2021), 92% Fe-56, Fe-54 (5.85%), Fe-57 (2.1%)
Calib. in ToF to En: L0=184.28 m
5x1017 protons

## Slide 15
RF @ EAR1: natFe
natFe, 65x0.2 mm (2021), 92% Fe-56, Fe-54 (5.85%), Fe-57 (2.1%)
Calib. in ToF to En: L0=184.28 m
5x1017 protons

## Slide 16
RF @ EAR1: natSi
natSi (Si wafer), 50x0.5 mm (2021)
Calib. in ToF to En: L0=184.28 m
2.87x1017 protons

## Slide 17
RF @ EAR1: natSi
natSi (“Si “wafer”), 50x0.5 mm (2021)
Calib. in ToF to En: L0=184.28
2.87x1017 protons

## Slide 18
Beam interception factor: Phase4
20 mm and 30 mm BIF diameter calculated with gold samples from full intercepting 45mm + 48 mm samples: values inferior,although similar, to Phase3
(differences possibly to no WF applied -> No correction due to changes in det. eff.)
 | Ph4 (exp.) | Ph2 (simulations)
20 mm | 0.6312 | 0.68
20mm (176Yb) | 0.6344 | 0.68
30 mm | 0.9168 | 0.96
Similar results in the posterior 176Yb(n,γ) campaign
Pending: verification with simulations/TC once WF applied

## Slide 19
RF @ EAR1: Backgrounds
natPb and natC samples:
  natPb same sample as 2018; natC 6.6 mm thickness (2021) vs 5 mm (2018)

## Slide 20
RF @ EAR1: Backgrounds
natPb and natC samples:
  Beam-off subtracted; natC scaled by thickness ratio (6.6 mm vs 5 mm)
In-beam gamma rays reduction?
Neutron scattering scaling
with increasing neutron flux

## Slide 21
RF @ EAR1: Backgrounds
natPb and natC samples:
  Beam-off subtracted; natC scaled by thickness ratio
In-beam gamma rays reduction?
Neutron scattering scaling
with increasing neutron flux (as expected)

## Slide 22
γ-Flash and prepulses
Au-197, 20x0.1 mm (2021)
Au-197, 20x0.125 mm (2018)

## Slide 23
γ-Flash and prepulses
Au-197, 20x0.1 mm (2021)
Au-197, 20x0.125 mm (2018)

## Slide 24
γ-Flash and prepulses
Au-197, 20x0.1 mm (2021)
Au-197, 20x0.125 mm (2018)

## Slide 25
γ-Flash and prepulses
Au-197, 20x0.1 mm (2021)
Au-197, 20x0.125 mm (2018)

## Slide 26
Preliminary: Tγ-Flash and prepulses
TFlash distribution time distribution in C6D6, 2018 vs 2021
Comparison dedicated vs parastic to see impact, 176Yb(n,γ) campaign
<1% of pulses
6.8% of pulses
Possible improvement in γ-Flash identification by PSA routine?
50 ns
50 ns

## Slide 27
Preliminary: Tγ-Flash and prepulses
TFlash distribution time distribution in C6D6, 2018 vs 2021
Comparison dedicated vs parastic to see impact, 176Yb(n,γ) campaign
6.8% of pulses
natFe resonance at 80 keV: width ~300 ns
Effect on keV range resonances? To be studied…
Possible improvement in γ-Flash identification by PSA routine?
50 ns

## Slide 28
Summary and first conclusions
Measurements with big, massive samples allowed to gather useful capture resonance data up to 100 keV and beyond for EAR1 RF verification in the full capture energy range
  achieved in the limited amount of time availabe
  Natural samples, more complex analysis
Measurement of several 197Au samples for verification of beam interception factor, and activation measurement of the flux absolute value at 4.9 eV sat. res.
Preliminary background comparison (lead and carbon): more flux but less (or similar) in-beam gamma rays
Next steps:
  Study of effect of prepulses in TFlash identification and the possible effects in En resolution: PSA, Dedicated vs Parasitic
  Usual capture analysis (fine data sorting, WF, simulations, etc.) to extract yield / w. CR
  First comparison: Ph2 RF
  RF verification, and comparison with Ph3 RF, expected to be completed in the first half of 2022

## Slide 29
Thank you for your attention!