## Slide 1: Round Robin Test for Ic measurements of Nb-Ti and YBCO conductors (EUCARD WP7 task 4)WP7
Jerome Fleiter (CERN, University of Grenoble)
CERN, EUCARD WP7-Task4

## Slide 2
Laboratories involved
  INFN-LASA (Milano)
  UNIGE (Geneva)
  CEA (Saclay)
  CNRS (Grenoble)
  CERN (Geneva)
2 Samples measured at 4.2K
  Nb-Ti LHC strand mounted on a LHC barrel at CERN.
  YBCO tape from SuperPower, mounted on a barrel at INFN- LASA.

## Slide 3: Nb-Ti @ 4.2K
  No correction of field
  No correction of temperature
Chart: Ic using 0.1μV/cm criterion (Sample 02R00056A01UX.129)
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
UNIGE 4.19K 0.1μV/cm | 974.7 | 777 | 645.5 | 527.6 | 419.7 | 313.2 | 207.9 | 109.7
INFN LASA 4.22K 0.1μV/cm | 100.2695 | 99.2241 | 197.7154 | 304.2103 | 411.8705 | 520.6646 | 637.52 | 772.767 | 968.9739 | 1342.726 | 1801.561
CERN 4.26K 0.1μV/cm | 303.3 | 404 | 515 | 627 | 762 | 959
32 mm

## Slide 4: Nb-Ti @ 4.2K
Chart: B (T)
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
UNIGE | 47.9 | 42.9 | 43.7 | 43.8 | 44.3 | 34.9 | 27.2 | 18.5
INFN | 17.7716 | 17.4065 | 25.2767 | 33.3449 | 39.8716 | 41.1719 | 47.0066 | 44.3084 | 42.544 | 44.4905 | 72.4408
CERN | 34 | 41 | 44 | 44 | 45 | 45
  Without temperature correction, measurements are already within 4%.
Chart: Ic using 1μV/cm criterion (Sample 02R00056A01UX.129)
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
UNIGE 4.19K 1μV/cm | 1016.3 | 820.1 | 677.7 | 556.8 | 444.5 | 334.5 | 226.3 | 124.3
INFN LASA 4.2K 1μV/cm | 114.1402 | 113.2575 | 216.5722 | 325.9594 | 436.3563 | 550.613 | 669.526 | 813.9874 | 1022.862 | 1414.048 | 1859.745
CERN 4.26K 1μV/cm | 323 | 432 | 542 | 660 | 802 | 1009

## Slide 5: YBCO @ 4.2K
  No correction of field
  Engineering Critical current density at 19T is 2400 A/mm2 (B//ab plane)
Chart: Ic data using 1μV/cm criterion compared to a fit (Spool SP-20100108-1 sample01)
| 6.2169229999999996E-4 | 1.000572 | 2.001385 | 4.0021129999999996 | 6.0026169999999999 | 8.0029699999999995 | 10.001899999999999 | 12.001670000000001 | 13.001099999999999
LASA 4.25K 1μV/cm | 1436.672 | 1398.596 | 1342.943 | 1268.928 | 1219.365 | 1177.864 | 1138.252 | 1101.266 | 1083.698
UNIGE 4.19K 1μV/cm | 995.8 | 988.5 | 980.5 | 971.1 | 959.6
Fit | 1512.1318 | 1432.6635 | 1399.5951 | 1376.3189 | 1357.4462 | 1341.0936 | 1326.3797 | 1312.8201 | 1300.12 | 1288.0865 | 1276.5865 | 1265.5243 | 1254.8285 | 1244.4447 | 1234.3303 | 1224.4516 | 1214.7809 | 1205.2959 | 1195.9778 | 1186.8107 | 1177.7815 | 1168.8786 | 1160.0922 | 1151.4136 | 1142.8355 | 1134.3511 | 1125.9548 | 1117.6412 | 1109.4058 | 1101.2444 | 1093.1532 | 1085.129 | 1077.1687 | 1069.2694 | 1061.4286 | 1053.6441 | 1045.9137 | 1038.2354 | 1030.6075 | 1023.0283 | 1015.4962 | 1008.0098 | 1000.5679 | 993.1691 | 985.8124 | 978.4966 | 971.2208 | 963.984 | 956.7854 | 944.2772 | 909.135 | 874.8172 | 841.2646 | 808.4317 | 776.283 | 744.7905

## Slide 6: YBCO @ 4.2K
Chart: Spool SP-20100108-1 sample01 (Ic @77K,s.f. =102A)
| 17.5 | 18 | 18.5 | 19
UNIGE n-value | 51.2 | 51.5 | 50.7 | 51.4
LASA n-value | 79.2599 | 77.9036 | 67.5312 | 67.8598 | 62.6497 | 57.3581 | 57.8302 | 57.7163 | 57.1208
  n-values are high. Sharp transition takes place.

## Slide 7: Large spread of Ic within same spool (SP-20100108-1) for YBCO tape
Chart: Ic performance of different samples from same spool @ 4.2K
| 0 | 0.5 | 1 | 1.5 | 2 | 3 | 4 | 6 | 8 | 10 | 12
CERN | 1301.1 | 1297 | 1269 | 1238.7 | 1212.7 | 1171 | 1134.1 | 1084.5 | 1041.3 | 1004.2 | 969.7
LASA | 1436.672 | 1398.596 | 1342.943 | 1268.928 | 1219.365 | 1177.864 | 1138.252 | 1101.266 | 1083.698
YBCO @ 4.2K

## Slide 8: Conclusions
Critical current Ic measurements of Nb-Ti strand and YBCO tape performed at 3 different laboratories are in good agreements.