## Slide 1: Progress on coupling MPGD-based PDs with nanodiamond photocathodes
Fulvio Tessarotto
  On behalf of the nanodiamond-THGEM group
  (Trieste-Bari-CERN)
Motivation for the R&D
H-ND: production and spray coating
Photoemission measurements
THGEMs with nanodiamond coating
Measurements in different gas mixtures
Aging studies
Conclusions

## Slide 2: THGEM+CsI: 8 years of dedicated R&D
Effective gain = 0.91 · 106
Ar/CH4: 50/50
PARAMETERS:
Diam. = 0.4 mm
Pitch = 0.8 mm
Thickn. = 0.4 mm
Rim = 10 μm
a. u.
Photon yield & Charged Particles vs Drift Field
Difficult mastering of the technology to cover very large area: el. stability and reproducibility issues

## Slide 3: COMPASS Hybrid THGEM+MM PDs
successfully
implemented in 2016 on
COMPASS RICH-1
active area = 1.4 m2
eff. gain ~ 15000,
gain stability 5% (p/T corr.)
single γ ang. res. 1.8 mrad
Hybrid PD scheme
N(ϴCh) = p0sin2ϴCh+ p1ϴCh
Standard Bulk Micromegas (CERN)
at ß=1 Nph = 12.9
S = 11.5, B = 1.4
THGEMs
S. Dalla Torre, NIM A 970 (2020) 163768
IBF = 3%
staggered
THGEMs
CsI

## Slide 4
Modular Hybrid THGEMs + Micromegas Minipad prototype
JINST 15 (2020) C09052
AIDA-2020-NOTE-2020-006
“after the positive experience with COMPASS RICH”
Minipad Hybrid THGEM+MM PDs
Prototype with 10x10 cm2 active area.
1024 square pads of 3x3 mm2 with 0.5 mm inter-pad space

## Slide 5: Nanodiamonds: the beginning in Bari
Antonio Valentini studies on carbon nanotubes led to identify diamond nano-grains as the most promising robust and efficient photoconverter

## Slide 6: Rich-graphite Nano-Diamond film
Photocatodes: diamond film obtained with Spray Technique
Spray technique: T ~ 120° (instead of ~800° as in standard techniques)
CsI bandgap: 6.2 eV; electron affinity: 0.1 eV; hygroscopic; ages by ion bombardment (~mC/cm2)
Diamond bandgap: 5.5 eV; chemically inert and robust; if hydrogenated: electron affinity -1.27 eV
Hydrogenated chemical vapor deposited diamond films (4-6 μm) known to have QE ~15% @ 140 nm.
47 %
L.Velardi, A.Valentini, G.Cicala, Diamond & Related Materials 76 (2017) 1
Heterogranular-structured diamond-gold nanohybrids proposed as stable field emission cathode material
Nano-Diamond grains (size: ~250 nm), with variable sp2 (graphite phase) and sp3 (diamond phase) hybridized carbon contents treated in H2 microwave plasma show large QE: >40% @ 140 nm

## Slide 7: Nano-Diamond Q.E.
PCD_MgF2
Photocathode area  6 mm
Spectral response 110 220 nm
Phototube
Hamamatsu
Hydrogenated powder
Hydrogenated powder
Phototube Hamamatsu
As received ND powder
L. Velardi, et al
Diamond & Related Materials 76 (2017) 1–8

## Slide 8: Hydrogenated Nano-Diamond

## Slide 9: Hydrogenation: MWPECVD setup in Bari
Hydrogenation details:
Vacuum: ~6.5x10-6 mbar.
70 mm between H2 source and ND powder.
H2 gas generated by electrolysis from distilled water. H2 gas flow rate controlled to 200 sccm.
Hydrogenation of ND powder: 1 hour at 43 mbar.
1380 W microwave power  1000 0C temperature with 830 0C substrate holder temperature.
1250 W microwave power  810 0C temperature with 650 0C substrate holder temperature.

## Slide 10: Different ND and H-ND
As-Rec_ND_BDD 0.5 Micron
Hydrogenated_ND_D&T
Hydrogenated_ND_E6

## Slide 11: Pulsed spray coating in Bari
Spray Control Interface
Low Temperature Deposition ( 120 °C)
Good reproducibility technique
Scalable to cover large areas

## Slide 12: Basic questions about H-ND and TGEM coupling
Is the H-ND layer on the THGEM:
- Reducing the electrical stability?
Changing the gain response?
Providing the same PDE as on PCBs?
Uniform and stable?
- More robust with respect to CsI?

## Slide 13: “Thickness” study
Sufficient surface coverage is reached with “100 shots” thickness

## Slide 14: Q.E. measurement setup in Bari
Turbo Pumps
Deuterium Lamp
Monochromator
Collimator
Lock
Gate
NIST PD
or
Photocathode
Trap
RotativePump
Multiwire Proportional Chamber(MWPC)
Anodic wires
Insulating support
Photocathode
Support clips
NIST Photodiode

## Slide 15: Q.E. measurement setup in Bari

## Slide 16: Photoemission of different NDs and H-NDs

## Slide 17: Q.E. measurement setup with gas
ATEX compliant gas mixing unit

## Slide 18: Photocurrent vs λ in Ar/CH4 gases
Chart: H-ND_D&T_1000 0C_PCB4
| 1160 | 1180 | 1200 | 1220 | 1240 | 1260 | 1280 | 1300 | 1320 | 1340 | 1360 | 1380 | 1400 | 1420 | 1440 | 1460 | 1480 | 1500 | 1520 | 1540 | 1560 | 1580 | 1600 | 1620 | 1640 | 1660 | 1680 | 1700 | 1720 | 1740 | 1760 | 1780 | 1800 | 1820 | 1840 | 1860 | 1880 | 1900 | 1920 | 1940 | 1960 | 1980 | 2000 | 2020 | 2040 | 2060 | 2080 | 2100
Vacuum | 4.638e-13 | 5.8378e-13 | 8.2316e-13 | 1.1445e-12 | 1.623e-12 | 2.0903e-12 | 2.429e-12 | 2.803e-12 | 3.5185e-12 | 4.6385e-12 | 6.0854e-12 | 8.4364e-12 | 1.1374e-11 | 1.4439e-11 | 1.7505e-11 | 2.0869e-11 | 2.4092e-11 | 2.6235e-11 | 2.7147e-11 | 2.72e-11 | 2.827e-11 | 3.2528e-11 | 3.754e-11 | 3.735e-11 | 2.8487e-11 | 1.6108e-11 | 7.5349e-12 | 3.9697e-12 | 2.8833e-12 | 2.5342e-12 | 2.3811e-12 | 2.2834e-12 | 2.2075e-12 | 2.1438e-12 | 2.0691e-12 | 1.9783e-12 | 1.8648e-12 | 1.7446e-12 | 1.6172e-12 | 1.4992e-12 | 1.386e-12 | 1.2855e-12 | 1.1974e-12 | 1.1169e-12 | 1.0564e-12 | 9.9953e-13 | 9.492e-13 | 9.0114e-13
Ar_CH4_00_100 | 1.7688e-13 | 1.8793e-13 | 4.0699e-13 | 1.6929e-12 | 4.4574e-12 | 8.0058e-12 | 1.0538e-11 | 1.1604e-11 | 1.1817e-11 | 1.241e-11 | 1.4423e-11 | 1.6784e-11 | 1.6768e-11 | 1.2822e-11 | 7.2909e-12 | 3.4487e-12 | 1.8513e-12 | 1.3723e-12 | 1.235e-12 | 1.1883e-12 | 1.1633e-12 | 1.1445e-12 | 1.1276e-12 | 1.1056e-12 | 1.0744e-12 | 1.0249e-12 | 9.6746e-13 | 9.0457e-13 | 8.462e-13 | 7.9156e-13 | 7.3854e-13 | 6.9042e-13 | 6.5048e-13 | 6.1122e-13 | 5.8648e-13 | 5.6014e-13
Ar_CH4_25_75 | 1.6839e-13 | 1.8496e-13 | 4.8033e-13 | 1.9349e-12 | 4.8434e-12 | 8.346e-12 | 1.0739e-11 | 1.1709e-11 | 1.1904e-11 | 1.2498e-11 | 1.453e-11 | 1.6901e-11 | 1.6901e-11 | 1.294e-11 | 7.3494e-12 | 3.4778e-12 | 1.8644e-12 | 1.3808e-12 | 1.2495e-12 | 1.1935e-12 | 1.1696e-12 | 1.1571e-12 | 1.1448e-12 | 1.1237e-12 | 1.0886e-12 | 1.041e-12 | 9.8547e-13 | 9.2563e-13 | 8.6468e-13 | 8.0707e-13 | 7.5302e-13 | 7.071e-13 | 6.6284e-13 | 6.276e-13 | 5.9727e-13 | 5.6988e-13 | 5.4201e-13
Ar_CH4_50_50 | 1.7887e-13 | 2.0708e-13 | 6.0515e-13 | 2.2531e-12 | 5.2234e-12 | 8.5065e-12 | 1.0627e-11 | 1.1485e-11 | 1.1659e-11 | 1.2235e-11 | 1.4225e-11 | 1.6546e-11 | 1.6527e-11 | 1.2654e-11 | 7.1942e-12 | 3.4074e-12 | 1.8358e-12 | 1.3681e-12 | 1.2352e-12 | 1.1848e-12 | 1.1667e-12 | 1.1561e-12 | 1.1464e-12 | 1.1267e-12 | 1.0961e-12 | 1.0489e-12 | 9.9318e-13 | 9.3179e-13 | 8.7255e-13 | 8.1318e-13 | 7.629e-13 | 7.1535e-13 | 6.7419e-13 | 6.4237e-13 | 6.1429e-13 | 5.8838e-13 | 5.6253e-13
Ar_CH4_75_25 | 1.9166e-13 | 2.8122e-13 | 9.2335e-13 | 2.7078e-12 | 5.5049e-12 | 8.2094e-12 | 9.8177e-12 | 1.0431e-11 | 1.056e-11 | 1.109e-11 | 1.289e-11 | 1.4994e-11 | 1.5006e-11 | 1.1506e-11 | 6.5632e-12 | 3.1282e-12 | 1.6948e-12 | 1.272e-12 | 1.1517e-12 | 1.1137e-12 | 1.0978e-12 | 1.0879e-12 | 1.082e-12 | 1.0676e-12 | 1.0421e-12 | 1.0053e-12 | 9.5375e-13 | 8.9663e-13 | 8.4495e-13 | 7.996e-13 | 7.503e-13 | 7.0798e-13 | 6.673e-13 | 6.3706e-13 | 6.1021e-13 | 5.8288e-13 | 5.6423e-13
Ar_CH4_80_20 | 1.6663e-13 | 2.6674e-13 | 9.6261e-13 | 2.8427e-12 | 5.5803e-12 | 8.0947e-12 | 9.5586e-12 | 1.0117e-11 | 1.0237e-11 | 1.0755e-11 | 1.252e-11 | 1.4577e-11 | 1.4589e-11 | 1.1191e-11 | 6.3795e-12 | 3.0337e-12 | 1.6386e-12 | 1.2253e-12 | 1.1083e-12 | 1.0739e-12 | 1.0606e-12 | 1.0519e-12 | 1.0437e-12 | 1.0304e-12 | 1.0062e-12 | 9.6932e-13 | 9.2778e-13 | 8.7476e-13 | 8.2356e-13 | 7.7126e-13 | 7.2343e-13 | 6.8273e-13 | 6.5165e-13 | 6.2029e-13 | 5.9795e-13 | 5.7609e-13 | 5.5326e-13
Ar_CH4_85_15 | 2.1489e-13 | 2.1503e-13 | 2.1191e-13 | 2.1023e-13 | 2.1494e-13 | 2.1768e-13 | 2.1852e-13 | 2.2181e-13 | 2.2718e-13 | 2.2885e-13 | 2.3247e-13 | 2.4327e-13 | 3.9918e-13 | 1.1735e-12 | 2.9616e-12 | 5.4506e-12 | 7.6474e-12 | 8.9075e-12 | 9.3961e-12 | 9.4973e-12 | 9.9623e-12 | 1.1566e-11 | 1.3442e-11 | 1.3437e-11 | 1.0306e-11 | 5.9053e-12 | 2.8489e-12 | 1.6037e-12 | 1.2528e-12 | 1.1504e-12 | 1.1128e-12 | 1.092e-12 | 1.0797e-12 | 1.0702e-12 | 1.0557e-12 | 1.0218e-12 | 9.8201e-13 | 9.3537e-13 | 8.8295e-13 | 8.3243e-13 | 7.8189e-13 | 7.3854e-13 | 6.9846e-13 | 6.6666e-13 | 6.4158e-13 | 6.1046e-13 | 5.8973e-13 | 5.7062e-13
_Ar_CH4_90_10 | 1.8336e-13 | 3.8003e-13 | 1.2473e-12 | 3.0564e-12 | 5.2761e-12 | 7.1259e-12 | 8.165e-12 | 8.5786e-12 | 8.6705e-12 | 9.1072e-12 | 1.0598e-11 | 1.2337e-11 | 1.2341e-11 | 9.4695e-12 | 5.4102e-12 | 2.5894e-12 | 1.4168e-12 | 1.0642e-12 | 9.753e-13 | 9.4428e-13 | 9.3236e-13 | 9.2814e-13 | 9.2471e-13 | 9.1869e-13 | 8.9544e-13 | 8.6653e-13 | 8.2648e-13 | 7.8072e-13 | 7.3664e-13 | 6.9797e-13 | 6.5827e-13 | 6.1979e-13 | 5.9269e-13 | 5.6522e-13 | 5.411e-13 | 5.2198e-13 | 5.0134e-13
Ar_CH4_95_05 | 1.6085e-13 | 1.6065e-13 | 1.667e-13 | 1.6847e-13 | 1.6777e-13 | 1.6953e-13 | 1.653e-13 | 1.7256e-13 | 1.7537e-13 | 1.7681e-13 | 1.8244e-13 | 2.3907e-13 | 5.7402e-13 | 1.5051e-12 | 3.0437e-12 | 4.715e-12 | 6.0323e-12 | 6.7722e-12 | 7.0761e-12 | 7.1452e-12 | 7.5057e-12 | 8.7193e-12 | 1.0152e-11 | 1.0163e-11 | 7.8072e-12 | 4.4779e-12 | 2.1667e-12 | 1.1975e-12 | 9.1589e-13 | 8.3931e-13 | 8.1827e-13 | 8.1118e-13 | 8.0868e-13 | 8.0949e-13 | 8.0559e-13 | 7.8579e-13 | 7.6457e-13 | 7.3384e-13 | 6.9821e-13 | 6.5441e-13 | 6.2361e-13 | 5.8699e-13 | 5.5964e-13 | 5.3849e-13 | 5.1481e-13 | 4.9565e-13 | 4.8049e-13
Ar_CH4_97_03 | 2.4509e-13 | 6.3015e-13 | 1.5576e-12 | 2.9242e-12 | 4.247e-12 | 5.2609e-12 | 5.8442e-12 | 6.1057e-12 | 6.1627e-12 | 6.4787e-12 | 7.538e-12 | 8.7743e-12 | 8.784e-12 | 6.7494e-12 | 3.8726e-12 | 1.8659e-12 | 1.0352e-12 | 7.8613e-13 | 7.2474e-13 | 7.0304e-13 | 6.9738e-13 | 6.9837e-13 | 6.9499e-13 | 6.9208e-13 | 6.8044e-13 | 6.6125e-13 | 6.3224e-13 | 6.0025e-13 | 5.6978e-13 | 5.4146e-13 | 5.0973e-13 | 4.8895e-13 | 4.674e-13 | 4.4653e-13 | 4.3386e-13 | 4.2021e-13 | 4.0709e-13
Ar_CH4_98_02 | 2.7943e-13 | 6.9707e-13 | 1.5671e-12 | 2.7765e-12 | 3.8797e-12 | 4.7284e-12 | 5.2259e-12 | 5.4445e-12 | 5.4914e-12 | 5.7674e-12 | 6.7139e-12 | 7.822e-12 | 7.8384e-12 | 6.025e-12 | 3.4561e-12 | 1.6668e-12 | 9.211e-13 | 7.0116e-13 | 6.4365e-13 | 6.2623e-13 | 6.216e-13 | 6.2351e-13 | 6.2052e-13 | 6.1714e-13 | 6.0602e-13 | 5.8805e-13 | 5.639e-13 | 5.3832e-13 | 5.0882e-13 | 4.8087e-13 | 4.5893e-13 | 4.3639e-13 | 4.168e-13 | 3.9939e-13 | 3.8546e-13 | 3.7409e-13 | 3.6197e-13
Ar_CH4_99_01 | 3.8752e-13 | 8.4782e-13 | 1.5942e-12 | 2.4667e-12 | 3.232e-12 | 3.8444e-12 | 4.2243e-12 | 4.3957e-12 | 4.4379e-12 | 4.6604e-12 | 5.4254e-12 | 6.3193e-12 | 6.3322e-12 | 4.8695e-12 | 2.8035e-12 | 1.3595e-12 | 7.6252e-13 | 5.8303e-13 | 5.3632e-13 | 5.2219e-13 | 5.1715e-13 | 5.1977e-13 | 5.1853e-13 | 5.1626e-13 | 5.0849e-13 | 4.9611e-13 | 4.7371e-13 | 4.5309e-13 | 4.3061e-13 | 4.1027e-13 | 3.9078e-13 | 3.7116e-13 | 3.5764e-13 | 3.4239e-13 | 3.3211e-13 | 3.2242e-13 | 3.1296e-13

## Slide 19: λ = 162 nm H-ND D&T
Photocurrent vs electric field
Substrate holder for
photocurrent measurement
Gap between substrate and electric wire: 4.4 mm.
Scan performed with MgF2 window in vacuum and various Ar:CH4 gas mixtures

## Slide 20: Photocurrent vs E in Ar/CH4 gases

## Slide 21: ASSET at CERN RD51 Lab
Aging tests performed thanks to the possibility to irradiate the sample inside the setup

## Slide 22: First aging measurements
Q.E.
Chart: Aged Q.E./Original Q.E. [H-ND, 50 shots, 160 nm]
| 0 | 1 | 11 | 21 | 31
QE at 160 nm | 1 | 0.95842 | 0.55716 | 0.25187 | 0.25314
preliminary indication:
H-ND is at least ten times more robust than CsI
against irradiation and ion bambardment
300 nm CsI
H-ND D&T

## Slide 23: THGEMs + H-ND
First trial of THGEM + H-ND not successful: coated THGEMs lost electrical strength.
Second trial: both full and half-coated THGEMs

## Slide 24: THGEMs H-ND
First results were puzzling:
J. Agarwala, et al.
Nuclear Inst. and Methods in Physics Research, A 952 (2020) 161967
Systematic studies: 15 THGEM samples characterized

## Slide 25: THGEM characterization
Full characterization including charging up measurements before and after H-ND coating
Voltage Applied:
Drift = 2250 V
T_top = 1750 V
T_bot = 500 V
Anode = 0 V
Gas Mixture:
Ar = 70 %
CO2 = 30 %
@ 10 l/h flow rate
X-ray Source:
Fe-55
THGEM active area:
30 mm x 30 mm
Electronics used:
CAEN N1471H HV PS
CREMAT CR-110 Pre Amp
ORTEC 672 Amp
CREMAT CR-150 Eva Board
AMPTEK MCA 8000A

## Slide 26: THGEMs are H-ND compatible
The response of THGEMs as electron multipliers is unaffected by H-ND coating
heat treatment after coating
and p/T corrections + charging-up are essential

## Slide 27: CONCLUSIONS
Exploratory investigation on H-ND photocathodes
  Promising values of Q.E. in the far UV but no clear reproducibility.
  High robustness against moisture, light irradiation, ion bombardment
Perspective of coupling H-ND with THGEM-based PDs
  Full compatibility (same electron multiplication response if correctly coated)
  Systematic study for gas, HV config. and detector geometry started
Potentially interesting for windowless gaseous PDs, Picosec, fire detection, …