## Slide 1: CRAB cavity tuning system SPS and LHC
16/12/2021
K. Artoos

## Slide 2: Tuning principle, Range
FINE TUNING PRINCIPLE
Symmetric actuation through tuner frame and concentric tubes. Actuator outside cryostat and floating
RFD
512 kHz/mm*
E. Cano Pleite
Tuning range
± 1.22 MHz
± 2.38 mm*
elastic range at 2 K
Max. force:
± 6.7 kN
2.8 kN/mm*
*Measured as tuner stroke or Δ distance between 2 plates
DQW
318 kHz/mm*
S. Verdú Andrés
Tuning range
± 509 kHz
±1.6 mm*
elastic range at 2 K
Max. force:
± 3.8 kN
2.2-2.4 kN/mm*
Range in Spec ± 150 kHz (LHC) (± push/pull)
Notes: Remark about elastic limit used

## Slide 3: Linear drive set-up, Resolution, speed
Motor: 1.3 Nm Bipolar Nema 23 (1.8°)
HD: HFUS-20-100-2SO, ratio i:0.01
Roller screw Rollvis RV 12 x 1 mm lead
EDMS 1994455 SPS DQW Tuner documentation
Ratio Motor displacement/cavity deformation: 1.5
System stiffness: 1.6 kN/mm
DQW:
Measured at 2 K (SM18)
0.23 Hz/ step resolution
Driver 8000 microsteps
Maximum tuning speed: 2000 steps/ second, because not ramped
Hysteresis present
25-50 Hz

## Slide 4: Instrumentation
Load cell Kistler 5 kN (DQW, RFD 10 kN))
4576A55C1 class 0.1
Variables Timber:
Tuner force: ACF.POW.C% :SG_RAW in N
Position: ACF.TUNER.C%:Position in steps
Remark about zeroing position
Potentiometer Megatron RC13-25 M no longer used, resolution not good
Plan is to replace this by resolver with conditioner
Interlock tuner DQW SPS : limit switches, Load cell at ± 2 kN
Request: function to move to zero load position when T He tank > 100K

## Slide 5: Tuner (load cell) possible perturbations “Drift”
Feedthrough bellow: sensitive to pressure (atm)
1% change atm >> ~ 3N *
Temperature gradient, change will change tuner load
Thermalisation temperature not measured, refer to FPC thermalisation
He Tank feedthrough + cavity surface
Sensitive to He tank pressure (and temperature when saturated liquid)
1 mbar >~3 N*
* Rough estimate

## Slide 6
Cavities at 2 K, no tuner changes
Temp He tank
So also pressure
Temp Thermalisation FPC
2nd ripple frequency fits with cryo cycle?

## Slide 7
Temperature tunnel
Cavities at 2 K, no tuner changes

## Slide 8
CM stable at room temperature
Check T influence on load cell + cables
No thermal gradient inside CM means no change of tuner force with tunnel temperature

## Slide 9
15-18 October
? Feedback with tuner
? Input
Comment: Orange curve are steps sent to the tuner. If no steps are sent there are no points stored in timber. So there was clearly a loop running.

## Slide 10
15-18 October
Influence tunnel temperature
+ indications cryo cycle on temperature tunnel
Influence thermalisation temperature
What was the input of the FB control ? Frequency, Power 2 K cryo?
This would indicate that the frequency drift is related to tuner drift and hence temperature drift

## Slide 11: MD

## Slide 12
Temperature Cavity
MD

## Slide 13
MD
Tunnel temperature

## Slide 15: Conclusions
Drift seen on tuner load cells, possibly related to observed frequency drift of cavity, can be attributed to tunnel temperature and to a lesser extend to the 2 K cryo cycle.
Some of the 2 K cryo cycle is found back in the tunnel temperature.
Drift can be compensated with the tuner as demonstrated on 15th to 18th October
Requests to implement in Tuner control: back to zero > 100 K, tuner speed ramp

## Slide 16: Thank you for your attention!

## Slide 17: Spare slides

## Slide 18: Reminder functioning tuner

## Slide 19: Reminder end of Run1 DQW SPS 2018
Sudden increase of stiffness 2-3 October after thermal cycle
no real blockage
One motor-gear coupling started slipping 19/10
Tuner heaters broken + wires damaged
Possible ice formation
Very hard to dismount motor in-situ or to replace the heater
Impossible to retighten the coupling without full actuator disassembly
No access to set limit switches and hard stop
Potentiometer not needed
Important observations:

## Slide 20: Testing
Elastomer
cycles Up to 7 kN
SPS Room temperature
Same stiffness as SM18

## Slide 21: SPS DQW and RFD Upgrade
D-type slip-free Oldham coupling with set screw
Introduction connection clamp
Creation more space
Lowered actuator height
Removed potentiometers
Moved limit switches to front
Only 1 hard stop in front