## Slide 1
Medica-Plus: a Micromegas detector for sub-Bq tritium activity assessment at the service of oncological research
A. Cools, E. Ferrer Ribas, T. Papaevangelou, M. Vandenbroucke
CEA-Saclay/IRFU/DEPHYS
F. Beau, C. Malgorn, F. Malloggi
Institut Joliot Curie and Iramis
CEA-Saclay
1

## Slide 2
PLAN
2
CONTEXT
Biological context
Radiolabelling techniques
Medica-Plus steps
EXPERIMENTAL SET-UP WITH A MICROMEGAS DETECTOR
RESULTS
First experiment
Second experiment
First activity measurement method
Second activity measurement method
CONCLUSION and PERSPECTIVES

## Slide 3
Recent progress in oncology : tumoral heterogeneity !
Variability between tumours
Different cell types
Decrease targeting drug efficiency
Oncological research issue : heterogeneity effect on drug targeting ?
Tumoral heterogeneity requires better detection sensibilities
3
Tumour heterogeneity

## Slide 4
Molecule labelling and tracking
Measure the drug concentration per single cell (estimated 0.5 Bq)
Several techniques available
Fluorescence → large groups
Isotopic marking does not modify drug’s chemical behaviour!
Pharmaceutical needs at the cell level for drug development :
Assess the drug distribution within cancer vs. host cells
Evaluate the impact of the cell heterogeneity on drug biodistribution
4

## Slide 5
Molecule labelling and tracking
Medica-Plus detection intention
At the cell level : Quantification of 3H concentration in single cell samples
An innovative cooperation between the involved disciplines !
Measure the drug concentration per single cell (estimated 0.5 Bq)
Several techniques available
Fluorescence → large groups
Isotopic marking does not modify drug’s chemical structure!
5

## Slide 6
Medica-Plus steps
6
Tumoural cells collected from an animal – Tritium tracking
Cellular culture
Cell deposit by microfluidics techniques
Tritium activity counting with gaseous detector

## Slide 7
Low energy threshold
Good spatial resolution
Good background rejection capabilities
For this application: 3H samples inside the gas volume
Detection of 3H with micromegas
7
Active zone 12 x 12 cm2
Unidirectional copper strips, 500 µm pitch
Gas : Ar-5% Isobutane
DREAM electronics : self-triggering mode
Acker et al., NIM A 957 2020 163423
β
Theoretical 3H energy spectrum
3H Beta range
Medica-Plus detector
Micromegas for 3H counting : samples deposited on replaceable drift electrode

## Slide 8
Medicaplus results
1st deposit with tritiated glucose : 0.4 Bq – 40 Bq
Proof of concept
2nd deposit with low activity cells and glucose : 0.3 Bq – 100 Bq
Clustering
1st method : Integration algorithm
2nd method : Peak association algorithm
Methods comparison
8

## Slide 9
Cathode
12 drops from 0,4 Bq to 40 Bq
Glucose drops
Tritiated glucose signals positions
Glucose drops deposit on the cathode
9
1st deposit with tritiated glucose : 0.4 Bq – 40 Bq
1
3
5
7
9
11
2
4
6
10
11
8

## Slide 10
1st deposit with tritiated glucose : 0.4 Bq – 40 Bq
Common noise suppression
Low amplitude cut (4σ)
ADC
ADC
Time samples (8 ns)
Time samples (8 ns)
Amplitude signals for one event
Amplitude signals for one event
10
Medica-Plus detector
Measured 3H amplitude spectrum

## Slide 11
Tritiated glucose events positions
Basic algorithm without clustering
Activity measurement for Micromegas :
Better error and linearity than β-imager in full range mode
Similar results compared to reduced range mode
Proof of concept
11
F. Jambon et al. NIMA 1027 (2022) 166332

## Slide 12
0.3 Bq cells
High activity
Low activity
5 mm
1 cm
12
Cathode
2nd deposit with low activity cells and glucose : 0.3 Bq – 100 Bq
Medicaplus results
Glucose and cell deposit on the cathode
Medica-Plus mylar drift electrode with samples

## Slide 13
Clustering
Raw data : signals positions
6 high activity and 6 low activity
6 very low activity cells
Clustering
13
Clusters centroid
Counts
Counts
Strip number
Strip number

## Slide 14
Event multiplicity
Histogram of the number of clusters per event
Number of clusters
Events
Events which don’t fulfil cluster conditions
Single cluster events
Multi clusters events
Tritium multi clusters probability : 0.002%
14
Log scale

## Slide 15
Activity measurement – 1st method : Integration
Blue and green lines : left and right limits of the gaussian pic→ ± 4σ.
Blue curve : Histogram of the number of events.
Red curve : Gaussian fit of the histogram.
Activity measurement : Ratio between the integral of the fit and the time of the full acquisition.
Log scale
15
Gaussian fit and integration for activity measurement
Single gaussian fit

## Slide 16
Activity measurement – 2nd method : association Algorithm
The algorithm evaluates a peak activity at every step of the events loop
Initialisation : Creates a peak with the first identified cluster
Event loop
Peak
- Position
- Number of clusters
- Time
- Activity =
If the event cluster’s position is close enough to a peak position :
→ Associate the cluster to the peak
If the cluster is too far from the peak
→ Creates a new peak → Initialisation
Data from the fit
Data from the algorithm
Number of clusters
Strip number
16
Activity measurement by peak association

## Slide 17
At the beginning : Reliable activity and position but some peaks are missing.
At the end : Activity and position have a better precision.
Number of clusters
Strip number
Activity measurement – 2nd method : association Algorithm
17

## Slide 18
Integration (Bq) | Integration statistical error | Association algorithm (Bq) | Association statistical error | Methods accordance
59,07 | 0,003 | 65,04 | 0,003 | 90.8 %
31,24 | 0,004 | 40,66 | 0,004 | 76.8 %
51,74 | 0,003 | 54,03 | 0,003 | 95.7 %
64,33 | 0,003 | 67,89 | 0,003 | 94.7 %
62,12 | 0,003 | 64,66 | 0,003 | 96 %
57,87 | 0,003 | 60,47 | 0,003 | 95.7 %
0,38 | 0,042 | 0,43 | 0,040 | 88.3 %
0,37 | 0,043 | 0,4 | 0,041 | 92.5 %
0,26 | 0,051 | 0,31 | 0,048 | 83.8 %
0,23 | 0,054 | 0,24 | 0,053 | 95.8 %
0,32 | 0,046 | 0,35 | 0,044 | 91.4 %
0,31 | 0,047 | 0,33 | 0,045 | 93.4 %
Activity measurement – Methods comparison
18
Methods agree at 90%
Methods match 0,3 Bq

## Slide 19
CONCLUSION AND PERSPECTIVES
Medica-Plus with standard micromegas
Validation of an experimental set-up with tritiated cells and glucose deposit
Efficient clustering and several activity measurement methods validated
Activities below 1 Bq were measured over 6 samples
Next steps :
Explore the lowest activity detection threshold of the detector
Use other gases (CF4) to limit transverse diffusion over a factor 2
Perspectives : Optical readout
19

## Slide 20
OPTICAL READOUT
Transmission X-ray radiography of a small bat with glass Micromegas detector,
F.M. Brunbauer et al. NIMA 955 (2020) 163320
Glass Bulk detector for Micromegas detector built at CEA-Saclay
Record scintillation light emitted during electron avalanche multiplication
Full 2D pixelized readout for high spatial resolution
Availability of commercial suitable cameras
Flexibility with lenses and mirrors for adjustable magnification and camera positioning
20

## Slide 21
BACK UP
21

## Slide 22
22

## Slide 23
Medicaplus results - CLustering
Strip number
Cluster Position Histogram
Clustering process
Trigger on one of the 256 strips
Records 25 samples for each channel
Analyses the signal amplitude of each sample for every strips
Associates the signal to a cluster if the amplitude is above the threshold
Compute the cluster position, amplitude, centroid...
Amplitude cut-offs
Common noise suppression
Online
Offline
Log scale
23

## Slide 24
Activity measurement – 3rd method : Time difference
Strip number
Number of clusters
Measurement of the mean time difference between two consecutive events of a same peak.
The inverse of the time difference gives the activity.
24