## Slide 1
Qualification and Modeling of the Non-Linear Response of the First Large-Area DEPFET Pixel Sensors with Internal Signal Compression
A. Castoldi, C. Guazzoni, M. Ghisetti, Politecnico di Milano & INFN
S. Maffessanti, K. Hansen, DESY
S. Aschauer, L. Strüder, PNSensor GmbH
Y. Ovcharenko, D. Lomidze, M. Porro, European XFEL GmbH
+ acknowledments to many people at XFEL and DSSC Consortium

## Slide 2
Module Box
P < 4.5 W
Board Cooling
P < 27 W
Chip Cooling
Ladder
256x128
Quadrant
Power
Signals
& Clocks
x-y Gap
128 x 256 Pixel Sensor
21 cm (1024 pixels)
unscattered beam
Main features
based on non-linear DEPFET sensors
200 µm hex pixel (1024 x 1024)
photon energy 0.5-6 keV
single-photon sensitivity
dynamic range 10,000 ph
frame rate 0.9 – 4.5 MHz
digitization/storage during bunch train
~800 stored frames/macrobunch
Power cycling
11 kW peak power
240 W average power
DEPFET Sensor with Signal Compression (DSSC) 1 Mpix camera
 accurate calibration of multiple DSSC properties is requested (gain, offset, QE, charge-sharing, etc)
 of particular importance is the qualification of the NL system response over the full dynamic range

## Slide 3
DSSC adaptation of the DEPFET concept:
 internal gate extends under the source
 at high levels, signal charges gradually spread under source
 non-linear ΔI/Qsig curve
 gain curve engineering by dose & geometry of implantations
CMOS-foundry DEPFET technology (PNSensor GmbH)
Aschauer, et al. JINST (2017)
Signal compression at sensor level
Low noise (low capacitance)
Fast collection
Signal compression
(but internal gate not accessible from ASIC)
DePFET:
NL pixel response
ASIC:
linear response
DSSC:
NL system response
High dynamic range & single photon sensitivity
136 µm
DEPFET Active Pixel

## Slide 4
Gain and offset adjustable pixel-wise
11 settings of Cfcf (coarse gain) for photon energy/dyn. range selection
6 bits of ADC gain fine trimming (nominal accuracy 2%)
4 bits for offset trimming (1.5 LSB range with 8% granularity)
Pixel readout channel
IRAMP (64)
CFCF (11)
PxDelay (16)
CS&H
Eph
Readout concept
Full-parallel readout
Analog shaping
In-pixel 8-bit digitization (9-bit @ f<4.5MHz)
In-pixel SRAM (800 frames/burst)
Data transmission during gaps
  possibility of fine gain trimming @pixel level
  limited ADC resolution

## Slide 5
Single photon sensitivity & high dynamic range
Under the constraints of maximum dynamic range and 8-bit resolution, single-photon counting is possible by setting DEPFET signal for 1 photon equal to 1 (or n) ADC bin
  signal compression will extend the DR
offset and gain in the read-out ASIC must be set for the correct incident photon energy before a scientific experiment is performed
  not possible to re-gain single photon counting by a re-calibration of the data
inevitable process variations in the sensor as well as in the read-out ASIC
  each DSSC pixel must be calibrated individually (~1% accuracy)
0
1
2
3
# photons
Energy
Analog signal
ADC bins (ADU)
gain=1 ADU/Eph
Linear region
Signal compression
n
(n-1)
(n+1)

## Slide 6
Calibration strategy
Two-step procedure:
calibration of gain/offset in the linear region for all settings
qualification of the shape of non-linear (NL) DePFET response.
Detailed tasks:
hard X-ray spectra, high-gain mode setting (Cfcf=1)
  DSSC-PulXar setup
LookUpTable of ASIC gain values (gain ratios)
  charge-injection circuit (13bit)
  Cfcf=11 settings, ADC fine gain (2x64 settings)
qualification of the NL response
  few options: optical (LED/laser), leakage current, XFEL beam
  only 1 gain setting
# photons
ADU
# photons
ADU
# photons
ADU
Cfcf=1
Cfcf=2
Cfcf=3
ADCFine=1:64
...
  validation of the calibration procedure at ladder level

## Slide 7
Cu Kb
Cu Ka
Calibration campaign of the 1st DEPFET ladder with X-rays
X-ray spectra fitting @ high gain
8-bit resolution impacts on accuracy of gain/offset determination
characterization/modeling of DEPFET spectral response (high-res 14-bit test bench)
  charge-sharing (i.e. low-energy peak tail)
  bremsstrahlung
  bin integration
gauss fit (magenta):
a1=740541 ± 317
sig1=0.293
xc1=30.657
gauss+binint (green):
a1=864818 ± 0.5
sig1=0.2005
xc1=30.718
impact of bin-integration in Pedestal peak fit
DN160 flange + 2m flight tube in front of DEPFET ladder
PulXar source @XFEL DET Group
600 pulses @ 50ns @ f=2.25 MHz
Cu anode, 25 kV
Al filter 50 mm
beam
FENICE vessel/DEPFET ladder
DEPFET ladder (128 x 512)
128 x 256 pixels
128 x 256 pixels
Cu Ka, Cu Kb:
(gauss + tail + shelf)
bremsstrahl.
+ beam filtration
Pedestal
Cfcf=1
Id=100mA
T=18 C
2.25 MHz

## Slide 8
gain
kbrems (a.u.)
gain (ADU), rows 0:4
kbrems (a.u.), rows 0:4
bremsstrahlung (pixel intensity)
gain (ADU)
Gain / noise on DEPFET ladder @ high gain (PulXar spectra)
Gain (ADU)
mean=39.2 ADU
(205 eV/ADU)
Cfcf=1
Id=100mA
T=18 C
2.25 MHz
Gain (untrimmed)
small gain spread (3.5%)
std_error(gain) <1%
Noise
ENC 14 ± 4 el rms @2.25 MHz
noise
ENC (electrons)
ENC (el)
mean=14.4 el
st.dev=4.3 el
Cfcf=1
Id=100mA
T= 18 C
2.25 MHz
st.dev=1.4 ADU (3.5 %)

## Slide 9
Qualification of the Non Linear DEPFET response - technique
gas
monitor
(GM)
(XRF)
sample
variable
attenuator
[k]
pulse energy
(E0=N0*Eph)
=
DSSC Ladder/Camera
# photon
ADU
~1000x1keV
XFEL beam on target: rationale
  XRF emission proportional to beam intensity
  pulse of N low-energy XRF photons (e.g. 1 keV) produce a “reference line”
  to scan the NL curve we change beam attenuation (i.e. gas transmission)
  irradiation of one full quadrant
Constraints
  150nm Al layer on entrance window (optical block)
  technique applicable to 1 Mpix camera
Feasibility tests @ XFEL SQS beamline
  SDD + pinhole to define photon flux
  high spectroscopy resolution allows accurate intensity calibration (photon# peaks)
  verified linearity range of (XRF counts vs. beam intensity)
SDD spectra at different beam intensity
gasTx=10%
gasTx=30%
Al target
(Al Ka 1.486 keV)
SDD @ 1m (target) 100mm pinhole
setup @ SQS

## Slide 10
Beam time @ SQS beam line (Oct. 2021)
Goal: verification of NL response measurement technique
irradiation profile of Al Ka photons
nice flat profile
 possible irradiation of 1 quadrant (=4 ladders)
DSSC mounted @ 40cm from Al target to maximize photon rate
SQS instrument must provide a stable beam
reference monitor of beam intensity (x-ray gas monitor)
Tests at low intensity & high gain setting (linear region)
Energy spectrum Al Ka photons
SQS 1.6 keV Al Ka (1.486 keV)
Cfcf=1
2.25 MHz
n=0
1
2
3
single-photon resolution (S/N ~ 30:1)
Al target
FENICE vessel
(DEPFET ladder)
Al Ka
XFEL beam
single-photon resolution: Al Ka photon peaks well separated (signal/noise ~30:1)
ENC from SQS spectra (14.1 el.) confirm values obtained from PulXar (14.4 el.)

## Slide 11
Non linear response of DEPFET pixel
acquisition of flat-field images of Al Ka photons vs. gasTx
X-ray Gas Monitor tracks beam intensity incident on the target
  XGM data (electron-based): fast (single-frame)
  XGM data (ion-based): slow (20s), accurate (down to few keV/pixel)
  non-negligible beam intensity variations (“blue” weeks @ XFEL)
energy calibration obtained by matching the response in the linear region with the measured DEPFET gain (from X-ray spectra)
DEPFET
XGM
(ion-data)
XGM (el-data)
run 54-84
pixel (10,400)
gain setting Cfcf=3 readout rate 2.25 MHz
Cfcf=3, Id=100mA T=18 C
pixel (10,222)
histogram
centroid (mean)
histogram
st.dev.
fit line
histograms of a single pixel (all runs)
Intensity ”scan”
successful qualification of the whole compression curve
maximum deposited energy about 3 MeV (~2000 Al Ka/pixel)
single-frame data (50 trains/run)
NL response curve vs. deposited energy

## Slide 12
Fitting model of the NL response
A single gain step described by an elementary function with 3 fit variables (Ea, Eb, n), representing energy boundaries and amplitude of the gain step
good-of-fitness: rmse <0.02 ADU, R2bar=1.0000, std error of fit variables <1%
link from NL response model to DEPFET technology qualification
E1a
E1b
E2a
E2b
E3a
gain (running avg.)
deposited energy (keV)
DePFET Gain + fit line (green)
E1a
E1b
E2a
E2b
E3a
deposited energy (keV)
signal current (ADU)
DePFET Signal + fit line (green)
Id=100 µA
Vds=-3 V
Vback=-200 V
T=-20 C
Ea
n
gain step parameters
xof1
xof2
xof3
qDj1
qDj2
qDj3
-qj(x)
Dx1
Dx2
Dx3
1d cut
along the
IG implant
(qualitative)
CMOS-DEPFET technology
fit line
Eb

## Slide 13
NL response – all DEPFET ladder
Ek1 (keV):
mean=81.1
st.dev.=5.82
(7.2%)
data analysis/calibration over all pixels
non-linear fit function
extraction of all fit parameters pixel-wise
mean value (81.1 keV) in line with lab tests
nice uniformity: 7% spread over the ladder
Ek1 (keV)
first compression “kink” (Ek1) of the NL curve
entries: 63429 pixels
gain: Cfcf=3, Id=100mA
T=18 C
average curve
NL response curves of all ladder

## Slide 14
Dynamic range
NL response more compressed than ideal shape (sqrt(x)), marginal increase of quantization noise
single-photon counting -> gain = 1 photon/ADU (0.5 keV/ADU in the linear region)
 DR = 2820 photons @ 4.5 MHz (7500 photons @ 2.25 MHz)
Eph=0.5 keV
gain=1 ph/ADU
ID=100uA
ADC bin width (#ph rms)
Poisson noise (#ph rms)
~7500 ph @ 285 bin
~2820 ph @256 bin
  maximum #photons at full scale of ADC, e.g. for 8 bit resolution @256ADU (or 512ADU if f ≤ 2.25 MHz)
  maximum #photons for which ADC resolution equals Poisson noise
E.g. photon energy 0.5 keV
ADC bin width =f(E)
1410 keV/0.5 keV =2820 ph
bin=256
gain=1 ph/ADU
pix 10,400
ID=100uA

## Slide 15
Conclusions and outlook
combination of single-photon counting & high dynamic range with limited ADC resolution reflects in a significant challenge for calibration, directly linked to the performance of the DSSC detector
the experimental qualification of first DEPFET ladder gave a very successful feedback:
  noise level of 10-20 el rms
  gain/noise accurate calibration from PulXar spectra @ Cu Ka (8 keV), confirmed at SQS @ Al Ka (1.48 keV)
qualification technique of the NL response @ SQS (XFEL beam on Al target) successful
  fulfilled required intensity (~3 MeV/pixel), quadrant-size field of illumination
  DEPFET response qualified over the whole ladder
  beam intensity variations observed (“blue week” not “user week”), XGM data and correction techniques under study to minimize impact
dynamic range @ 0.5 keV photon energy: DR~3000 ph @ 4.5 MHz or ~8000 ph at 2.25 MHz.
so far so good, but a lot to do next...
  sensor response as a function of photon energy (< 1 keV), correction for ADC DNL, optimization of the compression shape vs. DEPFET operating conditions, imaging properties, etc.

## Slide 16
spares

## Slide 17
Imaging properties at 1.48 keV (Al Ka photons @ SQS) - preliminary
Slit 0.36 mm
Gap 0.65 mm
Slit 0.60 mm
Gap 0.90 mm
Slit 0.85 mm
Gap 1.16 mm
2.5 mm
2 mm
1.1 mm
1. mm
0.750 mm
0.373 mm
Pix 37449 (73,73)
Run161
(~31 keV/pxl)
Run167
(~2 MeV/pxl)
Run175
(~5.4 MeV/pxl)
gain Cfcf3
~46 ph/pxl @gasTx=1%
Steel, 0.5 mm
image details well resolved for all intensities up to 5.4 MeV (colormap rescaled)
work in progress
average of 25 cols
erf fit: s=0.31 pix
Run 160 gasTx 0.1% (~7keV/pix)
Run161, E~31 keV/pixel
Run167, E~2 MeV/pixel
Run175, E~5.4 MeV/pixel