## Slide 1: Gain, noise, and collection efficiency of GaAs SAM-APDs with staircase structure by means of synchrotron radiation
Colja M.*,1, Cautero M.1 , Steinhartová T.2,8, Pilotto A.3,6, Antonelli M.4 , Cautero G.4,5 , Menk R. H.4,5,7,
Arfelli F.4,9, Dal Zilio S.2 , Biasiol G.2 , Gianoncelli A.5 , Bonanni V.5 , Palestri P.3 , Driussi F.3
1 DIA, University of Trieste, 34127 Trieste, Italy
2 IOM CNR, Laboratorio TASC, Area Science Park Basovizza, 34149 Trieste, Italy
3 DPIA, University of Udine, Udine, Italy
4 INFN-Trieste, Padriciano, Trieste, Italy
5 Elettra Sincrotrone Trieste S.C.p.A., Area Science Park Basovizza, 34149 Trieste, Italy
6 C2N, Université Paris-Saclay, 10 Boulevard Thomas Gobert, 91120 Palaiseau, France
7 Department of Medical Imaging, University of Saskatchewan, Saskatoon, Canada SK S7N 5A2
8 Infineon Technologies, Siemensstraße 2, 9500 Villach, Austria 9. DF, University of Trieste, 34127 Trieste, Italy
matija.colja@elettra.eu
28/06/2022

## Slide 2: Introduction
increasing demand for fast and efficient X-ray detectors
silicon-based detectors widely employed, but are not efficient in hard X-ray range
we are developing
GaAs SAM-APDs
with staircase structure
research of new materials
(Ge, compound semiconductors, …)

## Slide 3: Why GaAs?
 | Si | Ge | GaAs
density [g/cm3] | 2.33 | 5.323 | 5.32
effective atomic number | 14 | 32 | 32
electric-breakdown field [V/cm] | 3×105 | 1×105 | 4×105
electron mobility [cm2/Vs] | 1350 | 3900 | 8000
band-gap [eV] | 1.12 | 0.66 | 1.42
GaAs SAM-APDs with staircase structure
shorter attenuation length
shorter response time
possibility to operate at room temperature

## Slide 4
low fluxes / single photon detection
very weak signals (thousands of electrons)
need for amplification
impact ionization = internal amplification
Why avalanche?
GaAs SAM-APDs with staircase structure

## Slide 5: Why Separate Absorption and Multiplication?
photon absorption in a region with a high electric field = noise contribution
separation is obtained by a
highly doped layer
CHARGE LAYER
Depth
absorption
region
mult
photon
charge
sub
Electric field
GaAs SAM-APDs with staircase structure
region for photon absorption (with low electric field)
ABSORPTION LAYER
region for electrons multiplication (with high electric field)
MULTIPLICATION LAYER

## Slide 6: Why staircase structure?
multiplication in a region with similar electron/hole ionization coefficients
=
noise contribution
𝛼/𝛽GaAs ≅ 1 high noise
to reduce noise: band-gap engineering
staircase structure
(impact ionization of electrons at discrete locations)
[David, J. The staircase photodiode.
Nature Photon 10, 364–366 (2016)]
GaAs SAM-APDs with staircase structure

## Slide 7: The detector – cross section
Substrate
Multiplication region
Absorption region
Charge layer
Cr/Au contact
p-GaAs
GaAs
GaAs
AlxGa1-xAs
Al0.45Ga0.55As
GaAs
n-GaAs
n-GaAs
5×1018
2×1016
-5×1017
-2×1018
200 nm
300/4500/15000 nm
35 nm
20 nm
25 nm
35 nm
…
×12
100 nm
500 um
Ge-Au/Ni/Au n-contact
1 um
ẟ p-doping
Doping concentration [cm-3]
Digital alloy
aluminum content changes
from 45 % to 1 %

## Slide 8: The detector – cross section
Substrate
Multiplication region
Absorption region
Charge layer
Cr/Au contact
p-GaAs
GaAs
GaAs
AlxGa1-xAs
Al0.45Ga0.55As
GaAs
n-GaAs
n-GaAs
5×1018
2×1016
-5×1017
-2×1018
200 nm
300/4500/15000 nm
35 nm
20 nm
25 nm
35 nm
…
×12
100 nm
500 um
Ge-Au/Ni/Au n-contact
1 um
ẟ p-doping
Doping concentration [cm-3]
Digital alloy
aluminum content changes
from 45 % to 1 %
PROTOTYPES
not sufficiently thick for hard X-rays
expressly developed to investigate recombination in:
  metal-semiconductor interface
  absorption region

## Slide 9: Characterization with laser
Eph = 2.33 eV
initial exponential current growth due to lowering of the charge layer potential barrier
Vbreakdown = 37 V
multiplication from 25 V to breakdown
single e/h pair production
no photogeneration in the multiplication layer
exponential trend
exponential trend
breakdown
multiplication
200 um

## Slide 10: Gain
Excess Noise Factor
a and b are extracted from I-V measurements interpolating the initial exponential growth
Mmax = 50
Si: measured current spectral density
Iph: DC value of photo-current with M = 1
B: system bandwidth
𝛼/𝛽 = k ∊ [0.25, 0.35]

## Slide 11: Synchrotron radiation measurements
investigation of recombination in:
  metal-semiconductor interface
  absorption region
TwinMic beamline (Elettra Sincrotrone)
  photon energies = [400 eV, 2200 eV]
  sub-micrometric monochromatic beam
APDs
TwinMic experimental chamber
beam

## Slide 12: Metal-semiconductor interface (1)
chemical characterisation (As fluorescence)
identification of areas not covered by top contact
photo-current
As fluorescence
optical
Entry window only partially covered by the top contact

## Slide 13: Metal-semiconductor interface (2)
current measured when the radiation entered:
measured and theoretical transmissions have a similar trend and their values are comparable
traps and defects at the interface have a small effect in the charge loss
  through Cr/Au
  directly in GaAs

## Slide 14: Recombination in absorption region (1)
5 energies (940 eV, 1090 eV, 1500 eV, 1705 eV, 2010 eV)
measurements on 4.5 and 15 um-thick devices only
expected current
Φ0: photon flux [photons/s]
Eph: photon energy [eV]
Ee-h: average energy to produce e-h
transmittance as a function of depth
carriers produced at different distances from the multiplication layer

## Slide 15: Recombination in absorption region (2)
no dependence on the attenuation length -> negligible recombination during electron travelling through the absorption region
variations due to systematic error (reposition of the sample)

## Slide 16: Conclusion
GaAs valid alternative to Si, but careful design required for noise reduction
ENF can be greatly reduced by band-gap engineering
negligible effect in charge collection efficiency of:
  recombination in metal-semiconductor interface
  recombination in absorption region

## Slide 17
Thank you for your attention