## Slide 1: Spectral and phase-contrast imaging: from crystal-based synchrotron setups to detector-based compact systems
Luca Brombal – INFN Trieste
28/06/2022
iWoRiD 2022 – Riva Del Garda (TN), Italy – 26-30 June 2022

## Slide 2: OVERVIEW
Principles
X-ray spectral imaging
  Crystal-based systems
  Detector-based systems
Phase-contrast imaging
  Analyzer-based imaging @ synchrotron
  Edge illumination in the lab
PEPILab: a compact spectral and phase-contrast imaging tool
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022

## Slide 3: X-rays in matter
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
Attenuation
Phase shift
In the wave model the interaction of X-rays with matter is described through the complex refractive index n
3
= incoming plane wave at energy E
= object-modulated wave at energy E
= wave number
Spectral imaging
attenuation
coefficient
Phase-contrast imaging

## Slide 4: X-rays in matter
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
Why spectral imaging?
Why phase-contrast imaging?
4
For light materials (e.g. soft tissues) and energies of radiological interest (10 – 100 keV) phase effects are much larger than attenuation
to increase visibility of soft tissues!
Energy dependence of the attenuation coefficient is sensitive to chemical composition (i.e., difference in atomic number Z)
to separate and quantify chemical elements!
δ ~ phase
β ~ attenuation
δ ~ phase
β ~ attenuation
β ~ attenuation

## Slide 5: OVERVIEW
Principles
X-ray spectral imaging
  Crystal-based systems
  Detector-based systems
Phase-contrast imaging
  Analyzer-based imaging @ synchrotron
  Edge illumination in the lab
PEPILab: a compact spectral and phase-contrast imaging tool
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022

## Slide 6: X-ray spectral imaging
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
6
Spectral imaging requires to probe the attenuation properties of the sample (at least) at 2 different energies
Images acquired at different energies are processed through matrix inversion algorithms to extract (quantitative) maps of elements of interest
Water iodine solution
m
Acquired @ Elettra – bent Laue crystal
Decomposition
algorithm
Low energy
image
High energy
image
Low
High
Water
Iodine
Water-iodine solution
5 mm

## Slide 7: Basis material decomposition
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
7
Attenuation images at each energy bin are written as a linear combination of basis materials mass attenuation coefficients
By inverting the linear system, density maps of basis materials are obtained
Decomposition matrix
Low energy
image
High energy
image
Decomposed maps

## Slide 8: Multi-basis material decomposition
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
8
The algorithm can be extended to multiple energy channels and multiple decomposition materials
energy channels
materials
Inversion performed through least-squares

## Slide 9: Applications: imaging with contrast medium
bone
iodine
Acquired @ PEPI Lab INFN Trieste
IODINE (33 keV)
BARIUM (37 keV)
GADOLINIUM (50 keV)
Panahifar, Arash, et al. Physica Medica 32.12 (2016): 1765-1770.
C T Badea et al 2019 Phys. Med. Biol. 64 065007
1 cm

## Slide 10: Spectral imaging systems
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
10
X-ray spectrum-based
  2 X-ray tubes with different voltages
  Voltage switching
  Dual layer detectors
Detector-based: energy discrimination on pixel-by-pixel basis
Crystal-based: energy dispersive systems based on bent crystals (@ synchrotrons)
Th1<E<Th2
E>Th2
Fascio
Campione
Rivelatore
Beam
Sample
Detector

## Slide 11: CRYSTAL-BASED SPECTRAL IMAGING
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022

## Slide 12: Laue diffraction
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
12
Polychromatic (white) beam
= Bragg’s angle
= speed of light
= Plank’s constant
= lattice spacing
Bragg’s law
Silicon crystal
Detector
Monochromatic beam
Sample
Iodine

## Slide 13: Energy dispersive bent Laue diffraction
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
13
Energy dispersed beam
Detector
By bending the Laue crystal, the impinging beam is focused
Due to curvature Bragg’s law is satisfied at different angles
The beam is energetically dispersed in space!
Wide energy bandwidth
Focal distance
Sample
Polychromatic (white) beam
Bent silicon crystal
Iodine

## Slide 14: Spectral system @ SYRMEP beamline
Fixed radius frames
Radius = 0.5 m Focus ~ 0.25 m
SYRMEP beamline
Energy
Horizontal position
Iodine
Xenon
ΔE = 5 keV
1 cm
dE = 0.1 keV
Largest bandwidth
so far!
14
Pixirad 8 - detector
Sample stage
Bent Laue crystal
White beam
34.6 keV
33.2 keV

## Slide 15: Tomographic multi-material imaging
Xenon
Barium
Iodine
Angular position
(sample rotation)
Xenon
Barium
Iodine
Xenon
Barium
Iodine
Xenon
Barium
Iodine
Tomographic reconstruction
Multi-material decomposition
Goal: simultaneous cardiovascular + respiratory + inflammatory imaging!
iodine
barium
xenon
water
Multi-material density maps
1 cm
Energy
Linear Position

## Slide 16: DETECTOR-BASED SPECTRAL IMAGING
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022

## Slide 17: Detectors in spectral imaging
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
17
Spectral detectors can acquire multiple images over different energy channels in a single shot
High-Z sensors (CdTe, CZT, GaAs) are used for spectral imaging due to their high efficiency at high energies
Modern small-pixel chips feature charge sharing compensation mechanisms ( crucial for spectral imaging applications!)
Spectral µCT setup @ INFN - Trieste
Chip/Producer | Pixel size (µm) | N.o Threshold
Medipix3 | 55 | 2 (8 in 2x2 binning)
Pixirad – PixieIII | 62 | 2
Direct Conversion | 100 | 2
Dectris - Eiger 2 | 75 | 2
… |  | 
Pixirad - PixieIII
F Brun, et al. Phys. Med. & Biol. 65.5 (2020): 055016.
Funded by INFN

## Slide 18: SYSTEM CHARACTERIZATION
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
18
Object composition
Detector response
X-ray spectrum/ Threshold selection
Pixirad1-PixieIII, NPISUM mode:
Differential threshold scan with 33 keV monochromatic X-ray @ SYRMEP
Th1<E<Th2
E>Th2
Input energy spectrum, threshold selection and attenuation coefficients
Spectrum recorded in each channel
Brombal, L., et al. JINST 17.01 (2022): C01043.
Di Trapani, V., et al. NIM-A 955 (2020): 163220.
Material decomposition matrix
Poster
Optimization of energy threshold for fluorescence rejection in K-edge subtraction in synchrotron-based imaging with a spectral CdTe detector
Pasquale Delogu et al.

## Slide 19: Quantitative multi-material µ-CT
Unpublished data – manuscript in preparation
19

## Slide 20: Quantitative multi-material µ-CT
Unpublished data – manuscript in preparation
Composite image
iodine
barium
gadolinum
bone
soft tissue

## Slide 21: Crystal vs Detector-based spectral imaging
 | Crystal-based | Detector-based
Energy resolution | High (~100 eV) | Low (~3 keV)
Multi-material decomposition | Yes (nearby K-edges) | Yes (well separated K-edges)
Sensitivity | High – gold standard (<1 mg/ml) | Medium (function of detector/spectrum)
Setup complexity | High | Low
Scalability | Only in synchrotrons | Everywhere
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
21

## Slide 22: OVERVIEW
Principles
X-ray spectral imaging
  Crystal-based systems
  Detector-based systems
Phase-contrast imaging
  Analyzer-based imaging @ synchrotron
  Edge illumination in the lab
PEPILab: a compact spectral and phase-contrast imaging tool
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022

## Slide 23: Phase effects
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
23
Refraction
(= differential phase contrast)
Ultra-small angle scattering
(= dark field)
Sample
Microstructured sample
y
z
x
×
Within the ray-optical approximation phase effects = refraction
Refraction is proportional to the gradient of δ  strong at the edges
Refraction angles range 1-100 µrad
In microstructured samples multiple-refraction occurs, causing a diffusion of the beam in the range 1-100 µrad
The “amount of diffusion”, i.e. scattering signal, depends on sample’s properties at a scale smaller than the system’s spatial resolution
Phase-contrast imaging = to transform phase effects into detectable intensity modulations

## Slide 25: Analyzer-Based Imaging (ABI)
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
25
An analyzer crystal is introduced downstream of the object to highlight refraction and scattering effects
When perfectly aligned with the monochromator, the analyzer crystal diffracts the beam with 100% efficiency
The diffracted intensity as a function of the misalignment defines the rocking curve

The presence of the sample modifies the rocking curve
Transmission
Refraction
Scattering
Chapman, D. et al., Phys.Med.Biol. 42 2015 (1997)
Arfelli, F. et al., Sci. Rep. 8 362 (2018)
analyzer

## Slide 26: Edge Illumination (EI)
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
26
  In edge illumination (EI) the X-ray beam is structured into mutually independent beamlets by a structured absorbing mask
  The beamlets are analyzed by a second mask close to the detector or by the detector itself
  The detected intensity as a function of the misalignment between the two masks is the Illumination curve

DEI |  | EI
Monochromator |  | Sample mask
Analyzer crystal |  | Detector mask
Rocking curve |  | Illumination curve
M Endrizzi et al. App. Phys. Letters 104.2 (2014): 024106.
Poster
Unified Modulation Pattern Analysis algorithm (UMPA) for 1D sensitive X-Ray Phase Contrast Imaging techniques
Vittorio Di Trapani et al.
mask displacement

## Slide 27: DEI & EI: pros and cons
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
27
DEI
Pros:
High sensitivity (gold standard)
Quantitativeness/monochromaticity
Cons:
Complex setup
Requires laminar beam
Limited to synchrotron facilities
EI
Pros:
Works with spatially incoherent sources
Compact
Cons:
Mask alignment is challenging
“photon hungry”: most photons absorbed by the masks

## Slide 28: DEI: Lung imaging
Menk, R. H., et al. "A novel approach to separate absorption, refraction and scattering in analyzer based lung imaging." Journal of Instrumentation 15.05 (2020): C05069.
Transmission
Refraction/differential phase
Scattering/dark-field

## Slide 29: EI: soft tissue imaging and industrial inspection
Virtual histology
Attenuation
Scattering
Material science/industrial inspection
L Massimi et al. IEEE TMI 41.5 (2021): 1188-1195.
Matsunaga et al., J. Phys. D: Appl. Phys. 53 095401 (2020)

## Slide 30: OVERVIEW
Principles
X-ray spectral imaging
  Crystal-based systems
  Detector-based systems
Phase-contrast imaging
  Analyzer-based imaging @ synchrotron
  Edge illumination in the lab
PEPILab: a compact spectral and phase-contrast imaging tool
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022

## Slide 31: Photon-counting Edge-illumination Phase-contrast Imaging
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
31
In collaboration with:
PEPI is a project funded by the Young researchers grant programme (2020/22260) – National Scientific Committee 5 – INFN AIM: to develop a new laboratory phase-contrast imaging system based on a photon counting detector and the edge-illumination technique with 2 key features:
Flexibility: multiple imaging modalities (spectral/phase-contrast) on a broad energy range (from soft tissue to material science samples)
Compactness: overall dimension < 1m
https://web.infn.it/PEPI/index.php

## Slide 32: Flexibility
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
32
X-ray beam
Sample
Sample mask
Detector mask
Detector
Spectral
Single-Mask EI
Double-Mask EI
Attenuation imaging
Spectral imaging
Sensitivity to refraction (phase) and attenuation
Single shot imaging
Faster (2x or more) than double-mask
Sensitivity to refraction (phase), scattering and attenuation
High accuracy
Features
Features
Features

## Slide 33: Simulation and Design
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
33
A Geant4 simulation platform accounting for the whole imaging chain including:
X-ray refraction
Detector’s energy response
Masks, material and geometry
…
The simulation was key in the choice of:
X-ray source (Hamamatsu L10101 – 40 – 100 kV)
Mask material/thickness (gold 250 µm)
Geometry (magnification 2)
… Open source: https://web.infn.it/PEPI/index.php/it/#simulation
Brombal, L., et al. Journal of Physics D 55.4 (2021): 045102. Brombal, L., et al. JINST 17.01 (2022): C01043.

## Slide 34: Realization
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
34
Photo by Luigi Rigon
How it started…June 2021
How it is…June 2022

## Slide 35: Setup
L. Brombal |23rd iWoRiD 2022 |Riva Del Garda (TN), Italy | 26-30 June 2022
35
X-ray source
Filters for spectral tuning
Sample mask
Detector mask
Pixirad
Sample stage
Mask alignment + slide in/out
Mask alignment + slide in/out
LATERAL VIEW
TOP VIEW
65 cm
50 cm

## Slide 36: Spectral CT
Ex-vivo murine model perfused with iodine-based contrast medium
Bin1 [27, 33] keV
Bin2 >33 keV
Bone map
Iodine map
In collaboration with
1 cm
Voxel size 38 µm

## Slide 37: Double mask
Single mask
Flower low energy – 60 kVp
Smartphone camera high energy – 100 kVp
Transmission
Refraction
Preliminary results
5 mm
5 mm
5 mm

## Slide 38: THANK YOU!
Luigi Rigon (University of Trieste & INFN)
Fulvia Arfelli (University of Trieste & INFN)
Francesco Brun (University of Trieste & INFN)
Ralf H. Menk (Elettra & INFN)
Vittorio Di Trapani (University of Trieste)
Paola Perion (University of Trieste)
Sandro Donato (University of Calabria & INFN)
Alessandro Olivo (University College London)
Marco Endrizzi (University College London)
…
Acknowledgements
Next steps
Spectral phase-contrast imaging
Applications…
Open to collaborations!