## Slide 1
Electro – optical DC Beam Position Monitor
Antonio Cristiano, Michał Krupa
SY-BI-BP

## Slide 2
Outline
Introduction: Evaluation of unbunched beam position
Working principles:
- Pockels effect
- DC electric field detection in the frequency domain
Proof of concept laboratory test bench and its results
Current status
Summary

## Slide 3
Outline
Introduction: Evaluation of unbunched beam position
Working principles:
- Pockels effect
- DC electric field detection in the frequency domain
Proof of concept laboratory test bench and its results
Current status
Summary

## Slide 4
Standard BPM
Screens
Schottky monitor
Evaluation of unbunched beam position

## Slide 5
Standard BPM
Screens
Schottky monitor
DC Insensitive
Evaluation of unbunched beam position

## Slide 6
Standard BPM
Screens
Schottky monitor
DC Insensitive
Destructive interaction with the particle beam
Evaluation of unbunched beam position

## Slide 7
Standard BPM
Screens
Schottky monitor
DC Insensitive
Destructive interaction with the particle beam
Complex
Costly
Evaluation of unbunched beam position

## Slide 8
Outline
Introduction: Evaluation of unbunched beam position
Working principles:
- Pockels effect
- DC electric field detection in the frequency domain
Proof of concept laboratory test bench and its results
Current status
Summary

## Slide 9
V
Electro-Optical crystals and Pockels effect

## Slide 10
V
Electro-Optical crystals and Pockels effect
E=V/d
d

## Slide 11
E=V/d
d
Input
Polarization state
EOpt
V
Output
Polarization state
Laser beam
Electro-Optical crystals and Pockels effect

## Slide 12
Input
Polarization state
EOpt
V
Output
Polarization state
x
z
EOpt
Random polarization
Transmission
direction
Laser beam
Electro-Optical crystals and Pockels effect
E=V/d
d

## Slide 13
Input
Polarization state
V
Output
Polarization state
x
z
EOpt
Transmission
direction
Random polarization
Photodetector
Transmission
direction
EOpt
Laser beam
Electro-Optical crystals and Pockels effect
E=V/d
d

## Slide 14
x
z
EOpt
Transmission
direction
Transmission
direction
Photodetector
V
EOpt
Output
Polarization state
Input
Polarization state
Random polarization
Laser beam
Electro-Optical crystals and Pockels effect
E=V/d
d

## Slide 15
Outline
Introduction: Evaluation of unbunched beam position
Working principles:
- Pockels effect
- DC electric field detection in the frequency domain
Proof of concept laboratory test bench and its results
Current status
Summary

## Slide 16
Bias field [kV/m] | 85.3
External field [kV/m] | 0
Feedback field [kV/m] | 0
x
z
EOpt
Transmission
direction
Transmission
direction
Photodetector
DC Bias
Data processing system
Modulation
signal
Feedback signal
DC electric field detection

## Slide 17
Bias field [kV/m] | 85.3
External field [kV/m] | 0
Feedback field [kV/m] | 0
x
z
EOpt
Transmission
direction
Transmission
direction
Photodetector
DC Bias
Data processing system
Modulation
signal
Feedback signal
DC electric field detection

## Slide 18
Bias field [kV/m] | 85.3
External field [kV/m] | 0
Feedback field [kV/m] | 0
x
z
EOpt
DC Bias
Data processing system
Modulation
signal
Feedback signal
Transmission
direction
Transmission
direction
Photodetector
DC electric field detection

## Slide 19
Bias field [kV/m] | 85.3
External field [kV/m] | 0
Feedback field [kV/m] | 0
x
z
EOpt
DC Bias
Data processing system
Modulation
signal
Feedback signal
Transmission
direction
Transmission
direction
Photodetector
DC electric field detection

## Slide 20
x
z
EOpt
DC Bias
Data processing system
Modulation
signal
Feedback signal
Transmission
direction
Transmission
direction
Photodetector
y
External field to measure
DC electric field detection

## Slide 21
Bias field [kV/m] | 85.3
External field [kV/m] | 30
Feedback field [kV/m] | 0
External field
to measure
x
z
EOpt
DC Bias
Data processing system
Modulation
signal
Feedback signal
Transmission
direction
Transmission
direction
Photodetector
DC electric field detection

## Slide 22
Bias field [kV/m] | 85.3
External field [kV/m] | 30
Feedback field [kV/m] | 0
x
z
EOpt
DC Bias
Data processing system
Modulation
signal
Feedback signal
Transmission
direction
Transmission
direction
Photodetector
External field
to measure
DC electric field detection

## Slide 23
Bias field [kV/m] | 85.3
External field [kV/m] | 30
Feedback field [kV/m] | 0
x
z
EOpt
DC Bias
Data processing system
Modulation
signal
Feedback signal
Transmission
direction
Transmission
direction
Photodetector
External field
to measure
DC electric field detection

## Slide 24
Bias field [kV/m] | 85.3
External field [kV/m] | 30
Feedback field [kV/m] | -38
x
z
EOpt
DC Bias
Data processing system
Modulation
signal
Feedback signal
Transmission
direction
Transmission
direction
Photodetector
External field
to measure
DC electric field detection

## Slide 25
Bias field [kV/m] | 85.3
External field [kV/m] | 30
Feedback field [kV/m] | -38
x
z
EOpt
DC Bias
Data processing system
Modulation
signal
Feedback signal
Transmission
direction
Transmission
direction
Photodetector
External field
to measure
DC electric field detection

## Slide 26
Bias field [kV/m] | 85.3
External field [kV/m] | 30
Feedback field [kV/m] | -38
x
z
EOpt
DC Bias
Data processing system
Modulation
signal
Feedback signal
Transmission
direction
Transmission
direction
Photodetector
External field
to measure
DC electric field detection

## Slide 27
Bias field [kV/m] | 85.3
External field [kV/m] | 30
Feedback field [kV/m] | -38
x
z
EOpt
DC Bias
Data processing system
Modulation
signal
Feedback signal
Transmission
direction
Transmission
direction
Photodetector
The Feedback field to apply, in order to minimize the odd harmonics, is proportional to the External field to be measured.
The proportionality factor depends on the properties (geometry, material etc..) of the crystals.
External field
to measure
DC electric field detection

## Slide 28
Outline
Introduction: Evaluation of unbunched beam position
Working principles:
- Pockels effect
- DC electric field detection in the frequency domain
Proof of concept laboratory test bench and its results
Current status
Summary

## Slide 29
With the collaboration of Franck Guy Guillot-Vignot and Romain Ruffieux
Proof-of-concept laboratory test bench

## Slide 30
With the collaboration of Franck Guy Guillot-Vignot and Romain Ruffieux
Proof-of-concept laboratory test bench

## Slide 31
With the collaboration of Franck Guy Guillot-Vignot and Romain Ruffieux
Proof-of-concept laboratory test bench
External crystal modulated through PCBs

## Slide 32
With the collaboration of Franck Guy Guillot-Vignot and Romain Ruffieux
Proof-of-concept laboratory test bench
External crystal modulated through PCBs
LiNbO3 5%MgO doped
50x2x5mm

## Slide 33
Experimental results

## Slide 34
Outline
Introduction: Evaluation of unbunched beam position
Working principles:
- Pockels effect
- DC electric field detection in the frequency domain
Proof of concept laboratory test bench and its results
Current status
Summary

## Slide 35
Off-centered unbunched beam
…but where is the BPM?

## Slide 36
Electromagnetic DC field lines
Off-centered unbunched beam
…but where is the BPM?

## Slide 37
Electromagnetic DC field lines
Off-centered unbunched beam
Birefringent crystals
Polarizer
Polarizer
…but where is the BPM?

## Slide 38
Electromagnetic DC field lines
Off-centered unbunched beam
Birefringent crystals
Polarizer
Polarizer
Bias voltage
+
Modulation voltage
…but where is the BPM?

## Slide 39
Electromagnetic DC field lines
Off-centered unbunched beam
Birefringent crystals
Polarizer
Laser
Polarizer
Bias voltage
+
Modulation voltage
30 dB dynamic range on first harmonic power compensation.
…but where is the BPM?
Acquisition and post processing system

## Slide 40
Electromagnetic DC field lines
Off-centered unbunched beam
Birefringent crystals
Polarizer
Laser
Polarizer
Bias voltage
+
Modulation voltage
Feedback voltage 1
Photodetector and real time processing system
…but where is the BPM?

## Slide 41
Electromagnetic DC field lines
Off-centered unbunched beam
Birefringent crystals
Polarizer
Laser
Polarizer
Bias voltage
+
Modulation voltage
Feedback voltage 1
Photodetector and real time processing system
…but where is the BPM?

## Slide 42
Electromagnetic DC field lines
Off-centered unbunched beam
Birefringent crystals
Polarizer
Photodetector and real time processing system
Laser
Polarizer
Bias voltage
+
Modulation voltage
Feedback voltage 1
Feedback voltage 2
…but where is the BPM?

## Slide 43
S122C GE Photodiode
Oscilloscope
Signal analyzer
Raw data
Postprocessing
~32 dBm of peak dynamic range
…but where is the BPM?

## Slide 44
DET20C2
Photodiode
24-bit ADC
VFC-HD
1st Harmonic
24-bit DAC
A
f
Control signal
…but where is the BPM?

## Slide 45
Outline
Introduction: Evaluation of unbunched beam position
Working principles:
- Pockels effect
- DC electric field detection in the frequency domain
Proof of concept laboratory test bench and its results
Current status
Summary

## Slide 46
Custom acquisition chain and feedback loop for two optical branches



In progress
Summary
Theoretical study of detection technique
Development of one optical chain mathematical model
Design of a laboratory test bench with one optical chain
Experimental results

Model validation through experimental results

## Slide 47
Thank you for the attention!

## Slide 48
Backup slides

## Slide 49
Search for hidden particle (SHiP) experiment at CERN
Fig.1 Overview of the proposed facility1
Table 1 Cycle parameters foreseen for SHiP1
Fig.2 Position of the proposed facility
1-Ahdida, C., et al. "SPS Beam Dump Facility--Comprehensive Design Study." arXiv preprint arXiv:1912.06356 (2019).

## Slide 50
x
z
EOpt
Transmission
direction
Transmission
direction
Photodetector
DC Bias
Data processing system
Bias field [kV/m] | 0
External field [kV/m] | 0
Feedback field [kV/m] | 0
Modulation
signal
Feedback signal
Length = 40 mm
Cross section = 3x3 mm
CL ≃ 10 pF
Eπ = 198.7 kV/m
Vπ = 597 V
Length = 50 mm
Cross section = 2x5 mm
Eπ = 159 kV/m
Vπ = 318 V
One branch optical chain

## Slide 51
Seller : Photonics Solutions Ltd
Length : 50 mm
WxT : 5x2 mm
Material : MgO LiNbO3(5%)
Coating : AR coating @1550nm at input and output surfaces
Seller : United Crystals Inc.
Length : 40 mm
WxT : 3x3 mm
Material : MgO LiNbO3(5%)
Coating : AR coating @1550nm at input and output surfaces
Coating : Cr, Au electrode coatings on z-planes
Seller : MolTech
Length : 40 mm
WxT : 2x2 mm
Material : LiTaO3
Coating : AR coating @1550nm at input and output surfaces
Coating : Cr, Au electrode coatings on z-planes
Electro-Optical crystals purchased

## Slide 52
DC Field sensing
Different non-contact methods are available to measure DC electrostatics fields
Fig.3 Induction probes
Fig.4 Field mill
Fig.5 Cylindrical field mill
Fig.6 Oscillating plate sensors
Frequent re-zero going needed
More complex and expensive than induction probes
Large field’s source needed
Large size of the mills
More complex and expensive than induction probes
Large dimension ( min150mm x 22mm)
Complex structure
Highly sensitive to charged particles accumulation
Metallic sensors disturbs the electric field being measured
Not compatible with the accelerator environment

## Slide 53
Jones calculus
The state of polarization of a plane wave in the x-y plane, propagating along z-axes can be described as:
When the wave enters in the matter, if the reference crystal is different, the new components will be:
At the end of the matter it accumulates a phase delay proportional to the refractive indices of the material as:
In the reference system of the input light:

## Slide 54
Through the matrix:
The relation between the input light and the output light can be rewritten as :
The values of the components of the matrix M depend on several factor. They are listed for different optical components, as shown in the next figure.
Table 2. Jones matrices
Jones calculus

## Slide 55
Pockels effect
In the case in exam, the electric induction of the material can be written as:
Instead of the permittivity, it is studied the impermeability tensor, that is:
The tensor can be expanded, with respect to the field applied, as :
And can be easily rewritten as:

## Slide 56
Birefringency and Pockels effect
Fig7. LiNbO3 crystal crossed by an optical beam
Fig.8 Ellipsoid of indices of LiNbO3

## Slide 57
Fig.9 Pockels crystal crossed by an optical beam and subject to a slow varying electric field
Electro-optic tensor
Optical beam propagating along y-direction
An important parameter to consider for the geometry of the crystals is:
In case of fig.9
Birefringency and Pockels effect

## Slide 58
Fig.11 The inverse of the sensitivity of the system with respect to the optical beam’s wavelength
Fig.10 Refractive indices value of LiNbO3 with respect to optical beam’s wavelength computed through Sellmeier equation7
7-Zelmon, David E., David L. Small, and Dieter Jundt. "Infrared corrected Sellmeier coefficients for congruently grown lithium niobate and 5 mol.% magnesium oxide–doped lithium niobate." JOSA B 14.12 (1997): 3319-3322..
Optical beam’s wavelength choice

## Slide 59
Fig.13 Absorption coefficient versus optical beam’s wavelength9
Fig.12 Responsivity of commercial photodiodes with respect to the wavelength8
8 - https://www.thorlabs.com/newgrouppage9.cfm?objectgroup_id=3328
9 - Yang, Yunping, et al. "Photorefractive properties of lithium niobate crystals doped with manganese." JOSA B 20.7 (2003): 1491-1502.
1. Responsivity of the photodiode
2. Photorefractive effect
3. Availability of optical components
After a commercial survey, 1550 seemed to be the wavelength with the most of the components available on the market.
Optical beam’s wavelength choice

## Slide 60
Eπ1
Eπ2
Eπ3
A lower value of Eπ has benefits for the crystal:
Requires a lower voltage to be modulated
It has an increased sensitivity to the electric field in which is submerged
Fig.14 The value of Eπ with respect to the voltage applied for different crystal’s lengths
Fig.12 Optical path including polarizer, one LiNbO3 crystal, one analyzer
Electric field
Crystal’s length choice

## Slide 61
Modulation’s amplitude