## Slide 1
Magnetic field penetration method for developing superconducting thin films for SRF
Daniel Turner
Lancaster University
Daresbury Laboratory
The 5th ARIES Annual Meeting, 03-02 May 2022

## Slide 2
SRF limitations
SRF cavities have a large Q0 (1010) and are more efficient than normal conducting cavities for continuous wave applications.
Currently the material used in SRF cavities is bulk Nb.
  Element with the largest Tc.
  Critical fields (approx.) :
    Bc1 =170 mT
    Bc2 =240 mT ≈ Bsh – The superheating field.
  Imperfections on the surface cause localised heating due to the B field, and therefore local quenches.
  Theoretical accelerating gradient limit is 57 MV/m.
  Starting to reach the limits of Nb cavities ~ 40 MV/m.
Erk JENSEN, RF Basics and TM Cavities, SRF 2019 Tutorials
Increasing Eacc allows smaller accelerators to be built, therefore saving money on infrastructure.

## Slide 3
Increasing SRF accelerating gradient (Eacc)
Gurevich proposed to use the multilayer structures to increase Eacc:
  Superconducting-Superconducting bi layer – Consists of a ‘dirty’ SC layer with a larger λL, followed by SC bulk with smaller λL.
  Superconducting Insulating Superconducting structure – SC layers separated by a dielectric/insulating layer.
Kubo, T. (2017). Multilayer coating for higher accelerating fields in superconducting radio-frequency cavities: a review of theoretical aspects. In Superconductor Science and Technology (Vol. 30, Issue 2). https://doi.org/10.1088/1361-6668/30/2/023001
Bi-Layer
SIS Structure
Schematic
Gurevich, A. (2006). Enhancement of rf breakdown field of superconductors by multilayer coating. Applied Physics Letters, 88(1), 012511. https://doi.org/10.1063/1.2162264
Schematic of a SIS multilayer in a VSM setup.

## Slide 4
Increasing SRF accelerating gradient (Eacc)
Gurevich proposed to use the multilayer structures to increase Eacc:
  Superconducting-Superconducting bi layer – Consists of a ‘dirty’ SC layer with a larger λL, followed by SC bulk with smaller λL.
  Superconducting Insulating Superconducting structure – SC layers separated by a dielectric/insulating layer.
Kubo, T. (2017). Multilayer coating for higher accelerating fields in superconducting radio-frequency cavities: a review of theoretical aspects. In Superconductor Science and Technology (Vol. 30, Issue 2). https://doi.org/10.1088/1361-6668/30/2/023001
Bi-Layer
SIS Structure
Schematic
Gurevich, A. (2006). Enhancement of rf breakdown field of superconductors by multilayer coating. Applied Physics Letters, 88(1), 012511. https://doi.org/10.1063/1.2162264
Schematic of a SIS multilayer in a VSM setup.
But first – These structures need to be tested:
  Depositing small samples is faster, easier and cheaper than cavities.
  Flat samples are easier to deposit than 3D structures.
  DC can provide insight into critical fields of superconductors.
  Allows coating parameter optimization, and insight for materials that are potentially useful for SRF applications.
What do we want from a local magnetometer:
To imitate cavity operation – Magnetic field applied from one side of the sample.
Applied field much smaller than the sample to eliminate edge effects.

## Slide 5: Field penetration concept
DC magnetic field parallel to the surface.
Field local to the sample surface:
  Avoid edge effect.
  Allow possibility if sample scanning.
Magnetic field applied from one side of the sample to the opposing side, similar to an SRF cavity.
Applied and penetrated field measured by Hall probe sensors.
HP1
HP2
Sample
Length
Width

## Slide 6: Cryogenic facility
6
Stage 2 – Sample testing stage, T≈2.6 K
Stage 1 – Thermalisation of wires, heat sink for heat shields, and normal conducting to HTS join
Thermal radiation shields
LTS to HTS join

## Slide 7: Type I results - Pb
Method:
Zero field cool-down from above Tc (≈7.2 K).
Increase the applied magnetic field (B1).
Small increase in Hall probe sensor 2 (B2) indicating B leaking around the sample, indicated by line B1K2.
Sharp increase in B2, indicates the field of full flux penetration (Bfp).
After Bfp, the area the magnetic field has penetrated becomes normal conducting, such that B2(B1) lies on B1K1.
Magnet is degaussed, the sample is heated above Tc, and the test is repeated at the next set T.
B1K2
Goodfellows Ltd
Purity of 99.99 %
10 µm thick
50 × 50 mm2 cross section
Note: Bfp is the field at which the applied field has broken through the sample, and not Bc1.

## Slide 8: The effect of geometry
Starting with a sample size of 50 × 50 mm2, the sample was slowly reduced.
Comparing the raw data (left) it can be seen that reducing the sample size increases B1K2.
Reducing the sample ‘length’ has a greater affect on B1K2 than the width.
B1K2 is not always linear, which can skew Bfp results using the normalisation method.
It is still possible to extract Bfp.
Length Width
Multiple Pb sample sizes tested at 3.5 K

## Slide 9: Finding Bfp
The rate of change of B2 is plotted against B1.
The greatest change is defined as Bfp.
Smaller peaks represent smaller jumps – likely the normal conducting area of the sample increasing, allowing more magnetic field through.
Noisy data is smoothed using robust linear regression.
Second derivative method
Goodfellows Ltd
Purity of 99.99 %
10 µm thick
50 × 50 mm2 cross section

## Slide 10: Bfp as a function of T2
Bfp as a function of T2 has a linear trend.
Bfp(0 K) and Tc(0 mT) can be extracted from this graph.
7% decrease in Bfp when the sample size is reduced to 40×30 mm2 (from 50×50 mm2).
10% decrease in Bfp when the sample size is reduced to 30×30 mm2 (from 50×50 mm2).
Therefore sample size should be kept as large as possible to ensure reliable results.
Bfp taken using the second derivative method
10%
7%
Length Width

## Slide 11: Type II results - Nb
(Left) Raw data for 3.6 µm of Nb on a Cu disk with a diameter of 50 mm.
Three Nb samples of varying thickness were deposited at STFC Daresbury Laboratory courtesy of Reza Valizadeh, with thickness’ of 1.3, 3.6 and 5.1 µm to determine the effect of thickness on Bfp.
The transition for Nb is much more gradual than compared to Pb.

## Slide 12: Type II results - Nb
The data is much noisier for the Nb samples.
Thicker samples have an increased Bfp.
More thickness’ need to be tested to try and determine a relationship in Bfp and thickness.

## Slide 13: ARIES WP15 – SRF thin film program
The aim of the program was to develop superconducting cavities using thin film coatings:
  Cu is cheaper than bulk Nb.
  Improved thermal stability.
  Insensitive to the earths magnetic field, so magnetic shielding not required.
  Deposited Nb free from impurities unlike bulk Nb sheets.
The surface quality of the substrate affects the growth of the Nb film.
Cu substrates were polished using; Chemical polishing (SUBU) at either CERN or INFN LNL, electropolishing (EP), both EP and SUBU or Tumbling.
Nb thin films were deposited on the surface of the substrates at INFN Legnaro, University of Siegen and STFC Daresbury laboratory.
Superconducting properties of the samples were then tested using vibrating sample magnetometry (VSM) at the Institute of Electrical Engineering in Bratislava, and a new technique developed at STFC Daresbury.

## Slide 14: The effect of polishing – ARIES samples
13 samples were tested using the field penetration method.
All deposition run had a varying factor.
  STFC samples were 10µm instead of 3µm.
  INFN samples had been laser treated post deposition.
STFC samples (due to a greater thickness) have a greater Bfp.
Samples deposited at INFN LNL have the next best Bfp performance, followed by the samples deposited at University of Siegen.

## Slide 15: The effect of polishing – Aries samples
STFC |  | Siegen |  | INFN (with laser polishing sample surface) | 
Polishing method | Bfp(0 K) [mT] | Polishing method | Bfp(0 K) [mT] | Polishing method | Bfp(0 K) [mT]
EP | 224.7±8.5 | EP | 177.0±6.1 |  | 
 |  |  |  | Tumbling | 205.0±7.8
SUBU CERN | 222.6±0.9 | SUBU INFN | 169.9±5.1 | SUBU INFN | 178.9±2.2
SUBU INFN | 204.8±6.9 | SUBU CERN | 159.7±1.4 | SUBU CERN | 176.7±4.3
 |  | Tumbling | 159.8±1.6 |  | 
EP + SUBU | 191.0±4.3 | EP + SUBU | 134.2±1.1 | EP + SUBU | 148.0±3.7
Largest Bfp(0 K)
Lowest Bfp(0 K)

## Slide 16: Conclusion
An idea for a field penetration measurements has been realised in the facility;
  The facility has been designed, built, commissioned and is now in full time operation.
    A cryogen free system has been designed, built and tested at Daresbury Laboratory
    Applies a local DC field from one side to the other, Tmin = 2.5 K, Bmax ~ 612 mT at I = 8 A
    Data acquisition is fully automated.
The effect of sample geometry has been tested using a Pb sample.
The effect of (Type II) sample thickness has been tested using Nb samples deposited.
Multiple Nb samples deposited by ARIES WP15 partners with various substrate treatments have been tested:
  Magnetic Field Penetration of Niobium Thin Films Produced by the Aries Collaboration (vrws.de)
The ARIES thin film Nb samples have been laser treated and tested, with recent surface analysis measurements. Correlation between the magnetic field penetration method and surface analysis are ongoing.
This work has led to my PhD thesis.
Non-Nb films such as NbTiN and Nb3Sn.
Future plans for the facility within IFAST WP9:
  SIS structures – Test thin films shielding substrate
  Compare DC magnetometry with RF tests
  Final stage – Implement the best films into cavity.
Notes: This is a unique facility built on a novel idea, and it has already proven that it is a useful facility probing into the surface quality of superconductors.
PhD thanks to ARIES collaboration.

## Slide 17: Acknowledgements
CERN: A. Sublet
IEE: E. Seiler, R. Ries
INFN: C. Pira, E. Chyhyrynets
RTU: A. Medvids, P. Onufrievs
Siegen University: M. Vogel, S. Leith
STFC/CI: O.B. Malyshev, R. Valizadeh, K. Dumbell, J.T.G. Wilson, J. Conlon,
L. Smith, F. Lockwood Estrin, F. Walk, N. Pattalwar, S. Pattalwar, A. May
Lancaster University/CI: G. Burt
University of Victoria: T. Junginger
This research has been supported by European Commission’s ARIES collaboration H2020 Research and Innovation Programme under Grant Agreement no. 730871.

## Slide 18
Published papers
Daniel A Turner et al, “No interface energy barrier and increased surface pinning in low temperature baked niobium”, Scientific Reports 12 (1), 1-9 (2022).
Proceedings
Daniel A Turner et al, “Characterization of flat multilayer thin film superconductors”, in proc SRF 2019
Daniel A Turner et al, “Magnetic field penetration of niobium thin films produced by the aries collaboration”, in proc SRF 2021
Oleg B. Malyshev, Daniel A Turner et al, “Main Highlights of ARIES WP15 Collaboration”, in proc SRF 2021
R Valizadeh, AN Hannah, S Aliasghari, OB Malyshev, GBG Stenning, Daniel A Turner, K Dawson, VR Dahnak “PVD Depostion of Nb3Sn thin film on copper substrate from an alloy Nb3Sn target”, Proc. IPAC'19, 2818-2821
Submitted papers
Daniel A Turner, Oleg B Malyshev, Graeme Burt, Tobias Junginger, Reza Valizadeh, Lewis Gurran, “A facility for the characterisation of planar multilayer structures with preliminary Niobium results” – Submitted to Superconducting science and technology (2022)
Soon to be submitted papers
Daniel A Turner et al, “The effect of laser treatment on sputtered niobium thin films on copper substrates using a magnetic field penetration facility” – Submission by the end of this month

## Slide 19
Thank you for your attention

## Slide 20: Finding Bfp
Ratio method:
An infinitely large sample can be described as no B leaking around the sample:
However, as the samples are a finite size B1K2 must be taken into account:
Which can be simplified to:
However, this method is only reliable for samples with a sharp transition for Bfp
Normalisation method

## Slide 21
Data for the Pb sample
A comparison for the leakage factor and extrapolated values for Bfp and Tc using the Normalisation method (Method 1) and the second derivative method (method 2)

## Slide 22
Bfp(0K) as a function of sample length
Pb - Bfp taken using the second derivative method