## Slide 1: CI Update
25/02/2022

## Slide 2: NEG property

## Slide 3: Saturation Estimates
Where
Skimmer | Chamber |  | Saturation time
0.7x9mm | Interaction | 9.45E-09 | 7.7 days
0.3x9mm | Interaction | 3.50E-09 | 20.8 days
0.7x9mm | Dump | 9.50E-08 | 18.4 hours
0.3x9mm | Dump | 4.36E-08 | 40.0 hours
Skimmer | Chamber |  | Saturation time
0.7x9mm | Interaction | 9.45E-09 | 7.7 days
0.3x9mm | Interaction | 3.50E-09 | 20.8 days
0.7x9mm | Dump | 9.50E-08 | 18.4 hours
0.3x9mm | Dump | 4.36E-08 | 40.0 hours
Sorption capacity = 2.26mbar L
Pumping speed
= 360 L/s

## Slide 4: NEG subsystem

## Slide 5: NEG subsystem

## Slide 6: Neg pump installed in the Dump Chamber

## Slide 7: Saturation Data (dump chamber)
Chart: 
Before Jet on
Data missing for weekend due to Computer issue
Jet off

## Slide 8: Saturation Data (interaction chamber)
Chart: 
Before Jet on
Outgas rate of the chamber is dropping
Jet off

## Slide 9: Nozzle chamber
Chart: 
TMP in the nozzle chamber is warm but ok. The pressure even dropped a bit. ??Might be from the fluctuation of the main regulator or gas supply.

## Slide 10: Saturation Data
Chamber | Pressure / mbar | Pressure Change
Interaction | 3.80E-09 | 2.4%
Dump | 6.84E-08 | 27.5%
Chamber | Pressure / mbar | Pressure Change
Interaction | 6.45E-09 | 73.9%
Dump | 1.97E-07 | 268.1%
Predicted Saturation Time (T+40hours)
Most Recent Measurement (T+170hours)

## Slide 11: Summary
Week long data shows the NEG is not saturated for 0.3*9 mm 3rd skimmer (5bar inlet pressure).
The saturation time is longer than the estimation
It is Ok to use the NEG pump in the dump chamber for the smaller skimmer.
Future work:
  Larger 3rd skimmer (0.7*9 mm)
  Preparation for shipping.