## Slide 1: CI Experimental results and plan
BGC CI team

## Slide 2: Section I: Summary from v2 setup
OTR verification
Development of a hybrid simulation code
New configuration for 20 mm long uniform gas jet

## Slide 3: V2 gas jet monitor at the Cockcroft Institute
HIPACE 300L/S
Turbo-molecular pump
HIPACE 700L/S
Turbo-molecular pump
Nozzle Chamber
Skimmer Assembly
Gas Separator
Skimmer Chamber
Imaging System
Third Skimmer
Interaction Chamber
Dump Chamber
Moveable gauge 282 mm from nozzle

## Slide 4: Geometry used for V2

## Slide 5: Gas jet monitor Vs OTR
X
Y
X
Y
Y (mm)
X (mm)

## Slide 6: Hybrid gas jet code from continuous flow to molecular tracking
Gas jet density distribution verified with old configuration at 264 mm from nozzle:
    Nozzle : 30micron; Skimmer I : 180 micron; Skimmer II: 2mm
FWHM = 13.6 mm
FWHM = 15.8 mm
FWHM = 12.5 mm
FWHM = 13.3 mm
Neon
Nitrogen
Experiment
Simulation
X
Y

## Slide 7: Configuration suggested by simulation to generate a uniform distribution gas jet for ~20 mm
Old configuration
  30 micron nozzle
  180 micron first skimmer
  2mm second skimmer
New configuration
  30 micron nozzle
  400 micron first skimmer
  2mm second skimmer
Simulation suggest: uniform distribution after 2nd skimmer at 264 mm from nozzle.
It is verified by the experiment.
Simulation
Simulation
Experiment
Nitrogen
X
Y

## Slide 8: Gas jet density with new configuration
Experiment agree with simulation very well.
Neon
Nitrogen
Decision: New configuration will be used for future V3 test.

## Slide 9: Pressure in each chamber (v3)
Configuration | Nozzle chamber | Skimmer chamber I | Skimmer chamber II | Interaction chamber
Nozzle :30µm 1st skimmer: 180µm 2nd skimmer: 400µm | 6.03e-3 | 1.69e-05 | 5.33e-07 | <5e-9
Nozzle :30µm 1st skimmer: 180µm 2nd skimmer: 2mm | 5.9e-3 | 7.91e-6 | 8.03e-7 | <5e-9
Nozzle :30µm 1st skimmer: 400µm 2nd skimmer: 2mm | 5.56e-3 | 1.13e-5 | 1.52e-6 | <5e-9
Pressure in each chamber given in (mbar) at the specified configuration.
The last configuration offered the highest gas-curtain density(tested at two locations- before and after third skimmer), as well as a more uniform jet.

## Slide 10: Summary for section I
Gas jet image was verified with OTR.
Developed a reliable gas jet density distribution measurement tool.
Developed a hybrid simulation code can simulate gas jet from continuous flow to molecular flow with a good agreement with experiments.
Gas jet curtain criteria could be met by new configuration.
for v4

## Slide 11: Commissioning of v3 gas jet monitor
Setup description
Gas jet beam profile measurement
Gas jet density distribution simulation and experiment

## Slide 12: V3 gas jet monitor at CI
New configuration used
 | Nozzle (0 mm) | 1st skimmer (4.49 mm) | 2nd skimmer (33.59 mm) | 3rd skimmer (168.2mm) | Interaction (393.7 mm)
1 | 30 um | 400 um | 2 mm | 0.7 * 9 mm2 or 0.3 * 9 mm2 | 

## Slide 13: Vacuum control system
Rough pump
Electrical valve
TMP
Pressure Gauge
PC
Solenoid valve (48V control)
Gate valve
power outrage
Normally closed
Normally closed
RS485 protocol
RS232 protocol
USB/Arduino
Pneumatic
USB/Arduino
Status
Gauge
Controller

## Slide 14: Main control GUI

## Slide 15: NEG VS TMP installed in interaction chamber (Nitrogen)
Condition | CH1 | CH2 | CH3 | CH4 | CH5
E-gun & Jet off TMP | 3.8e-7 | 1.2e-8 | 3.1e-8 | 3.2e-8 | 1.3e-8
E-gun off Jet on TMP | 4.9e-3 | 1.03E-05 | 1.16E-06 | 3.97E-08 | 1.37E-07
E-gun & Jet on TMP | 5.5e-3 | 5.3e-6 | 1.6e-6 | 4.58e-8 | 1.2e-7
3rd skimmer
0.7 x 9mm

## Slide 16: NEG VS TMP installed in interaction chamber (Nitrogen)
3rd skimmer
0.3 x 9mm
Condition | CH1 | CH2 | CH3 | CH4 | CH5
E-gun & Jet off TMP | 6.90E-09 | 1.90E-10 | 3.70E-09 | 1.66E-09 | 3.20E-09
E-gun off Jet on TMP | 5.5e-3 | 1.2e-5 | 1.9e-6 | 1.14e-8 | 3.2e-8
E-gun & Jet on TMP | 5.5e-3 | 1.2e-5 | 1.9e-6 | 3.66e-8 | 3.1e-8

## Slide 17: Gas jet measured beam profile
Photon number per second: Neon: ~22s Nitrogen: ~300s
The ratio between Nitrogen and Neon signal level consistent with V2, but is at least 5 times higher
Condition:
0.7 mm *9 mm 3rd skimmer
E-beam conditions: 0.63mA at 5keV
Exposure time per picture : 0.1s
Total integration time :
  Neon: 2000s
  Nitrogen: 200s
Improvement for S/N ratio from V2 to V3

## Slide 18: New Movable gauge system
Pinhole = 1 mm
Y
X

## Slide 19
3rd skimmer
 | Nozzle (0 mm) | 1st skimmer (4.49 mm) | 2nd skimmer (33.59 mm) | 3rd skimmer (168.2mm) | Interaction (393.7 mm)
1 | 30 um | 400 um | 2 mm | 0.7mm*9mm | 

## Slide 20: Convoluting simulation with pinhole Vs Measurement (large skimmer)
Simulation (convoluted)
Measurement
Gaussian fit
Flat top
Sigma = 0.95 ± 0.05 mm
FWHM = 2.2 ± 0.1 mm
Sigma = 0.87 ± 0.07 mm
FWHM = 2.1 ± 0.2 mm

## Slide 21: 3rd skimmmer
 | Nozzle (0 mm) | 1st skimmer (4.49 mm) | 2nd skimmer (33.59 mm) | 3rd skimmer (168.2mm) | Interaction (393.7 mm)
1 | 30 um | 400 um | 2 mm | 0.3mm*9mm | 
Measured distribution

## Slide 22
Simulation (convoluted)
Measurement
Gaussian fit
No Flat top
Sigma = 0.497 ± 0.009 mm
FWHM = 1.17 ± 0.02 mm
Sigma = 0.50 ± 0.02 mm
FWHM = 1.18 ± 0.05 mm

## Slide 23: Convergent-Divergent (CD) Nozzles
CERN workshop EDM and Focused Ion Beam (FIB) divergent nozzle development
FIB nozzle now with Cockcroft institute for tests
Pictures Courtesy A. Cherif

## Slide 24
Old vs New nozzle

## Slide 25: Summary for section II
Design, procurement and commissioning of V3 gas jet monitor
Demonstrated a beam profile measurement with much higher SNR than v2.
New configuration of nozzle-skimmer assembly shows a good gas jet property and easy of alignment.
The difference of CD nozzle and flat nozzle shows little difference.

## Slide 26: Section III: NEG Experiments

## Slide 27: NEG Pump Introduction
CapaciTorr Z1000 installed in the Dump chamber and later re-located to the Interaction chamber, replacing the TMP’s installed in each chamber.
Once activated, the NEG pumps at 360L/s but slows as the getter material becomes coated with the sorbed molecules. Once fully saturated, all pumping will stop.
When reactivated, sorbed molecules are released or diffused into the bulk of the getter.

## Slide 28: NEG Subsystem
NEG inside DN100 pipe
Gate valve to isolate TMP from NEG during normal operation
TMP used to reach activation pressure and pump released molecules during activation.
100mm

## Slide 29: NEG Subsystem
100mm

## Slide 30: NEG Pump Installed in the Dump Chamber
All Jet tests use N2 at 5 bar

## Slide 31: Dump Chamber Pressure
Chart: Total Jet Time / Hours
Data failed to save over weekend. Problem fixed.
0.7x9mm skimmer saturates much quicker than 0.3x9mm skimmer due to increased volumetric flowrate to dump chamber.
  NEG pumping at max 360L/s
  NEG pumping speed reduced as getter surface becomes coated
  Pumping tends towards an equilibrium as NEG saturates and pumping is provided from diffusion effect through the 4th skimmer.
  Using the saturation curve from the larger skimmer, It is predicted that the smaller skimmer will reach an equivalent saturation after 13 days of continuous jet.
1
2
3

## Slide 32: Interaction Chamber Pressure
Chart: Total Jet Time / Hours
Interaction chamber follows a similar pressure curve to the Dump chamber.
There is a clear dependency between the interaction and dump chambers as a result of the diffusion effects through the 4th skimmer.
A maximum pressure of 1.75E-8 mbar is seen compared to the dump chamber maximum pressure of 5.73E-7 mbar.

## Slide 33: Skimmer Chambers I and II Pressure
Chart: Total Jet Time / Hours
Chart: Total Jet Time / Hours
Near constant skimmer II chamber suggests NEG effects are independent of pressure data from nozzle to skimmer II chamber.
The skimmer I chamber is also near-constant with an increase in noise experienced.

## Slide 34: Nozzle Chamber Pressure
Chart: Total Jet Time / Hours
Drop in nozzle chamber pressure of approximately 50% from initial jet.
Anticipated not a function of NEG saturation as nozzle chamber should be fully independent from NEG located in the dump chamber.
Likewise, pressure decrease between different 3rd skimmers should be relatively similar in the nozzle chamber. However, we see a discrepancy of 5 days.
Further testing required.

## Slide 35: NEG Pump Installed in Interaction Chamber

## Slide 36: Pump Down Pressure Data
Chart: Total Pumping Time / Hours
E-7 mbar reached within 15 minutes of gauge on.
Full activation and NEG cooldown takes 3 hours and reaches E-9 mbar.
Anticipated minimum pump down from 1 bar & NEG activation should take 4 hours total to reach E-9 mbar.
Left pumping over weekend. Maximum pressure required to activate NEG is E-4 mbar.
NEG activation and TMP gate valve closed.

## Slide 37: NEG Initial Results
0.7x9mm 3rd skimmer
100s integration time
0.60mA E-gun Emission
N2 @ 5bar
1mm

## Slide 38: NEG Beam Measurement
 | NEG | TMP
Photon Rate / s-1 | 290 | ~300

## Slide 39: Dump, Interaction and Nozzle Chamber Pressures
Chart: Total Jet Time / Hours
Chart: Total Jet Time / Hours
Only 54 Hour Jet Time due to shipping time constraints.
Little to no noticeable change in Interaction and Dump chamber pressures over such a small time.
As with previous measurements, the Nozzle chamber appears to reduce again. Further evidence of a NEG independent trend.

## Slide 40: Future Pressure Tests
Jet will be ran until pressure reaches a steady state and NEG can clearly be defined as saturated.
More pressure data collected for NEG located in the interaction chamber.
Pressure will be recorded with Jet for exclusively TMP pumping to observe nozzle chamber without a NEG installed.