## Slide 1
Fluorescence Monitors at GSI
--- Lessons learned ---
Collaboration Meeting the 30th of March 2022
Peter Forck & Serban Udrea, GSI
Goal: Preparation of experiments at HEL-test-bench and LHC
Experiences from single pass ion beams:
University München TANDEM: 3 MeV/u < Ekin < 5 MeV/u
GSI LINAC: 5 MeV/u < Ekin < 11 MeV/u
GSI synchrotron: 100 MeV/u < Ekin < 800 MeV/u
For peace
and freedom
Violence must
be rejected

## Slide 2
BIF-Monitor: Technical Realization at UNILAC@GSI
Example: BIF station at GSI UNILAC:
Using transition with N2
N2 + Ion  (N2+)*+ e- + Ion  N2+ +  + e- + Ion
Insertion length 25 cm
2 x image intensified CCD cameras
Optics with reproduction scale typ. 0.2 mm/pixel
Gas inlet + gauge, pneumatic actuator for calibration
Advantage: Compact insertion, resolution adaptable
of-the-shelf components
Beam
Vertical BIF
Image Int. CCD
Horizontal BIF
Photocathode
Phosphor
double
MCP

many 
e-
Image intensifier
for single photons
C. Andre (GSI) et al., Proc. IBIC 2014
 15 cm
Beam: Ar10+@4.7 MeV/u, I=2.5 mA1011 ppp, p=10-5mbar

## Slide 3
BIF-Monitor: N2 Fluorescence Spectra for different Ions
C. Andre (GSI) et al., Proc. IBIC 2014
For N2 working gas the spectra for different ion impact is measured:
Fluorescent transitions:
N2 + ion(N2+)* +e-+ ionN2++  + e- + ion
N2+ : B2Σu+ (v’’)  X2Σg+ (v’),
 = 391, 427 & 471 nm,
life time   60 ns
GSI UNIAC: ion energy  5 MeV/u
Results:
Comparable spectra for all ions
Small modification due to N2+
dissociation by heavy ion impact
 Stable operation with
constant spectrum for highly ionized ions
vib. 0  0
vib. 0  1
vib. 0  2
N+ trans.
N+ trans.
1 0

## Slide 4
Nitrogen Molecule Spectra with Ion Impact: Cross Section
N2 as working gas for process:
N2 + ion  (N2+)*+ e- + ion
 N2+ +  + e- + ion
Result for cross section:
Ionic lines: B2Σu+ (v’’=0)  X2Σg+ (v’=0)
 Constant  over entire range
Neutral lines: C3u(v’’=0)  B3g(v’=0)
Increasing 
as generated by secondary electrons only
(spin forbidden by proton impact)
Estimation of cross section:
Ionic: S(2.7 MeV/u) = 4  10-16 cm2
 Proton p(2.7MeV/u) = S(2.7MeV/u) / q2 = 6.3  10-18 cm2
 Bethe-Bloch scaling p(7TeV) = 0.03  p(2.7MeV) = 1.8  10-19 cm2
Serban’s estimation: p(7TeV) = 3.710-20 cm2  factor 5 too large
Neutral: Can not be excited by proton & ion impact!

## Slide 5
Observation: Trans. of ionic states e.g. N2+  profile width independent on pressure
Trans. of neutral states e.g. N2  width strongly dependent on pressure!
Ionic transitions =391 nm:
N2 + ion(N2+)* +e-+ ionN2++ +e- + ion
N2+ @391nm: B2+u(v’’=0)  X2+g(v’=0)
large σ for ion-excitation, low for e-
N2+ trans.@391 nm
N2 trans.
@337 nm
p = 0.1 mbar
N2
p = 30 mbar
N2
F. Becker et al., IPAC’12 &HB’12
Neutral transitions =337 nm:
N2 + e-  (N2)* + e-  N2 +  + e-
N2 @337nm: C3u(v’’=0)  B3g(v’=0)
large σ of e- excitation., low for ions
at p  0.1 mbar free mean path 1 cm!
N2
p = 0.003 mbar
BIF Monitor: Different Gas Pressures and Profile Width
Result: Save operation if p < 10-3 mbar

## Slide 6
Neon Spectra with Sulfur Ion Impact at 3.1 MeV/u
Result for cross section:
Ne+ :  is independent on pressure
Ne & high pressure p > 10-2 mbar:
caused be secondary electrons
Ne & low pressure p < 10-3 mbar:
 is constant
i.e. reflects correct beam profile
Absolute systematic error  sys = 50 %
Estimation of cross section for neutral Ne:
Sulfur 16S8+ no Titanium foil
S(2.7MeV/u) = 2.5  10-17 cm2
dE/dx  q2 scaling with charge state q = 8 (Bethe-Bloch scaling)
 Proton p(2.7MeV) = S(2.7MeV/u) / q2 = 3.9  10-19 cm2
Energy loss from 2.7 MeV to 7 TeV by factor 0.03, but non Bethe-Bloch scaling for neutrals
 Bethe-Bloch scaling p(7TeV) = 0.03  p(2.7MeV) = 1.2  10-20 cm2
 Bethe-Born scaling p(7TeV) = 0.007  p(2.7MeV) = 2.8  10-21 cm2
Serban’s estimation: p(7TeV) = 4.710-22 cm2  factor 6 too large (25 for B-Bloch)
i.e. sec e- excitation
foil

## Slide 7
Beam: S at 3 MeV/u at TU-München TANDEM
100mm
10-2 mbar
rmfp~30 mm
10-1 mbar
rmfp~ 3 mm
30mm
10+1 mbar
rmfp~ 30 m
Pressure dependent Profile Reading
F. Becker et al. (GSI) , IPAC’12 and HB’12
Task: To which pressure the methods delivers
a correct profile reproduction?
Process:
1. Direct excitation e.g.: N2 + ion  (N2+)*+ e- + ion
2. Indirect processes e.g.: N2 + e-  (N2)*+ e-  N2 +  + e-
 No all spectral lines are correct
Results: Possible misreading for neutral lines
avoid 10-2 mbar < p < 10 mbar 
chose either rmfp >> rbeam or rmfp<< rbeam
use transition of the charged specious
all transitions
okay p< 10-3mbar
effect for Ne

## Slide 8
GSI UNILAC:
Ar10+@4.7 MeV/u,
2.5 mA
Distortions by Depth-of-focus, Space Charge and Curtain Thickness
P. Forck, A. Bank (GSI), EPAC 2002
Task: The depth-of-focus must fit to the beam size
Method: Determination of beam size
Care: Radial beam density distribution
Additionally: Vignetting  increase for open iris
Parameter:
Gas with lifetime =60 ns
Curtain: 1 mm thick curtain
Electron beam 5 A, mag. B=4 T
Hollow electron beam:
Start distribution
Space charge distortion
Curtain smearing
Smearing & distortion
Simulations by Ondrej Sedlacek (using GSI-code)
Smearing due to finite curtain size:
Thickness smaller than required resolution
Distortion due to space charge:
For N2 gas (fluorescence from ) not for Ne
Care: Non-uniform electron-beam distribution
For first prove-of-principle:
Thick curtain for large signal strength
O. Sedlacek et al, IPAC2021

## Slide 9
Image from 1·109 U
p= 2·10-3 mbar,
mounted ≈ 2 m
before beam-dump:
Ekin dependence for signal
& background close to beam-dump:
60 MeV/u
350 MeV/u
750 MeV/u
viewport
Background prop. Ekin2  shielding required
Background outside of viewport  not optical photons
Signal proportional to energy loss
Suited for single pass with ≥ 1010 ions/pulse
F. Becker (GSI) et al., Proc. DIPAC’07
Energy Scaling for N2 Fluorescence and Background

## Slide 10
Neutron induced Background
Neutron creation by nuclear interactions
Refection at concrete wall, i.e. ‘bouncing around’
Long ‘detention’ in tunnel; here  1 ms
Mitigation:
Image intensifier gate only during beam
However, at LHC not applicable
O. Sedlacek et al, IPAC2021
beam
Beam
dump
Concrete wall
Neutron flux by PHITS:
Beam: Xe48+ @200MeV/u stopped in dump
PHITS:
Particle and Heavy Ion Transport code System
(comparable to FLUKA)
Comparison PHITS & plastic scintillator:
(II comparable background decay):
Beam: Xe48+ @200MeV/u stopped in dump
Only relevant at LHC; no neutrons at HEL

## Slide 11
Fast Proton induced Background (LHC only)
Proton creation by nuclear spallation reaction:
Fast proton emitted in beam direction
Image Intensifier window with photo cathode
Cherencov radiation in glass
Sometimes ‘comets’ appeared
for parallel orientation of Image Intensifier
Mitigation:
Shielding by lead between vacuum window
Image Intensifier 900 rotation of II plane
beam
Beam
dump
Concrete wall
Neutron flux by PHITS:
Beam: Xe48+ @200MeV/u stopped in dump
100µm
stainless steel
vacuum window
Parallel II orientation
 ‘comets’
900 II orientation
 No ‘comets’
‘Comet’ due to fast proton
passage inside glass
with photo cathode
Only relevant at LHC; no fast protons at HEL

## Slide 12
Some experimental ‘Lessons learned’
Image Intensifier gate:
Voltage photo cathode  MCP follows gate TTL-gate,
input  1 M impedance
Maximum trigger rate 10 kHz
For long cable > 20 m (at LHC):
External 50  termination via T-piece required & 50  TTL driver electronics
Test: Background contribution must scale with gate length
Image Intensifier operation:
Avoid high illumination
Amplification by MCP voltage: Start with low value increase only to required voltage
Background:
Optical photons though viewport ( blackening methods detailed investigated by CERN)
Homogeneous distributed background by ionizing radiation (i.e. also outside of viewport)
Image Intensifier gate and camera exposure
II gate only if beam is present (HEL is pulsed, LHC quasi cw operation)
Camera exposure only slightly longer than II gate

## Slide 13
Conclusion
Working gas:
N2 and Ne properties fairly well investigated (more light from N2)
experimental knowledge for low electron currents at CI
Scaling with one order of magnitude for LHC beams
Proposal: If possible, start with N2 at HEL and large skimmer (‘first we have to see a signal…’)
Image intensifier, optics and camera:
Robust and well designed system, remote controlled at LHC, accessible at HEL test-bench
Proposal: Image Intensifier gate and camera exposure only during beam presence
LHC test operation:
Background measurement urgently required, dependence on optical filters
HEL test bench:
High beam current, high magnetic field, ‘exotic’ beam shape  surprises possible
High magnetic field and space charge:  surprises due to captured residual gas electrons or ion
Thank you for your attention!
Questions and comments?

## Slide 14
Backup slide

## Slide 15
BIF-Monitor: N2 Scaling for different Ions
C. Andre (GSI) et al., Proc. IBIC 2014
Scaling for different ions Aq+ beams:
GSI-UNILAC: 5 to 11 MeV/u
Care: recording within several years
 Systematic errors possible
Basically
Experiment: At GSI LINAC with 11.4 MeV/u
scaling confirmed within a factor of 3
 Bethe-Bloch scaling is reasonable
For N2 working gas the spectra for different ion impact is measured:
Ideally: constant
Result: Variation by factor 3

## Slide 16
UNILAC (4,5 MeV/u)
SIS-18 (200 MeV/u)
Pressure -Variation by 6 Orders of Magnituden
Notes: Original images show the signal increase
-> Energy variation – about the conversion factor

## Slide 17
gas | Y for p | Y for p/ne
Xe | 86 % | 22 %
Kr | 63 % | 25 %
Ar | 38 % | 30 %
He | 4 % | 26 %
N2 | 100 % | 100 %
Relative fluorescence yield Y (all wavelength):
ne: gas electron density  energy loss  beam influence
BIF-Monitor: Spectroscopy & Fluorescence Yield
Beam: S6+ at 5.16 MeV/u, pN2 =10-3 mbar
Results of detailed investigations:
Rare gases and N2: green to near-UV
Compact wavelength interval for N2
Fluorescence yield: N2  4x higher as rare gases
 N2 and Xe are well suited !
Photo-cathode:
low noise bialkali
F. Becker et al. (GSI, TU-München), IPAC’12 and HB’12

## Slide 18
Normalized profile reading for all :
Profile reading equal
for all gases except He
BIF-Monitor: Spectroscopy & Profile Reading
Results of detailed investigations:
Rare gases and N2: green to near-UV
Compact wavelength interval for N2
Fluorescence yield: N2  4x higher as rare gases
Same profile reading for all gas except He
 N2 and Xe are well suited !
Beam: S6+ at 5.16 MeV/u, pN2 =10-3 mbar
F. Becker et al. (GSI, TU-München), IPAC’12 and HB’12
Photo-cathode:
low noise bialkali