## Slide 1: CI Update
10/03/2022

## Slide 2: NEG property

## Slide 3: Saturation Estimates
Where
Skimmer | NEG Chamber |  | Saturation time
0.7x9mm | Interaction | 9.45E-09 | 7.7 days
0.3x9mm | Interaction | 3.50E-09 | 20.8 days
0.7x9mm | Dump | 9.50E-08 | 18.4 hours
0.3x9mm | Dump | 4.36E-08 | 40.0 hours
Skimmer | NEG Chamber |  | Saturation time
0.7x9mm | Interaction | 9.45E-09 | 7.7 days
0.3x9mm | Interaction | 3.50E-09 | 20.8 days
0.7x9mm | Dump | 9.50E-08 | 18.4 hours
0.3x9mm | Dump | 4.36E-08 | 40.0 hours
Sorption capacity = 2.26mbar L
Pumping speed
= 360 L/s
All data recorded here with a 5bar N2 Jet

## Slide 4: NEG Explanation
Sorbed Quantity / mbar.L PV = nRT PV = mbar.L
Pump lifetime is extended due to not always pumping at 360L/s especially at very low pressure, pumping speed is a function of sorbed quantity.
Saturation defined as when entire NEG surface is covered and no longer pumps. We can define as when pressure is too high.

## Slide 5: Neg pump installed in the Dump Chamber

## Slide 6: Saturation Data (dump chamber)
Chart: Dump chamber pressure(mbar) 0.7x9mm
Chart: Dump chamber pressure(mbar) 0.3x9mm
NEG pumping at 360L/s
NEG pumping speed reduced
Reaching equilibrium with diffusion pumping
0.3x9mm Skimmer, NEG at Dump
0.7x9mm Skimmer, NEG at Dump

## Slide 7: Saturation Data (interaction chamber)
Chart: Interaction chamber pressure(mbar) 0.7x9mm
Chart: Interaction chamber pressure(mbar) 0.3x9mm
0.3x9mm Skimmer, NEG at Dump
0.7x9mm Skimmer, NEG at Dump

## Slide 8: Saturation Data (Skimmer 2 chambers)
Chart: Skimmer II chamber pressure(mbar) 0.7x9mm
Skimmer chamber 2 has relatively constant pressures
Chart: Skimmer II chamber pressure(mbar) 0.3x9mm
0.3x9mm Skimmer, NEG at Dump
0.7x9mm Skimmer, NEG at Dump

## Slide 9: Saturation Data (Skimmer 1 Chamber)
Chart: Skimmer I chamber pressure(mbar) 0.7x9mm
Chart: Skimmer I chamber pressure(mbar) 0.3x9mm
Skimmer chamber 1 has relatively constant pressures with some increased noise towards the end of testing
0.3x9mm Skimmer, NEG at Dump
0.7x9mm Skimmer, NEG at Dump

## Slide 10: Saturation Data (Nozzle chamber)
Pressure in nozzle chamber drops over extended duration (as seen in previous test)
Chart: Nozzle chamber pressure(mbar) 0.7x9mm
Chart: Nozzle chamber pressure(mbar) 0.3x9mm
0.3x9mm Skimmer, NEG at Dump
0.7x9mm Skimmer, NEG at Dump

## Slide 11: NEG Re-installed in Interaction Chamber
Next results with 0.7x9mm skimmer

## Slide 12: Saturation Data (Dump Chamber)
Chart: Dump Chamber pressure(mbar)
Gate Valve open & degassing
Gas Jet turned on
0.7x9mm Skimmer, NEG at interaction chamber

## Slide 13: Saturation Data (Interaction Chamber)
Chart: Interaction Chamber pressure(mbar)
TMP1 controller overheat (extra peak)
Gate Valve open & degassing
TMP1 controller overheated due to airflow blockage
0.7x9mm Skimmer, NEG at interaction chamber

## Slide 14: Saturation Data (Skimmer 1&2 Chamber)
Chart: 
Chart: 
TMP controller overheat
0.7x9mm Skimmer, NEG at interaction chamber
*Note: the overheat is due to the controller is covered with the lab book not the pump itself.

## Slide 15: Saturation Data (Nozzle Chamber)
Chart: 
TMP1 controller overheat
Chart: 
Non-log scale to show pressure reducing after TMP controller cooled down
0.7x9mm Skimmer, NEG at interaction chamber

## Slide 16: 0.7*9 mm skimmer comparison (interaction chamber)
Chart: Interaction chamber pressure(mbar) 0.7x9mm
Chart: Interaction Chamber pressure(mbar)
NEG at Dump
NEG at Interaction chamber

## Slide 17: 0.7*9 mm skimmer comparison (dump chamber)
Chart: Dump Chamber pressure(mbar)
NEG at Dump
NEG at Interaction chamber
Chart: Dump chamber pressure(mbar) 0.7x9mm

## Slide 18: Shipping
We will pack on Tuesday (March 15th), ship on March 16th, Hopefully it will arrive at Friday (March 18th).

## Slide 19: Summary
Week long data shows the NEG is not saturated for 0.3*9 mm 3rd skimmer (5bar inlet pressure).
Future work:
  Preparation for shipping.