## Slide 1: BPM Processing for the HiLumi LHC Interaction Region BPMs
Douglas BETT, SY-BI-BP
BI Student and R&D Meeting #3, 18 March 2022
Michal Krupa
Manfred Wendt
Andrea Boccardi
Manuel Barros Marin
Irene Degl’Innocenti
Selim Ozdogan
Joël Daricou
Phil Burrows

## Slide 2: Contents
Introduction to the HL-LHC IR BPM upgrade
Simulation framework for BPM processing
Power compensation

## Slide 3: BPM upgrade
New stripline BPMs in the immediate vicinity of interaction regions 1 (ATLAS) and 5 (CMS), where the two beams coexist within a single pipe
Signals from all eight ports required to calculate the horizontal and vertical position of both beams
The distortion of the beam signals due to the presence of the other beam must be compensated for
type-A
type-B
BPM | Type | s [m] | bxt [ns]
Q1 | A | 21.853 | 3.92
Q2A | B | 33.073 | 3.92
Q2B | B | 43.858 | 6.82
Q3 | B | 54.643 | 9.72
CP | B | 65.743 | 10.52
D1 | B | 73.697 | 7.36
bxt = bunch crossing time i.e. beam-beam delay
A

## Slide 4: Zynq RFSoC: Candidate digital processor
Processing System Quad Arm Cortex-A53
Dual Arm Cortex-R5
Programmable Logic
RF-DAC
RF Out
DUC
RF-ADC
RF in
DDC
RF-ADC
RF-DAC
GTY Transceiver
Zynq RFSoC
Xilinx Zynq RFSoC (= radio frequency system on a chip)
Model ZU28DR (Zynq UltraScale architecture)
ZCU111 Evaluation Board

## Slide 5: Zynq RFSoC: Candidate digital processor
Xilinx Zynq RFSoC (= radio frequency system on a chip)
Model ZU28DR (Zynq UltraScale architecture)
Processing System Quad Arm Cortex-A53
Dual Arm Cortex-R5
Programmable Logic
RF-DAC
RF Out
DUC
RF-ADC
RF in
DDC
RF-ADC
RF-DAC
GTY Transceiver
Zynq RFSoC
ZCU111 Evaluation Board
Processing System Features | 
Application Processing Unit | Quad-core Arm Cortex-A53 up to 1.33 GHz
Real-Time Processing Unit | Dual-core Arm Cortex R5 up to 533 MHz
Programmable Logic | 
System Logic Cells (K) | 930
DSP Slices | 4,272
Memory (Mb) | 60.5
33G Transceivers | 16
Maximum I/O Pins | 347
Direct-RF Signal Chain Features |  | 
Max. RF input Frequency (GHz) |  | 4
Decimation / Interpolation |  | 1x, 2x, 4x, 8x
12-bit RF-ADC | # of ADCs | 8
 | Max Rate (GSPS) | 4.096
14-bit RF DAC | # of DACs | 8
 | Max Rate (GSPS) | 6.554
Asynchronous clocking
~1.5 counts of noise

## Slide 6: BPM signals
IR1 Q1 presqueeze orbit
Both beams close to center in horizontal
~6 mm off in vertical
First beam (2.3e10)
Second beam (2.3e11)
Gaussian (RMS 0.95 ns)
Bunch crossing timing 3.92 ns

## Slide 7: Signals from first beam
Port 1 (T)
Port 3 (L)
Port 5 (B)
Port 7 (R)
First beam “coupled”
signals

## Slide 8: Signals from first beam
Port 2 (T)
Port 4 (L)
Port 6 (B)
Port 8 (R)
First beam “isolated”
signals
Signals small at downstream end due to stripline directivity

## Slide 9: Signals from second beam
Port 1 (T)
Port 3 (L)
Port 5 (B)
Port 7 (R)
Second beam “isolated”
signals

## Slide 10: Signals from second beam
Port 2 (T)
Port 4 (L)
Port 6 (B)
Port 8 (R)
Second beam “coupled”
signals
Signals arrive later and are larger due to higher intensity

## Slide 11: Two beam signals
Port 1 (T)
Port 3 (L)
Port 5 (B)
Port 7 (R)
Signals from second beam visibly distort signals used to calculate position of first beam

## Slide 12: Two beam signals
Port 2 (T)
Port 4 (L)
Port 6 (B)
Port 8 (R)
Signals from first beam difficult to discern due to much lower intensity

## Slide 13: Simulation framework: analogue front-end
beam
coupled
isolated x20
beam on axis, intensity 1e11
Ideal induced signal from CST

## Slide 14: Simulation framework: analogue front-end
long cable
Stripline BPM
LCF12-50JFN
Length = 100 m
Velocity = 0.88c
Attenuation = 7.19 dB/100 m @ 1 GHz
Impulse response of coaxial cable
Constructed signals
Cable in
Cable out

## Slide 15: Simulation framework: analogue front-end
long cable
Low-pass filter
Stripline BPM
“absorptive” filter (low reflections)
Constructed signals post-filter

## Slide 16: Simulation framework: analogue front-end
long cables
Low-pass filter
variable attenuator
Digitizer
Stripline BPM
12-bits @ 4.096 GSPS
2 dB step size
Asynchronous clock

## Slide 17: Simulation framework: digital processing
Digitize waveforms
Compensate for other beam
Integrate channels
Calculate position
Correct non-linearity
Calibrate, account for BPM roll
algorithm
Polynomial adjustment (2D)
e.g. root of sum of squares (rss)
Standard BPM processing operations

## Slide 18: Digitized waveforms
Stable orbit assumed
Position calculated from of waveform power
Shot-by-shot resolution ~10 µm
Channel 1 @ 4 GHz
12-bit, ±0.5 V
2,760 acq.
Systematic error on mean position: ~10 µm horizontal [1.5%]
~200 µm vertical [3%]
Unwanted signal from counter beam

## Slide 19: Power compensation I
+
=
: signal induced at upstream end by reference on-axis first beam
: signal induced at upstream end by reference on-axis second beam
scale factor sets amplitude of each term according to intensity, position of beam. Higher intensity or closer to stripline  larger
Each waveform can be expressed in the form:

## Slide 20: Power compensation II
Taking the sum of the samples squared for both sides:
Can be solved for
By replacing with , the above expression becomes a quadratic equation whose coefficients are functions of scalars that are calculated in advance from the reference waveforms or calculated in real time from the measured waveforms:

## Slide 21: Power compensation III
By replacing with , the expression becomes a quadratic equation whose coefficients are functions of scalars that are calculated in advance from the reference waveforms or calculated in real time from the measured waveforms:

## Slide 22: Power compensation
Mean offset
10.0 µm  0.22 µm
First beam, horizontal
First beam, vertical
Mean offset
201 µm  0.36 µm

## Slide 23: Measurement campaign
Gateware and firmware designed by Irene Degl’Innocenti to digitize 4 of 8 ADCs
Companion board produced by Selim Ozdogan
BPM 2L5
~200 ns
Actual two-beam waveforms

## Slide 24: CONCLUSION
New stripline BPMs for HiLumi interaction regions require new processors; direct digitization with RFSoC is ideal
Primary challenge for is accurately recovering position of each beam in presence of unwanted signal from the other
Power compensation scheme proposed and found to deliver best results for simulated beam waveforms
Future studies involving beam time and position scans will test performance of the power compensation method with realistic beam signals