## Slide 2
Compound Annual Growth Rate (CAGR)
High quality, low defect density SiC is available up to 200 mm wafers

## Slide 3: Many potential applications in sensors
Nuclear fusion reactors
-Plasma diagnostic
Aerospace
-Sensors and electronics
Medical
-Dosimetry in FLASH therapy and microdosimetry

## Slide 4: Advantage of using SiC
Wide bandgap that reduces the leakage current, maintaining low noise levels even at high temperatures.
Insensitive to visible light.
High atomic displacement threshold (~20 eV for C, and ~40 eV for Si), which should make the material more radiation resistant;
Fast saturated electron drift velocity (2x107 cm/s at room temperature), twice faster than silicon;
High thermal conductivity (490 Wm-1K-1), which is three times higher than that of Si, and ten times higher than that of GaAs.

## Slide 5: SiC detectors fabricated at the IMB-CNM clean room
Technology based on :
SiC rectifiers for space
  Bepi-Colombo & Solar orbiter missions
  Main contractors EADS, Airbus, ALTER
DOI: 10.1109/TDMR.2009.2029090

## Slide 6: Detector technology SiC
SiC based “P on N” diode fabricated on 4H-SiC on a 50um high resistivity n-type doped epi-layer. 100 or 150mm wafers.
The technology has been optimized with high temperature processers and metal contacts.
J.M. Rafí et al 2018 JINST 13 C01045

## Slide 7: Testing at different facilities
SiC measurements at high temperature at the 3 MeV Tandem.
TPA-TCT, with two photons producing one electron-hole pair, and the energy is deposited at the laser focal point, so this permits a 3D spatial resolution.
FLASH effect conditions* : Irradiations with con Ultra-High Dose Rate pulsed radiation reduce adverse effects in healthy tissues. Tested in PTB, Germany electron FLASH.
*Favaudon et al., Sci Transl Med 6 (2014)

## Slide 8: CCE homogeneity study
Homogeneity IBIC measurements on the microprobe line at CNA with He2+ @ 3.5 MeV.
The mean standard deviation is ~1%.
M.C. Jiménez-Ramos, J. García López , A. García Osuna
Jiménez-Ramos M.C et al., Radiation Physics and Chemistry 177 (2020) 109100b

## Slide 9
High temperature setup
Vacuum chamber for high temperature measurements.
Furnace accommodated inside the chamber (RT-500 ºC)
TO257 package
Ag Sintering
Gold bondings
Detectors mounted on a special package certified for high temperature.
Tested with thermal cycles during various days.
Calibration with He2+ ions

## Slide 10
Alpha source measurement at RT for setting the operation voltage
244Cm emisión ~ 5,8 MeV
Range ~ 19 mm
At 40 V the depletion zone is thick enough to completely stop the 5.8 MeV alpha particles and the CCE signal saturate.
Both electronic noise and reverse current increase at higher voltages during high temperature measurements, so it was decided to perform the experiments at 40 V.

## Slide 11: Measurements done at : TAMARA Temperature Analysis Measurements And Radiation Applications
IBIC using 3.5 MeV alphas
Measurements from RT to 450ºC

## Slide 12: Pair creation energy vs T
The increase of energy collected vs T is excepted:
Pressure and Temperature Dependence of Energy Gap in SiC and Si 1-x Ge x.
DOI: 10.33899/rjs.2019.163296

## Slide 13: Two Photon Absorption TPA
Esteban Curras et al., Caracterization of neutron irradiated IMB-CNM SiC planar diodes with TPA-TCT, RD50 workshop 2022
I. Vila et al., “TPA-TCT: A novel transient current technique based on the two photon absorption (TPA) process,” in Proc. 25th RD Workshop. Geneve, Switzerland: CERN, 2014, pp. 17–19.
Wave length (800nm) adapted to the SiC bandgap for the TPA
Metal opening

## Slide 14: First measurements of SiC diodes with TPA
C. Quintana, I. Vila, R. Montero
Ref: http://dpnc.unige.ch/seminaire/talks/fernandez.pdfhttp://dpnc.unige.ch/seminaire/talks/fernandez.pdf
Rise time on the order of few hundreds of ps

## Slide 15: Depletion width vs bias: non irradiated
Neff=1.3 E14/cm2 in agreement with the nominal doping
Diode fully depleted between 300-500 volts
Homogeneity in the sensor depletion

## Slide 16: Z-scan charge profiles: non-irradiated diodes
depletion
Detector is 50um thick, nominal value
50um
300um
Z-scan

## Slide 17: Z-scan charge profiles: irradiated diodes
The diode behavior lost!
Capacitor-like charge collection
No charge collection saturation with bias voltage
Neutron-irradiated (ATI Vienna) July/Aug 2021
Fluences are in 1MeV neutron equivalent /cm2

## Slide 18: Depletion width vs bias: irradiated
Both figures show that the depletion width is constant for the irradiated detectors.
The depletion width is different if we compare irradiated and non-irradiated detectors, but also between the irradiated ones

## Slide 19: Charge collection efficiency
Dependence with irradiation:
Charge collection increases with bias
Charge collection decreases with irradiation.
Lost of charge due to signal trapping in the resistive electrode (metal opening for TPA).

## Slide 20: FLASH radiotherapy: dosimetry
FLASH effect (Favaudon et al., Sci Transl Med 6 (2014)): Irradiations with con Ultra-High Dose Rate pulsed radiation reduce adverse effects in healthy tissues.
Need real-time, highly precise dosimeters
CNM is a partner in the European project EMPIR-UHDPulse (2019-2023) for metrology in UHDR beams
C. Fleta and G. Pellegrini
FLASH: h(> 40Gy/s) while keeps same tumor control & "need real-time, highly precise dosimeters since standard dosimeters saturate"

## Slide 21: First tests of CNM’s SiC diodes in conventional RT dosimetry
Linearity deviation with 50 kV X-rays integrated dose delivered is less than 1.5%
50kV X-rays, 0 V bias
(~ 18kGy max. dose @ 3.5h)
Medium-term stability with integrated dose up to 18 kGy X-rays is better than 0.3%
F. Gomez USC

## Slide 22: Electron beam
Linearity deviation with dose per pulse of 9 MeV electrons up to 2.5 mGy is less than 1.5%
9 MeV electrons, 40 V bias.
Reference dosimeter: PPC40 @ -400 V
SiC diodes fabricated at CNM show good performance as dosimeters in conventional RT beams where they are a promising alternative to other semiconductor materials.
Currently tests are being performed to evaluate their capabilities in Ultra-High Dose Rate beams.
F. Gomez USC

## Slide 23: 3D in SiC material
diameter :8mm.
depth :300mm.
SiC
Reduce electrode distance keeping a thick substrate
New Deep RIE for Si and SiC materials will be available soon in our Clean Room
Laser drilling seems a possible solution
Possibility to fabricate 3D detectors or other MEMS structure
G. Pellegrini, “Technology development of 3D detectors for high energy,” PhD Thesis, University of Glasgow, 2003.
S. Nida et al., “Silicon carbide X-ray beam position monitors for synchrotron applications,” Journal of Synchrotron Radiation, vol. 26, no. 1, pp. 28-35, 2019.
Laser drilling [1]

## Slide 25: Future work
Fusion :
Test irradiated detectors at high temperature
Understand the use of Graphene as a contact (TPA, high T, etc…)
Explore other application related to fusion (IFMIF –DONES)
Test sensors in fusion facilities (DIII-D San Diego, CA,).
Medical Flash:
Study the limit of SiC dosimetry in FLESH therapy (in collaboration with PTW)
Fabricate new sensors for micro-dosimetry to measure LET (National project submitted, NEWDOSI project started in 2022 and UHDR EU project TWAC for development of compact electron accelerators
Large strip SiC sensors and LGAD for beam monitoring (U. if Vienna and MedAustron)
HEP:
Test beam of irradiated detectors : strip configuration.
Understand the annealing at high temperature.
Fabricate SiC detectors with internal gain (LGAD)
Exploratory detector concepts based on 2D materials, e.g. grapheme

## Slide 29: Irradiation campaign
2 MeV electron irradiation (e/cm2) |  |  |  | 
1x1014 | 1x1015 | 1x1016 |  | 
Neutron irradiation (n/cm2) |  |  |  | 
5x1013 | 1x1014 | 5x1014 | 1x1015 | 2x1015
24 GeV/c proton irradiation (p/cm2) |  |  |  | 
8.6x1013 | 1.5x1014 | 1x1015 | 1.7x1015 | 2.5x1015
2 MeV electron irradiations were carried out at Takasaki-JAEA electron accelerator in Takasaki, Japan.
Neutron irradiations were performed at JSI research reactor in Ljubljana, Slovenia
4 GeV/c proton irradiations were carried out at CERN PS-IRRAD facility in Geneva, Switzerland.
This project has received partial funding from the European Union’s Horizon 2020 Research and Innovation programme under Grant Agreement no. 654168.

## Slide 30: Electrical characterization after irradiation
J.M. Rafí et al 2018 JINST 13 C01045

## Slide 31: CCE characterization
Detectors tested with 5MeV alpha particles and at room temperature
J.M. Rafí et al IEEE TNS, vol. 67, no. 12, december 2020

## Slide 32: Ecosystem
Materials Science Institute (CSIC)
IMB-CNM(CSIC)
Alba Synchrotron
ICN2
PRUAB Spin-off Incubator
Sciences (UAB)

## Slide 33
Clean Room
1,500 m2
total area
190
equipment units
3000
Wafers/year
40
self service
40
staff
2500
Hours self service
450
registered self service licenses
550
Runs/year
Largest R&D Clean Room in Spain

## Slide 34: Integrated Micro & Nano Fabrication Clean Room
Clean room
class 100 to 10.000
Technologies for Micro y Nano fabrication
3 Areas:
  CMOS (high purity, no contaminants)
  MNC (Noble Metals contaminants for Si)
  Nanolithography
Process size: 4'' and 6” wafers
Technologies: CMOS, MEMS/NEMS, power devices, radiation detectors
Silicon Micromachining
Microelectronic Packaging Area
200 m2, class 100
Wafer cutting, wire-bonding, bump-bonding
The main activities of IMB-CNM is basic and applied research and development, education and training in micro and nanotechnologies, components and systems.

## Slide 35: Clean Room Equipment
  Thermal processes, CVD and ALD
  Ion Implantation
  PVD and Metallisation (Sputtering and Evaporators)
  Optical Lithography:
    Proximity Aligners: single and double side
    Steppers: g-line and i-line
    Direct laser writing
  Nano-lithography (e-beam, NIL, FIB and AFM)
  Dry etching
  Wet and dry micromachining
  Wet etching and cleaning
  On-line test
  Conventional and advanced packaging
  Electrical characterization
Notes: 35

## Slide 37: Experience in production
SiC rectifiers for space
  Bepi-Colombo & Solar orbiter missions
  Main contractors EADS, Airbus, ALTER
  400k€/year (2,800 k€ accumulated)
Beam monitors for Synchrotron accelerators
  position and intensity thin sensors (Alba cell, Alibava)
Radiation detectors
  Main contractor CERN (ATLAS experiment, RD50, TOTEM
  Other contractors: INFN, BNL, DESY, GSI, Xian University, .
Four Quadrant diodes for space
  Main contractor non disclosed
  In orbit since February 2019, full contract 650 k€

## Slide 38
Change in the BandGap energy
Measurements from RT to 450ºC
https://www.iue.tuwien.ac.at/phd/ayalew/node61.html
Properties of Advanced Semiconductor Materials: GaN, AIN, InN, BN, SiC, SiGe.
John Wiley & Sons. (2001).
Change in the pair creation energy
Change in measured energy
THE INCREASE OF CHARGE COLLECTED VS T IS EXPECTED