# Review meeting 19/04/2022

- Talking Points

  - Minimum thickness of the vacuum chamber to avoid leak tightness

## Action Items

| Action | Responsible |
| --- | --- |
| - Ask to the vacuum group about the bake out temperature | (G & T) |
| - Find the alloy that better suits our requirements | (T) |
| - Check in ANSYS asymmetries (variable thickness) and misalignments | (T) |
| - Check grain size with ultrasonic technology after buying the material to machine | (G & T) |
| - Perform a cleaning operation after the manufacturing and ensure a clean storage before installation | (G & T) |

- Notepad

- Multidirectional forged aluminium alloy

- Aluminium alloy resistant to pitting corrosion to avoid leak tightness

- Aluminium alloy with good mechanical properties (tensile strength and yield strength) after baking out at 260 ºC during 100h (check with the vacuum group the baking out temperature)

- Avoid chlorides or fluorides in the aluminium alloy

- Fine grain: ten grades over the thickness (10 μm). Check the size with ultrasonic technology

- Apply a safety factor of 2 between standard and engineering properties

- Check in ANSYS asymmetries (variable thickness) and misalignments

- After manufacturing and before installation perform a cleaning process

2219 (Standard properties)

5083H0 (Standard properties)