## Slide 1
Passive Piezo-resistive Gauges
for Insulation Vacuum

## Slide 2
Overview
The different cases where piezo-resistive gauges are used
Low-radiation areas
High-radiation areas
Current gauge in high-radiation areas
Alternative passive gauges and their tests so far
Summary

## Slide 3
Piezo resistive cases based on LHC location
LSS
DS
ARC
High-radiation areas
(1MGy HL lifetime)
Only passive gauges
Low-radiation areas
(500Gy HL lifetime)
Active RadTol gauges
Balzers APR017
x65 pcs
HUBA 680.99413
x168 pcs

## Slide 4
Low-radiation areas

## Slide 5: Huba 680.99413
Pressure type | Absolute
Measurement range | 0 - 1.6 bar
Overpressure | 5 bar
Process connection | KF16 DIN 28403
Electrical connection | DIN EN 175301-803- A
Output signal | 0 – 10 V
Bridge excitation | ± 13.5 VDC
Linearity | < ± 0.25 % FS
Temperature range | 0 - 70°C
Seals | ??
168 already installed in low radiation areas
  LHC DS + ARC for IV measurement
Provides 0-10V for 0-1600mbar
  Requires +/-13.5VDC power supplies
Required radiation tolerance of 500Gy
  2.5x higher than expected HL-LHC level in the DS
  10x higher than expected HL-LHC level in the ARC
Only one active component inside
  Differential amplifier MC33178
Tested at:
  CC60 up to 2.4kGy
  CHARM up to 500Gy

## Slide 6: Huba 680.99413 irradiation at CC60

## Slide 7: Huba 680.99413 irradiation at CC60
Successfully accumulated 2.4kGy with gamma
<2.5% error at 2.4kGy
<1.5% of error at 500Gy
Active Piezo-Resistive Gauges in the gamma field of CC60 EDMS 2149211
Positive rail
Negative rail

## Slide 8: Huba 680.99413 irradiation at CHARM

## Slide 9: Huba 680.99413 irradiation at CHARM
Successfully accumulated 315Gy with mixed fields
<2.2% of error at 315Gy
<3.5% of error at 500Gy (extrapolation)
Active Piezo-Resistive Gauges in the Mixed Filed of CHARM EDMS 2149211
Positive rail
Negative rail

## Slide 10
High-radiation areas

## Slide 11
65 already installed in high radiation areas
  LHC LSS for IV measurement
Provides 0-600 mV for 0-2000 mbar
  Requires +4mA current excitation
Required radiation tolerance of max 1MGy
  Never tested (not to our knowledge)
Not manufactured any more:
  Spares for LHC operation
  HL-LHC needs (32 units in addition)
Current gauge in LSS – Balzers APR017

## Slide 12: STS TM series to replace Balzers APR017
Pressure type | Absolute
Measurement range | 0-2 bar
Overpressure | 6 bar
Process connection | KF16 DIN 28403
Electrical connection | Binder 723, 5-Pts
Output signal | 0 – 50 mV
Bridge excitation | 10V DC / 1mA DC
Accuracy | <±0.5 % FS
Temperature range | 0 - 70°C
Seals | EPDM

## Slide 13: Oil surrounding the STS TM series
Proposal for radiation resistant oil
Tested by ITER (awaiting for the results)
Increase of unit price by a factor 10!
Unit price 3500 CHF
https://www.stssensors.com/blog/2020/06/30/electronic-pressure-measurement-comparison-of-common-measuring-principles/
Standard oil
Unit price 350 CHF

## Slide 15: Irradiation at CC60
Potential radiation sensitive elements
  Seals (EPDM)
  Piezo-resistive element
    silicon diaphragm
  Oil surrounding the piezo-resistive element
    silicone-based oil
Active CC60 irradiation
  Pre/Post-irradiation characterization
  Pre/Post-irradiation leak test
  Active irradiation at atmospheric pressure
  1 reference gauge, 3 irradiated samples
2 references tested
  Voltage bridge excitation (STS TM-132960)
    up to 90kGy
  Current bridge excitation (STS TM-133427)
    up to 60kGy

## Slide 16: STS TM-132960 (Voltage excitation): Leak detection
Gauges | Pre-irradiation [mbar.l/s] | Post-irradiation [mbar.l/s] | Delta %
V - Reference | 3.2E-8 | 5.4E-8 | +69%
V - Sample 1 | 2.1E-8 | 5.1E-8 | +143%
V - Sample 2 | 1.0E-8 | 1.6E-7 | +1500%
V - Sample 3 | 3.0E-8 | 1.3E-7 | +333%

## Slide 17: Leak detection clarification

## Slide 18: STS TM-132960 (Voltage excitation): Characterization

## Slide 19: STS TM-133427 (Current excitation): Leak detection
Gauges | Pre-irradiation [mbar.l/s] | Post-irradiation [mbar.l/s]
C - Reference | 1.3E-7 | ?
C - Sample 1 | 9.2E-8 | ?
C - Sample 2 | 9.4E-8 | ?
C - Sample 3 | 1.2E-7 | ?

## Slide 20: STS TM-133427 (Current excitation): Characterization

## Slide 21: Irradiation at external facility – RADTEST company
Passive Gamma irradiation
  Pre/Post-irradiation characterization
  Pre/Post-irradiation leak test
  8 samples sent in total
    2 samples at 50kGy
    2 samples at 100kGy
    2 samples at 500kGy
    2 samples at 1MGy
Awaiting the samples to arrive at CERN
Test specifications EDMS 2621641

## Slide 22
Summary
STS gauges seem to be a good candidate to replace APR017
  Preferred solution is with the current excitation (long cables, integration w/ existing electronics)
  Noticeable relative error already from 60kGy
  CC60 campaign data require further analysis (offset, gain linearity, error linearity wrt TID)
  Based on CC60 sample analysis some mitigation actions can be envisaged to maintain errors in tolerable margins
  Awaiting for post radiation leak rate tests of CC60 samples (Q2 2022)
  Awaiting for the post radiation tests from the samples irradiated externally (Q3 2022?)
Not many other options apart from STS
  Pfeiffer APR262 active gauge by removing the internal electronics