## Slide 1: Carbon gantry project(s)
Marco Pullia

## Slide 2: Comparison of the depth dose profiles
Protons/Carbon
Electrons
Photons
Tumour
Beam

## Slide 3: Treatment delivery
(courtesy of Siemens Medical)

## Slide 4: What is a gantry
A gantry is a section of beamline that can rotate around the isocenter in order to direct the beam onto the patient from any direction

## Slide 5: Why a gantry?
To treat patients in supine position (eventually prone) in the same position in which CT, PET and MRI were acquired. Patient rotation only around gravity to preserve internal organs and soft tissue geometry.
To provide the maximum flexibility in selecting the irradiation direction when optimising the dose delivery.
To allow a “robust” treatment planning. Exploiting the sharp distal fall off can be risky in some cases and a gantry helps in avoiding fields directed towards an Organ At Risk (OAR).
Avoid density heterogeneities
Minimize SOBP extension (less energies required and better peak to plateau ratio)

## Slide 6: Why a gantry
Allows better, more robust planning:
e.g. minimize fields pointing towards OAR (Organ At Risk)
O.A.R.
With gantry
With horizontal line only

## Slide 7: Gantry in conventional radiotherapy
The whole linac is inside the gantry
The gantry head can pass between patient and floor for irradiation from below
(Varian Clinac IX)
2.6 m
3.5 m
Notes: 4-25 MeV
R ~ 1.3 m
Space around isocenter ~ 43 cm
The gantry is small!

## Slide 8: Size and magnetic rigidity
Conventional RT
Proton Gantry
Br < 2.4 Tm
Carbon Ion Gantry
Br < 6.4 Tm
Deep pit under the patient
Notes: Deep pit under the patient

## Slide 9: Proton gantry
Mitsubishi
Hitachi
IBA

## Slide 10: HIT
HIT C gantry: L = 25 m x f = 13 m, 600 t
(Udo Weinrich, GSI)
360° rotation
Parallel scanning
200 mm x 200 mm field
140 t magnets
120 t shielding-counterweight
600 t total rotating mass
Very large, very heavy,
very expensive

## Slide 11: Carbon Ion Gantry
The HIT Gantry
L = 25 m x f = 13 m,
600 t rotating mass

## Slide 12: HIMAC Superconducting rotating-gantry
Weight: order of 300 tons
Ion kind : 12C
Irradiation method: 3D Scanning
Beam energy : 430 MeV/n
Maximum range : 30 cm in water
Scan size : 200×200 mm2
Beam orbit radius : 5.45 m
Length : 13 m
The size and weight are considerably reduced
Use of superconducting (SC) magnets
Courtesy of Y. Iwata

## Slide 13: Layout of the SC gantry
Combined function SC magnets
（BM01～BM06）
→No quadrupole magnet required
Scanning magnets on top
→Large scan size
Combined function SC magnets
→Square irradiation field
→Parallel beam
Courtesy of Y. Iwata

## Slide 14
NIRS superconducting rotating gantry
HIT
HIMAC

## Slide 15: Yamagata
First treatment a few days ago

## Slide 16: Launch of a gantry design
CNAO collaboration with CERN – INFN – MedAustron
Design to be ready in 3 (4) years (when present expansion will be completed)
… find the money, design building and plants, build everything… to be ready in 8-10 years?
Advisor committee selected

## Slide 17: The starting point
Toroidal gantry operating in steady-state with a discrete number of irradiation directions
Rotating gantry attached to the wall with no counterweight and a partial rotation of ±110 degrees.
U. Amaldi et al, SIGRUM gantry, CERN-NIMMS-Note-002: https://cds.cern.ch/record/2766876
L. Bottura et al., GaToroid - Status and Perspectives,
CERN-NIMMS-Note-003: https://cds.cern.ch/record/2776464
GaToroid
SIGRUM

## Slide 18
Advisor committee, Dec 16th 2020: SIGRUM type gantry with recommendations:
Many alternative optics studies
Many alternative mechanical studies
Evaluation of the feasibility of SC dipoles for many different requirements
(embedded gradient, fixed vs variable gradient, field and gradient
level, ramp rate, aperture, technology, cooling, number of feedthroughs
and so on)
Study of the space available at CNAO for the gantry/layout linee
Gantry for carbon ions
revise specifications (180 vs 360, dipole specs, 3T vs 4T, scanning system, scan with quads, upstream scanning and so on)

## Slide 19: Many synergic projects
Sigrum
(CNAO - CERN – INFN – MedAustron) Signed 31 March 22
Hitriplus Started April 2021
Ifast Started May 2021