## Slide 1: September 2016-October 2018 Supporting the new initiatives of CNAO and AVO-ADAM
A. Garonna on behalf of the TERA technical team at CERN
30th anniversary of the TERA Foundation
15 September 2022
Notes: Disclaimer: please excuse the omissions (PHILA, PERLA, Centro Radiofarmacia, BISE2, etc) due to shortness of presentation and intent not to duplicate content presented by others

## Slide 2: What brought me to TERA? 2007-2011

## Slide 3: What brought me to TERA? 2007-2011

## Slide 4: The technical team in 2016-2018 (1/2)

## Slide 5: The technical team in 2016-2018 (1/2)
Left to right: G. Molinari, U. Amaldi, C. Cuccagna, P. Carrio Perez, V. Bencini, myself, E. Felcini

## Slide 6: The technical team in 2016-2018 (2/2)
Left to right: N. Alharbi, S. Allegretti, P. Pearce
Left to right: M. Varasteh, P. Riboni, M. Vaziri,
E. Benedetto, D. Arlunno, V. Ljubicic, R. Maccaferri, D. Bergesio

## Slide 7: CNAO Expansion

## Slide 8: CNAO Expansion

## Slide 9: Area A
Two single-room facilities
Comparison of technical specifications from VARIAN ProBeam, Mevion S250i, IBA ProteusONE, Protom Radiance330
Carbon Ion gantry
Comparison of technical specifications of the carbon ion gantries from HIT, NIRS, Riesenrad (design), FFAG (design) along with an analysis of the relevant innovations in field
For each vendor and C-gantry, a proposal of the new facility layout was made

## Slide 10: Area B
3´500 m2 dedicated to:
  Radiopharmacy (diagnostics & 11C) department based on a 18 MeV cyclotron
  Nuclear Medicine Diagnostics Facility
  Boron Neutron Capture Therapy facility (accelerator & beam shaping assembly designed by Pavia University & INFN)

## Slide 11: Layouts of underground and ground floors of Area B

## Slide 12: Status Update - 2022
One single-room from Hitachi
BNCT system from Tae Life Sciences

## Slide 13: FASST: the most beautiful gantry in the world
Notes: William of Ockham cites the principle of economy, Frustra fit per plura quod potest fieri per pauciora ("It is futile to do with more things that which can be done with fewer"; Thorburn, 1918, pp. 352–53; Kneale and Kneale, 1962, p. 243.)
    - Rotation max speed of 0.75 rpm
    - Max deviation of ±0.1° and deflection of optical axis ±0.5 mm

## Slide 14: FASST: the most beautiful gantry in the world
    Pierluigi Riboni´s mechanics concept based on Occam´s razor with:
      Structure made of 4 rods
      25 t rotating mass (no counterweight) supported on the wall
      Iso-centric gantry with 220° rotation
Notes: William of Ockham cites the principle of economy, Frustra fit per plura quod potest fieri per pauciora ("It is futile to do with more things that which can be done with fewer"; Thorburn, 1918, pp. 352–53; Kneale and Kneale, 1962, p. 243.)
    - Rotation max speed of 0.75 rpm
    - Max deviation of ±0.1° and deflection of optical axis ±0.5 mm

## Slide 15: Optics Design
    compact and light design, thanks to small transverse emittances of proton beam linac & downstream scanning
    Magnets & power supplies designed for beam energy pulse-to-pulse variation at 200 Hz:
    2% current variation in 4 ms, i.e 80% of longitudinal FWHM of Bragg peak
    - One optical solution for multiple spot sizes (all energies):
    -> Telescope with rotational invariance
CERN FeCo prototype (Tommasini et al.)
Notes: LIGHT proton beams are unique not only because of their rapid energy variations, but also
for their very small normalized emittances (Fig. 1.8): about 0.2 π mm mrad (normalized, 5-
rms), roughly ten times smaller than the emittances of the beam produced by synchrotrons
and cyclotrons. For example, the PIMMS synchrotron is characterised by an emittance of 2.6
π mm mrad (normalized, 5-rms) [PIMMS], while the standard C235 cyclotron has 19 π mm
mrad (normalized, 5-rms) [Walter2001].

## Slide 16: Fast Adaptive Spot Scanning Therapy
Notes: The reason behind this choice is that, by having two spot types, a tumour target can be “painted” twice: the central part with “large” spots and the boundaries with “small” spots so to obtain a steeper fall-off of the dose in a shorter time. This is another unique feature of the FASST gantry beam.

## Slide 17: Unique beam properties
Two spot sizes at isocenter (in air) to obtain steeper dose fall-off in short irradiation time
Constant spot size in air for most energies to facilitate treatment planning and QA
Nozzle geometry modeled with Monte Carlo to determine the scattering effect
Chart: E (MeV)
| 70 | 90 | 110 | 130 | 150 | 170 | 190 | 210 | 230
small spot | 15.2591 | 12.0304 | 10 | 8.5703 | 7.6322 | 7 | 7 | 7 | 7
larg spot | 15.2591 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14
Small spot
Large spot

## Slide 18: Following the proton from source to patient: «full Monte Carlo»

## Slide 19: Following the proton from source to patient: «full Monte Carlo»
LINAC
Beam transport lines
Beam interaction Dose delivery system
phantoms patient
FLUKA
PET tools
MC
Secondaries production
In-beam PET
with CABOTO
Accelerator
Time structure
FULL MC SIMULATION for TULIP
In-beam PET with CABOTO and
Fast Range Verification

## Slide 20: Fast range verification with in-beam PET
    Low duty cycle of carbon ion linacs
    Short half-lived + emitters production (8B most interesting isotope)
    Experimental/simulation data: carbon ion range can be determined with a precision of σ= 1 mm
Notes: Production of short lived (t1/2=9-800 ms) radioisotopes (13O, 12N, 9C, 8B) combined with 0.1% duty cycle of carbon ion linacs can be exploited to verify the particle range

## Slide 21: What happened for me after TERA?

## Slide 22: What happened for me after TERA?
Tomorrow
Today

## Slide 23: Enabling non-invasive & automated treatments of heart arrhythmias using proton beams

## Slide 24: Thank you TERA and happy birthday !

## Slide 25: Thank you TERA and happy birthday !