## Slide 1
E. Picatoste
PMTs Photon detectors and cables
13/12/2022
ECAL Upgrade II Workshop at ICJLab (Orsay)

## Slide 2: PMTs for the Calorimeter Upgrade II
Different detector zones, different needs (gain, ageing, geometry)
  Inner part:
    High doses
    2 channel technologies: SPACAL-W, SPACAL-Pb
  Outer zone
    Shashlik technology
    Somehow more relaxed requirements due to lower radiation doses
Stringent geometry in the innermost zone (15mm)
Ageing is an important limit
  Needs of high number of photoelectrons → maximum amount of integrated charge before degrading the device characteristics
  Total integrated charge ≥ 103 C (to be confirmed/evaluated for MCD)
Cell size case | Channel technology | High G (Imax lim.) | Low G (Imax lim.)
15 mm | SPACAL W | 4k | 1k
30 mm | SPACAL Pb | 4k | 500
40, 60, 120 mm | Shashlik | 100k | 11k
Gain values expected from ageing

## Slide 3: PMT Measurements
PMTs?
  R11187 (TILECAL), R14755U-100, R7600U-20 (MCD), FEU115M,…
Perform measurements to check PMT characteristics at detector expected gains
  Time resolution ~20 ps
  Energy resolution
  Signal linearity
Other measurements
  Ageing
  Radiation hardness
PMT: R11187
TILECAL
Good timing
Lower gain (8 dynodes)
Lower gain (6 dynodes)
Good timing
Smaller but still 1-2 mm large
PMT: R14755U-100

## Slide 4: PMT Laboratory Setup
Optical Beam Attenuator
Diffuser
PMT
Light tight box
Collimator
Single Mode Fiber
Picosecond pulsed Diode Laser
Wavecatcher 8ch
Laser trigger
Power Supply (0-5V)
Output
HV Power Supply
HV
Control + data

## Slide 5: PMT Measurements at the Laboratory
First good results on time resolution, BUT
Gain at laboratory using statistics and pulsed light method is not well understood
  Need to use same Nphe and PMT bias voltage (HV) as test beam for comparison
    Higher Nphe reduce time resolution!
Try 1-Phe method
  Use PACTA transimpedance amplifier (from CTA) to amplify PMT signal
  Analyze data with double Gaussian fit or Bellamy function (E.H. Bellamy et al.INucl. Instr. and Meth. inPhys . Res. A 339 (1994) 468-476)
  Estimate Nphe from the measured (1-Phe) gain?
Chart: V bias (V)
| 345 | 375 | 400 | 450 | 500 | 550 | 650 | 750
R11187 | 0.052663 | 0.03046 | 0.021933 | 0.01528 | 0.020425 | 0.013426 | 0.014539 | 0.020516
R14755U-100 | 0.011208 | 0.011358 | 0.013653 | 0.016012 | 0.021305 | 0.043532 | 0.055593 | 0.076868

## Slide 6: PMT Laboratory Setup 1-Phe
Optical Beam Attenuator
Diffuser
PMT
Light tight box
Collimator
Single Mode Fiber
Picosecond pulsed Diode Laser
Wavecatcher 8ch
Laser trigger
Power Supply (0-5V)
HV Power Supply
HV
Control + data
Diff Output
LV Power Supply
PACTA

## Slide 7: PMT Measurements: 1-Phe Gain
(TILECAL)
PMT R11187 – 800 V (HV)

## Slide 8: PMT Measurements: 1-Phe Gain
(round)
PMT R14755U – 1000V (HV)

## Slide 9: PMT Linearity Measurements at SPS Test Beam (2022)
R11187 (Tilecal)
  Max/Min response ~1% in the 20-100 GeV range
  8 dynodes, tapered board to help against space-charge effects
  More linear than Round MCD
R14755U-100 (Round MCD)
  Max/Min response significantly worse than Tilecal
    10-30% in the 20-100 GeV range at the useful HVs
  Max/Min response ~2.5% even in the 20-40 GeV range at reasonable voltages.
  6 dynodes
  R&D needed
    Tapered board?
    Transistorized base?
    Asking some more dynodes?
, where
Tilecal
Loris Martinazzoli

## Slide 10: PMT Linearity Measurements at SPS Test Beam (2022)
R11187 (Tilecal)
  Max/Min response ~1% in the 20-100 GeV range
  8 dynodes, tapered board to help against space-charge effects
  More linear than Round MCD
R14755U-100 (Round MCD)
  Max/Min response significantly worse than Tilecal
    10-30% in the 20-100 GeV range at the useful HVs
  Max/Min response ~2.5% even in the 20-40 GeV range at reasonable voltages.
  6 dynodes
  R&D needed
    Tapered board?
    Transistorized base?
    Asking some more dynodes?
, where
Round
Loris Martinazzoli

## Slide 11: PMT Linearity Measurements at SPS Test Beam (2022)
Round PMT non-linearity causes?
PMT DC output current should be lower than 1/100 the voltage divider current (Ibias)
  Ibias = 30-100µA
Estimate PMT Iout,DC from
  Amplitudes
  Rates: 10kHz to ~200Hz (at different E)
Iout,DC ≥ Ibias/100 for Vbias ≥ 600V
Plan:
  Reduce base resistors (2MΩ to 470kΩ or lower)
  Test transistorized base
Round PMT ampliudes Vs. HV Vs. E
Loris Martinazzoli

## Slide 12: PMT Linearity Measurements at SPS Test Beam (2022)
Round PMT non-linearity causes?
PMT DC output current should be lower than 1/100 the voltage divider current (Ibias)
  Ibias = 30-100µA
Estimate PMT Iout,DC from
  Amplitudes
  Rates: 10kHz to ~200Hz (at different E)
Iout,DC ≥ Ibias/100 for Vbias ≥ 600V
Plan:
  Reduce base resistors (2MΩ to 470M or lower)
  Test transistorized base
Only if resistive divider
Only if transistorized divider

## Slide 13: PMT Signal Conditioning
Amplifier + Shaper circuit included on the PMT base or FEB under consideration
  To compensate cable attenuation, improve SNR, if necessary, and reduce spill-over effort
  To act as a buffer to help split the signal between paths
  Different ASIC requirements (signal range, gain, noise, BW): add dedicated passive attenuator for each path.
  If at FEB, use differential outputs to ASICs
Detector cells
PMT
12-20m cable
50Ω
Front End Board
I
Energy ASIC
Time ASIC
opamp
?
?
FPGA
GBT
Front-End ASIC
Digital Back-End
Optical Link
ADC
Photodetectors readout solution follows the same scheme as in current ECAL:
  Minimal light transport with PMT sensors near modules,
  All electronics in crates on top of the detector (reduced radiation),
  Connection via analog link (coaxial) ~12m long (up to 20m considered).
ASIC/chipset in TSMC 65nm with separate energy and timing processing paths

## Slide 14: PMT signal Conditioning: Opamp Circuit
Two stage OpAmp based circuit on PMT divider:
  Transimpedance amplifier(OPA847) boost SNR.
  PZ cancellation network reduce pulse width.
  High slew rate coaxial cable driver(OPA695).
Radiation damage limits available commercial components:
  design dedicated ASIC or integrate in front-end board.
Board designed and produced, now under test.

## Slide 15: Summary
Different detector zones, different needs (gain, ageing)
Main PMTs under study: R11187 (Tilecal) and R14755U-100 (round)
PMT tests
  Laboratory: gain, linearity, time resolution
    Gain has been measured with 1-Phe method with reasonable results
    Need to define a method to define the Nphe to be able to compare results
  TB: energy resolution, time resolution, linearity, gain
    Energy and time resolution routinely measured at TB
    Linearity problems with R14755U-100 → review PMT base
  Other: ageing, radiation hardness
PMT signal conditioning is considered with operational amplifier-based circuit
  To compensate cable attenuation, improve SNR and reduce spill-over effort
  To act as a buffer to help split the signal between paths

## Slide 16: Thank you for your attention!

## Slide 17: PMT characteristics summary
(1) SBA: Super bialkali, BI: bialkali, ERMA: Extended red multialkali, EGBI: Extended green bialkali
PMT | Ø outer (mm) | Ø Eff area (mm) | Photo cathode (1) | λ range (nm) | λ peak (nm) | <QE> PMT*GFAG (%) | Gain ~800V | Dynodes | Dark Current ~800V (nA) | t rise (ns) | T.T.S. (ns) | Price (€)
R14755U-100 | 16 | 8 | SBA | 300-650 | 400 | 9.5 | 2.5 x 104 | 6 | 0.1 | 0.4 | ? | 798
R11187 (TILECAL) | 25.7x25.7 | 18x18 | BI | 300-650 | 420 |  | 1.0 x 105 | 8 | 0.25 | 1.5 | ? | 
R7600U-20 (MCD 2020) | 30x30 | 18x18 | ERMA | 300-920 | 530 | 15.3 | 1.0 x 106 | 10 | 20 | 1.6 | 0.35 | 1750
R12421 (2018) | 13.5 | 10 | EGBI | 300-700 | 420 | 10 | 2.0 x 106 | 10 | 1 | 1.2 | 1.4 | 
R7899-20 (ECAL) | 25 | 22 | BI | 300-650 |  | 10.3 | 2.0 x 106 | 10 | 2 | 1.6 | 0.6 | 
PMT: R11187, R7600U-100
PMT: R14755U-100
PMT: R7899-20
PMT: R12421

## Slide 18: PMTs being studied
PMT: R11187
PMT: 7600U-20
PMT: R7899-20
Used on ECAL/HCAL runs ½
Low timing uniformity over photocathode
10 dynodes
MCD 2020
Good timing
Relatively high gain (10 dynodes)
TILECAL
Good timing
Lower gain (8 dynodes)
Lower gain (6 dynodes)
Good timing
Smaller but still 1-2 mm large
PMT: R14755U-100

## Slide 19: PMT gain calculation for test beam at CERN
for digitizer
E is the beam energy (E=150 GeV)
ξ is the max fraction of energy in one cell (ξ=0.4)
Y is the photoelectron yield (Y=20000 GeV-1)
τ is the effective decay time of scintillation (τ=60 ns)
U is the signal amplitude (U=1V)
R is the digitizer input impedance (R=50 Ohm)
e is the electron charge (e=1.6·10-19 C)
G is the PMT gain
The signal charge
then
then, VR7600 = 410V (750V if -40dB attenuation)
for integrating ADC
E is the beam energy (E=150 GeV)
ξ is the max fraction of energy in one cell (ξ=0.4)
Y is the photoelectron yield (Y=20000 GeV-1)
Qmax is the max input charge (200 pC for LeCroy 1182)
e is the electron charge (e=1.6·10-19 C)
G is the PMT gain
The signal charge
then
then, VR7600 = 290V (530V if -40dB attenuation)
u < 1 V
τ = 60 ns

## Slide 20: PMT gain calculation for the LHCb operation
gain is limited by maximum anode current
considering the very central spacal cells ()
Imax is the max anode current (Imax=100 μA)
is the total integrated luminosity (=300 fb-1)
L is the instantaneous luminosity (L=1.5·1034 cm-2s-1)
D is the TID dose for 300 fb-1 (D=1 MGy)
M is the cell weight (M=0.5 kg)
Y is the photoelectron yield (Y=20000 GeV-1)
e is the electron charge (e=1.6·10-19 C)
G is the PMT gain
The dose rate
then
The max transverse energy
at the very central cell
the ≈600 GeV
ξ is the max fraction of energy in one cell (ξ=0.4)
Y is the photoelectron yield (Y=20000 GeV-1)
τ is the effective decay time of scintillation (τ=60 ns)
R is the input impedance (R=50 Ohm)
e is the electron charge (e=1.6·10-19 C)
G is the PMT gain
The signal charge
the signal amplitude:
Then, within the Spacal section, the PMT gain should scale proportionally to sinθ.
i.e. G=156*sinθ/0.03

## Slide 21: The Upgrade-2 FTDR configuration: gain limited by ageing
cell size (mm) | GPMT channel | GPMT REAR | GPMT FRONT
15 | 568 | 812 | 1894
30 | 381 | 544 | 1270
40 | 8210 | 11729 | 27367
60 | 21868 | 31240 | 72893
120 | 21367 | 30524 | 71223
cell size (mm) | U (V) channel | U (V) REAR | U(V) FRONT
15 | 0.42 | 0.60 | 1.41
30 | 2.76 | 3.95 | 9.21
40 | 2.30 | 3.28 | 7.66
60 | 7.39 | 10.56 | 24.63
120 | 3.24 | 4.63 | 10.80
cell size (mm) | LSB (V) channel | LSB (V) REAR | LSB (V) FRONT
15 | 1.03x10-4 | 1.48x10-4 | 3.44x10-4
30 | 6.75x10-4 | 9.64x10-4 | 2.25x10-3
40 | 5.61x10-4 | 8.01x10-4 | 1.87x10-3
60 | 1.80x10-3 | 2.58x10-3 | 6.01x10-3
120 | 7.91x10-4 | 1.13x10-3 | 2.64x10-3
GAIN
Max signal
LSB
PMT gain limited by Imax
The signal charge:
The signal amplitude:
cell size(mm) | Y (phe/GeV) | ξ | Emax (GeV) | τ (ns)
15 | 20000 | 0.4 | 698.4 | 60
30 | 20000 | 0.4 | 453.1 | 4
40 | 3000 | 0.4 | 291.5 | 10
60 | 3000 | 1 | 140.8 | 10
120 | 3000 | 1 | 63.2 | 10
Maximum signal amplitudes are orientative
  If too high, the gain of the PMT can be lowered and PMT lifetime extended
LSB is calculated dividint maximum signals by the minimum ADC step (4096 for 12 bit)
LSB in Calo I is 4.9·10-4 V
  and noise is about 1-2 LSB
Lumi=1.5·1034 cm-2s-1.
SPACAL light yield (zone 15mm and 30mm): 20 ph.el/MeV
Shashlik light yield: 3 ph.el./MeV

## Slide 22: W/Poly - Pulse shape
Studied the average pulse at 60 GeV
  R14755U pulses are shorter than TileCal’s with faster rise and decay time
    Better containment within the bunch crossing
  Average pulse shape affected by the readout configuration
    Slower pulse in single side readout due to the light reflected by the mirror in front
    Studies ongoing with Monte Carlo simulations
Single side readout | Rise time (10-90%) [ns]
R11187 (TILECAL) | 3.6
R14755U-100 (round) | 1.8
PRELIMINARY
PRELIMINARY

## Slide 23: Transient time measurements
Objective: study the transient time uniformity over the photocathode of R7600U-20
Measurements:
  PMT signal output amplitude
  Transient time: time between laser trigger and PMT output pulse
  Transient time spread: standard deviation at a given photocathode position
PMTs:
  R7600U-20: ZF0002, TS0340
  ECAL R7899-20
PMT: R11187
Grid of points, (laser)
12x12 matrix, 1.5mm separation
500 waveforms per point
Bias voltage = 800V
‘Cardboard’ collimator
Collimator
PMT
Laser
EXPERIMENTAL SETUP

## Slide 24: Light readout: PMTs time resolution and TTS uniformity
  Arrival time uniformity over R7600U-20
  CFD 0.5 → total std = 41 ps
  Non uniformity may increase the time resolution
  The use of light guides would reduce the uniformity requirement
  Scan over photocathode with laser light
Collimator
PMT
Laser
PMT: R7600U-20 or R11187
Grid of points, (laser)
  Arrival time uniformity over R11187
  CFD 0.5 → total std = 32 ps
  Arrival time uniformity over R7899-20
  CFD 0.5 → total std = 1.36 ns
PMT: R7899-20
Grid of points, (laser)

## Slide 25: Timing and amplifier at PMT base
20 ps time resolution is mandatory to distinguish interactions → maintain physics performance.
Time path self-trigger threshold could be set for 5GeV depositions in channel.
  Testbeam 2020 results assure 20 ps time resolution above 5GeV of beam energy (ET = 2GeV).
  Time measurement for 2x2 cluster or each channel, depending on occupancy.
Theoretically, time jitter above time resolution for smaller signals.
  Assume
    Dynamic Range=1V, for ET,max = 40GeV and ET,max = 2GeV → min signal 50mV ;
    Rise time (10%-90%) of SPACAL GFAG and MCD: 5ns ;
    noise = 1mV,
  therefore → jitter = noise/slope = 1mV/(50mV/5ns) = 100ps rms
Can we increase the Slope by a factor 5 to achieve 20ps?
  reduce scintillator rising time?
  increase gain, but then pulse saturates at SCA so no CFD correction.
    external time walk correction using energy path measurement?
    external baseline correction with initial SCA samples/event
Alternative to measure time with MCP-PMT layer.
  Sampling rate specification would increase by some factor.
  Sample pattern for digitization may change wrt module PMTs.