## Slide 1: Profile monitors of secondary lines: an urgent request for refurbishment of MWPCs & DWCs
Jocelyn TAN, Inaki ORTEGA RUIZ, Stephane DESCHAMPS
BI-XEI
EA-BI WG, 27th September 2022

## Slide 2: Outlook
Status and motivations
Proposed solutions
Resources and planning

## Slide 3: Multi Wires Proportional Chambers
Gaseous detector developed in early 70’s
Ar + CO2 (50% each)
Anode: 32 tungsten wires ( ~ 20-40 m)/plane
Spill 105 to 1011 particles per second : ionised gas
HV applied to wires :  e- avalanche
Resolution 0.5mm to 1mm
Motorised in/out
In air

## Slide 4: Delay Wire Chambers
Gaseous detector developed in early 80’s
Similar to MWPC
Spill up to 2x105 particles per second
Resolution 0.2 mm
In air

## Slide 5: Scintillating Fibre profile monitors : XBPF
Replaced the wire chambers in NP and EA
100 or 200 scintillating fibres
Light from every fibre read-out by indiv Si PM
No gas, no HV
Vacuum Compatible : reduced material budget
In/out motorisation
Spill : 100 to 107 particles/second/mm2
Resolution : 1 mm

## Slide 6: Problems (1)
Gaseous chambers : 40+ or 50+ years old
Radiation damage : PTFE
Loss of sensitivity  silicon deposit
Tooling and knowhow for repair are lost
Aug ‘22, two DWCs destroyed by users : overpressure of 3 bars
SEM image of a wire with silicon deposit caused by exposure to highly ionising beams
Only paper drawings available

## Slide 7: Problems (2)
Essential for sec. lines commissioning, and daily operation
NA-CONS work units
  Electronics boards : Phase 1
  Replacenent by XBPF : Phase 2, ie deployed in 2030 !
Limited operational spares
 | Operational | In stock (ready) | In stock (broken)
Small MWPC | 21 | 2 | 9
Large MWPC | 16 | 1 | 7
Delay Wire Chamber | 16 | 3 | 12

## Slide 8: Proposed solution
We propose a program to refurbish ten units of each type of wire chamber currently used in the North Area:
  10 small MWPC
  10 large MWPC
  10 Delay Wire Chambers
These quantities should cover eventual replacements of operational chambers until the new XBPF is fully deployed in LS4.

## Slide 9: Resources : 220 kCHF
Material : 60 kCHF
  2 kCHF per monitor : PCB wire planes and local electronics x30 monitors
  re-use mech. frame
Manpower : 160 kCHF, two man.year of an electro-mechanical technician (ORIGIN)
  Study the current design and identify possible improvements.
  Investigate the use of radiation-hard materials.
  Provide CAD models and drawings of the new design.
  Build/optimize a prototype of the new detector.
  Write a procedure for the assembly of the new detector.
  Follow-up of series production: Lead the assembly and quality control of the new detectors.

## Slide 10: Timeline
Due to the critical situation and increasing failure rate, it would be advisable to start in 2023.
The budget profile would be as follows:
  2023: 110 kCHF = 80 kCHF ORIGIN + 10 kCHF prototyping + 20 kCHF electronics.
  2024 : 110 kCHF = 80 kCHF ORIGIN + 30 kCHF series production.
To be urgently presented at the IEFC

## Slide 11: A consolidation request drafted