## Slide 1: Helium Cryostats for Superconducting Devices
Vittorio Parma
CERN – SY/RF Group
Technical Training: Cryostat Engineering for helium superconducting devices
CERN, 7-9 November 2022

## Slide 2: Content
Helium cryostats for SC devices, functions and requirements
Examples of cryostats
Mechanical design and construction of cryostats:
  Materials for cryostats and their properties
  Pressure/vacuum vessels, codes, and norms
  Supporting systems
Heat transfer mechanisms at cryogenic temperatures:
  Thermal radiation and thermal design solutions (thermal shielding, MLI)
  Thermal conduction and thermal design solutions (feedthroughs, heat intercepts)
Notions of cryogenic safety
Calculation tool: Cryostat Toolbox v.1.1
1 h
1 h
1 h

## Slide 3: Mechanical design and construction of cryostats

## Slide 4: Materials in cryostats
Helium tank (for SC device, ph.separator, etc.):
  austenitic st.steels (Fe-Cr-Ni): 304L(1.4307), 316L(1.4404), 316LN
  titanium alloys (Grade 7, Grade-5 (Ti-6Al-4V)) in SRF
Internal (cold) supporting system
  Composites (e.g. GFRE, CFRP, ULTEM)
  St.steel, titanium alloys (tie rods)
Thermal shielding:
  aluminum alloys (series 5xxx, 6xxx, 7xxx)
  Copper (Cu OF, Cu OFE)
Vacuum vessel:
  Low carbon steels (e.g. DIN GS-21 Mn5)
  st.steels (304L)
Cryogenic piping and expansion joints (bellows):
  st.steel (304L)
  Cu (HX tubes)
RF Couplers/HOM (for SRF):
  St.steel Cu plated, Nb, ceramics, etc.
Current leads (for SC magnets)
  Cu, HTS, st.steel, elect. insulating (Kapton), thermal insulating (G10),etc.
Magnetic shielding (for SRF, as needed)
  μ-metal, Cryoperm®, etc.

## Slide 5: Useful material standards for cryostats
Plates and sheets | • EN 10028-1:2007+A1:2009 Flat products made of steels for pressure purposes - Part 1: General requirements • EN 10028-3:2009 Flat products made of steels for pressure purposes - Part 3: Weldable fine grain steels, normalized • EN 10028-7:2007 Flat products made of steels for pressure purposes - Part 7: Stainless steels
Tubes | • EN 10216-5:2004 Seamless steel tubes for pressure purposes - Technical delivery conditions - Part 5: Stainless steel tubes • EN 10217-7:2005 Welded steel tubes for pressure purposes - Technical delivery conditions - Part 7: Stainless steel tubes
Forged blanks | • EN 10222-1:1998 Steel forgings for pressure purposes - Part 1: General requirements for open die forgings • EN 10222-5:1999 Steel forgings for pressure purposes - Part 5: Martensitic, austenitic and austenitic-ferritic stainless steels
Castings | • EN 10213:2007 Steel castings for pressure purposes
Pipe fittings | • EN 10253-4:2008 Butt-welding pipe fittings - Part 4: Wrought austenitic and austenitic-ferritic (duplex) stainless steels with specific inspection requirement
Bars | • EN 10272:2007 Stainless steel bars for pressure purposes
Aluminium | • EN 12392:2000 Aluminium and aluminium alloys – Wrought products – Special requirements for products intended for the production of pressure equipment (choose materials included in the list given in EN 13445-8 section 5.6)

## Slide 6: Relevant mechanical failure mechanisms in cryostats
Helium tanks:
  Rupture (rare!) or permanent deformation due to excessive mechanical stress (pressure loads)
  Helium leaks in welds or material micro-crack
Cryogenic lines/expansion joints:
  Buckling of expansions joints with or without rupture/leaks
  Helium leaks in welds or material micro-crack
Vacuum vessels:
  Buckling under external pressure
  Permanent deformations due to excessive stress concentrations
Internal supporting systems:
  Failure due to excessive mechanical stress
  Failure due to thermo-mechanical stress
  Buckling under compressive load
Thermal shields:
  Permanent deformation due to thermo-mechanical stress (CD/WU transients)
Alignment jacks:
  Break of floor/fixations due to excessive load

## Slide 7: Tensile test material properties
L
F/2
F/2
cross section s0
 = ΔL/L (%)
Stress
s=F/s0 (N/m²Pa)
Strain
Yield tensile strength
YS
Ultimate tensile strength
UTS
Fracture
Necking
Plastic deformation
Elastic deformation
Slope: Young modulus E
 stiffness k=EA/L
σ0.2
0.2% offset line

## Slide 8: Ductile vs. brittle
Ductile behaviour
  (think about lead, gold...)
Stress
Strain
Stress
Strain
  Brittle behaviour
  (think about glass)

## Slide 9: Low temperature effect of tensile properties
Ductile materials tend to become brittle (fragile) at cryogenic temperatures
Increase of the yield and ultimate strength
Young modulus also (marginally) increases.
F/S0
A%
F/S0
T
T1
T2
T3
Fragile fracture
T1
A%
F/S0
T2
>
A%
F/S0
T3
>
UTS
YS

## Slide 10: Mechanical Properties at cryogenic temperatures
Yield, ultimate strength
  Yield and ultimate strengths increase at low temperature
From: Ekin, J.W. Experimental Techniques for Low Temperature Measurements

## Slide 11: Mechanical Properties at cryogenic temperatures (cont.d)
Young modulus
From: Ekin, J. Experimental Techniques for Low Temperature Measurements
Young Modulus
1 : 2024 T4 aluminium
2 : copper-beryllium
3 : K monel
4 : Titanium
5 : SS 304
6 : Carbon Steal C 1020
7 : Steel 9% Ni
From: Technique de l’Ingénieur
Fracture toughness

## Slide 12: Thin shells under external pressure
Buckling:
Non-linear phenomenon. Actual critical pressure depends on initial imperfections: Safety factor needed!
Buckling critical pressure for a thin tube of infinite length
A conservative rule of thumb for stainless steel tubes under vacuum:
If we use a safety factor of 3:
Alternatively, we need to add reinforcements (e.g LHC vessels, t=12 mm)
Example:
r = 500 mm
t > 18.5 mm

## Slide 13: Thin shells under internal pressure
Example: cyl.helium tank
r = 400 mm
p = 4 barg (0.4 MPa)
σa= 160 MPa (304L, Rp0.2= 240 MPa, SF=1.5)
t ≥ 1 mm
(Note: under vacuum (previous calculation)  t ≥ 15 mm !)
For preliminary calculation
For detailed calculation see EN 13458-2 Cryogenic vessels – static vacuum insulated vessels

## Slide 14: Pressure vessel codes regulations
Article 4, paragraph 3
Category | Conf. assessment module | Comment
SEP | None | The equipment must be designed and manufactured in accordance with sound engineering practice. No CE marking and no involvement of notified body.
I | A | CE marking with no notified body involvement, self-certifying.
II | A1 | The notified body will perform unexpected visits and monitor final assessment.
III | B1+F | The notified body is required to approve the design, examine and test the vessel.
IV | G | Even further involvement of the notified body.
For vessels with non-dangerous gases (cryogenic liquids are treated as gas)
Pressure European Directive 2014/68/EC (PED) is a legal obligation in the EU since 2002
  Applies to internal pressure ≥ 0.5 bar gauge
  Vessels must be designed, fabricated and tested according to the requirements defined
  Establishes the conformity assessment procedure depending on the vessel category, which depends on the stored energy, expressed as Pressure x Volume in bar.l
Vacuum vessels
Pressure vessels

## Slide 15: Harmonised codes and standards
Harmonised standards give presumption of conformity with the PED, within their scope. Useful codes for cryostat design and fabrication, including safety devices:
Very useful
guidelines and design rules

## Slide 16: LHC dipole Vacuum Vessels
Main features:
Pipeline standard size: 36-inch OD (1013 mm), 12-mm thick, low carbon steel (DIN GS-21 Mn5) tubes
St. steel extremity flanges
Material resilience: > 28 J/cm2 at -70ºC
Forged cradles, welded rings reinforcements
Dimensional stability:
  Stress relieving
  Final machining to achieve tolerances at interface
Out-doors storage at CERN
Final lathe machining
Forged cradle
Stress relieving
Production:
1250 units
2 firms
4 yrs of production

## Slide 17: LHC thermal shields
Aluminium alloy 6063 extrusions and 1100 top sheets
Al 6063 extrusions
LHC Main Cryostat
(Cross-Section)
Al 1100 sheets
He cooling (50-65K)

## Slide 18
Thermo-mechanical considerations

## Slide 19: Thermal expansion of some materials

## Slide 20: Thermal stress in composite material assemblies: 3 cases
ΔT
A) Restrained component
B) Assembly of different materials
Material 1: α1, E1, A1
Material 2: α2, E2, A2
restrained
ΔT
C) Different cooling ΔT1(t) ≠ ΔT2(t) (different material diffusivity or different cooling)
Material 1: α1, E1, A1
Material 2: α2, E2, A2
restrained
ΔT2(t)
ΔT1(t)

## Slide 21: LHC cryostats: longitudinal thermal contractions
Cold mass, thermal shield, support posts and vacuum vessel must be free with each other to cope with longitudinal thermal contractions
One fixed point per each component
Leave plays to cope with all extreme T cases (ex. Cold mass cold, thermal shield warm)
Guided sliding of cold mass onto vacuum vessel  slot/key
Flexible thermalisation anchors
(example of an LHC special quadupole)
Fixed post
Sliding posts
Δl cold mass= 30 mm (10m x 3mm/m)
Δl thermal shield= 40 mm (10m x 4 mm/m)

## Slide 22
Supporting systems

## Slide 23: LHC Cryomodule, 400 MHz cavities (Nb on Cu) at 4.5 K

## Slide 24: LHC Cryomodule Supporting System
For each cavity assembly (about 500 kg):
7 adjustable st.steel tie rods:
4 vertical
2 horizontal
1 longitudinal

## Slide 25: LHC magnet Cryostats
Main dipole
5-10 K heat intercept
GFRE composite
column
Vac.vessel
Interface: 293 K
Cold mass interface flange: 2 K
50-65 K heat intercept
4-mm thickness, glass-fiber epoxy
Manufactured by Resin Transfer Moulding (RTM):
  Suited to a large-scale industrial production (4’700 units)
  High reproducibility in thermo-mechanical properties

## Slide 26: LHC Supporting system
The design is a trade-off between 2 conflicting requirements:
High flexural stiffness (for mechanical stability)  thick and bulky structure
Low heat in-leaks  thin and slender structure and low conductivity material
 Flexural stiffness/conductivity as figure of merit in the choice of the material
x
q
F
Q
(The acceptable x and θ
are alignment requirements)
(F is mainly the result of
interconnect forces
and gravity component)
F = 5 kN
x = 0.1 mm
ᶿ = 1 mrad
(Q is to be within budget)

## Slide 27: Resin Transfer Moulding
Cutting the fabrics to length
Lay-up stacking
Sealed closing of mould
RTM processing
(complete manufacture cycle: 4 hrs)
Columns as de-moulded

## Slide 28: Lay-up, calculated safety factors, material properties
layup
Triaxial braid and % fiber volumes
Bending cantilever 40 kN. (vacuum barrier load)
Tsai-Wu ply failure: Safety Factor = 2.7
Compression load: 175kN (175% gravity load, transport limit)
Tsai-Wu ply failure: Safety Factor = 6
E modulus: 24 GPa

## Slide 29: SC magnet supporting systems from other machines
Hera dipole
RHIC dipole
Tevatron
SCC

## Slide 30
Supporting system of Tesla/TTF/ILC Cryomodule
Fixed support
Sliding support
DN 300 as back-bone
Invar longitudinal positioner
External supports (jacks)
RF coupler (with bellows)
2 K

## Slide 31: Supporting system of SNS and ESS high β Cryomodules
SNS
“Space-frame” and tie rods
ESS

## Slide 32: Supporting system of PIP II Cryomodule
strongback
“strongback”:
rigid supporting plate in vacuum at RT

## Slide 33: SPL cryomodule: RF Coupler as support
Vacuum vessel
Vertical loading of cavities string

## Slide 34: Heat transfer mechanisms and thermal design

## Slide 35: Heat transfer mechanisms
Thermal radiation
Thermal conduction
Heat intercepts

## Slide 36: Thermal Radiation

## Slide 37: Thermal Radiation
Q
Stefan-Boltzmann’s law
Stefan-Boltzmann’s constant
Q
A, T
, 𝜀 emissivity
F, geometrical view factors(<1) between and between and ), with = .
E, effective emissivity, a function of view factors and emissivity and of the bodies
Figures source: “Heat Transfer”, A.Bejan, John Wilney & Sons, Inc.

## Slide 38: Black body radiation
Total emissive power (integrating over λ):
(k, Boltzmann constant, h, Planck constant, c, speed of light )
In cryogenics: far infrared region
(725 μm at 4 K)
hemispherical emissive power (monochromatic)
In practice: Blackbody emits ~ 420 W/m2 at 293 K , and only ~2 W/m2 at 77 K (200 times less!)
Figures source: “Heat Transfer”, A.Bejan, John Wilney & Sons, Inc.
Hemispherical monochromatic emissive power [W/m∙m2] (energy per unit time, wave-length, and surface area)
=
= 2 π h
=
Planck’s law

## Slide 39: Black and diffuse-gray model
Black surface: α = 1
Diffuse-gray model (good approximation for real surfaces in engineering):
Gray: (λ,T) ≈ (T)
Diffuse emitter/absorber (not directional)
Opaque ( = 0)
(T) = (T)
Figures source: “Heat Transfer”, A.Bejan, John Wilney & Sons, Inc.
Practical application in cryostats: every gap in shields is a black surface
(1 cm2 gap exposed to 293 K, ɛ = 0.2  ~10 mW)
Diffuse-gray model
Black model

## Slide 40: Emissivity
Ref. R.B.Scott, Cryogenic Engineering, (Van Nostrand, New York, 1959; Y.S.Touloukian, Thermophysical Properties of Matter, (Plenum Press, New York, 1995))
𝜀 depends on wave-length and direction (neglected in gray-diffuse model). In metals emissivity rises sharply and up to more than 40% above 50 deg. incidence from the surface normal
Clean, well-polished metallic surfaces have smaller 𝜀, non-metallic materials have higher 𝜀
Metals: 𝜀 decreases with decreasing temperature, almost linearly at cryo temperatures
Non-metals: 𝜀 may decrease or increase with decreasing temperature (increases for organic materials)
Engineering approach. Due to uncertainties on real 𝜀 values, take conservatively high values for preliminary design. Measurement validation may be necessary
Ref. https://neutrium.net/heat-transfer/emissivity/
Ref. Cryogenie, ses applications en supraconductivite.Techniques de l’ingenieur, Ed.1995
Directional distribution

normal

## Slide 41: Radiation between 2 diffuse-gray enclosures
Radiation balance between A1 and A2 surfaces:
For 2 enclosed cylinders or spheres (not necessarily concentric!), F12=1 :
A1
A2
Figures source: “Heat Transfer”, A.Bejan, John Wilney & Sons, Inc.
  For cryostats, to reduce heat load to inner cold surface:
  Limit size of A1 (SC device vessel)
  Reduce A2 (vac.vessel as small as possible)
  Small emissivities: ε1 reduced by low T; ε2 at RT & moderated by A1/A2
(see numerical application in Annex )

## Slide 42: Radiation between 2 diffuse-gray flat plates
Radiation balance between A1 and A2 (A1=A2=A), F12=1 :
T1
T2

## Slide 43: Radiation with an intermediate floating shield
Radiation balance between A1 and A2:
A1 to S gap
S to A2 gap
For flat surfaces approximation, and same ɛ :
 with 1 shield ½ of the rate without shield
=
=
Ts
T1
T2
1
 with N shield 1/(N+1) of the rate without shield
 Can be generalized to N shields
…
…N

## Slide 44: N shields  Multi Layer Insulation (MLI)
Working principles:
Low emissivity of aluminium layer
Multi-layer to enhance radiation protection:
  N+1 radiation shielding
Minimal thermal conduction between reflective layers  interposing of isolating layers (e.g polyester net):
Enhanced performance at low T  use actively cooled shield:
  Lower emissivity of reflective material layers at low T
  Reduce radiation from inner-most layers, cooled at T of shield
  Extract heat at higher thermal shield T  thermodynamic efficiency
Qc
Qr
Reflective film
Insulating spacer
Colder wall (actively cooled
thermal shield)
Warm wall
Tw
Tc < Tw
T
x
Tw
Tc
Text layer
Interleaved reflectors and spacers

## Slide 45: MLI material
LHC MLI:
Reflecting film: 6 µm thick polyethylene teraphthalate (PET) film coated with 400 Å minimum aluminium on each side (Physical Vapor Deposition process)
Spacer: polyester net of very low weight (< 5 g/m2)
10 or 15 layer pre-assembled blankets (outermost layer reinforced with a polyester net)
Blanket overall thickness: > 2 (3) mm for 10 (15) layers
Perforated layers for outgassing: 0.05-0.15 % surface
Interleaved reflectors and spacers
Velcro™ fasteners
Blanket manufacturing
MLI in residential construction
Decoration Al films
MLI in Space applications

## Slide 46: LHC MLI blankets
Features:
1 blanket (10 reflective layers) on cold masses (1.9 K)
2 blankets (15 reflective layers each) on Thermal Shields (50-65 K)
Reflective layer: double aluminized polyester film (2 km2)
Spacer: polyester net
Stitched Velcro™ fasteners for rapid mounting and quality closing
1 blanket on CM, 2 on thermal shield
Mock-ups for complex geometries
Diode box
Vacuum Barrier
MLI on Cold-to-Warm
transitions

## Slide 47: MLI: How many reflective layers (N)?
No straightforward answer ! Not just a matter of N !
Complex system to model, with parameters needing experimental qualification
Radiation vs. conduction
  Radiation reduces as 1/N
  Conduction proportional to packing density (N/mm)
Residual helium gas conduction
  Negligeable at res. pressure p < 10-4 Pa (10-6 mbar)
  Sizeable contribution at res. pressure 10-4 Pa < p < 10-3 Pa (10-5 mbar)
  Dominates at res. pressure p > 10-2 Pa (10-4 mbar)
Packing density should be limited  typically 20-25 N/cm
  Avoid “compressed” blankets (do not put as much MLI as you can fit!)
  Reserve space allocation for MLI blankets
  Consider differential thermal contractions wrt support (Al shields, cold mass…): blankets must remain quite loose at cold
Overall performance depends very much on:
  Application design solutions
  Craftsmanship quality
Radiation
Total heat flow
Conduction
Arbitrary units
No. layers/cm

## Slide 48: MLI application: important tips
Thermal path warm to cold layers
(images: courtesy of Meyer Tool & MFG.)
Handle MLI with gloves
(fingerprints are a black body!)
loose at RT
(MLI contracts more than metal support)
Not tool much MLI
(risk of thermal shorts)
Thot
Tcold
Velcro TM
Reflective layer
Thot
Tcold
Some tapes are
resin coated = black body !
Check conductivity

## Slide 49: Performance measurements on MLI samples
300 K - 77 K, abundant literature (see a few ref. [1], [2])
77 K - 4 K, fewer results. CERN experimental work (see a few ref.[3],[4],[5] )
Some values, in good vacuum (<10-4 Pa) [3]:
N | Tc = 5 K-80 K
10 | ~ 1.0 W/m2
20 | ~ 0.5 W/m2
30 | ~ 0.4 W/m2
[1] T.Nast, Multilayer Insulation Systems, in Handbook of Cryogenic Engineering, ed. J.G.Weisend II, Taylor & Francis, Philadelphia (1998),
[2] M. G. Kaganer, Thermal Insulation in Cryogenic Engineering, Jerusalem, Israel Program for Scientific Translation, (1969),
[3] L.Mazzone et al., Measurements of Multi-Layer Insulation at high boundary temperature, using a simple non-Calorimetric method. Proceedings ICEC 19, 2002, Grenoble, France.
[4] Ph.Lebrun et al., Investigation and qualification of thermal insulation systems between 80 K and 4.2 K, Proceedings ICEC/ICMC, 1992, Kiev, Ukraine
[5] V.Venturi, Thermodynamic and technological optimization of complex insulation systems. PhD thesis, 2019, WUST University, Poland.
N | Tc = 4.2 K
1 or more | ~ 30-60 mW/m2
Tw = 300 K
Tw = 80 K
Virtually no Tc dependency, Tw radiation dominates
Virtually no N dependency
(compare with 1 W/m2 calculated in exercise)
(compare with 110-140 mW/m2 calculated in exercise)

## Slide 50: Practical MLI performance values
MLI samples measured performance is considerably degraded in applications due to the engineered solutions and installation procedures:
  pre-assembled blankets vs. rolled MLI
  complex (non developable) geometries
  junctions, stitches, VelcroTM fasteners, etc.
  higher than ideal packing density (e.g. self-weight compaction)
  thermal shorts between cold and warm layers (e.g. overlapped blanket closure)
  quality of assembly (lab vs. industrial scale)
Engineering values: what values should one use for HL first estimates?
  Take generous margins. Rule of thumb  x 2 sample measurements
LHC cryostat (+) (a thermally well optimized large-scale application), average measure on 2.7 km length sector :
  293 K to 50 K thermal shield with 30 MLI layers:  ~1 W/m2
  50 K to 2 K magnet with 10 MLI layers:  ~ 54 mW/m2
(+) C. Maglioni, V. Parma: “Assessment of static heat loads in the LHC arc, from the commissioning of sector 7-8”, LHC Project Note 409, 2008.

## Slide 51: Summary of radiation HL exercises (for details see spare slides)
(th.shield at 260 K)
(th.shield at 50 K)
(th.shield at 143 K)
x 1000 reduction !
Notes: 45 minuti
Pausa QA: 10’

## Slide 52: Thermal Conduction

## Slide 53: Thermal Conduction
When a T gradient exists in a body, there is a heat transfer from the high T region to the low T region (Fourier Law):
For one-dimensional problems (ex. a bar or tube):
k is the thermal conductivity (W m-1 K-1), normally a function of P,T, material structure, non-homogeneity, anisotropy (ex. Composite materials).
k is strongly T-dependent and non-linear at low T
“good conductors” vs. “poor conductors”  k range ~ 5-6 orders of magnitude
Tw
Tc
6 decades

## Slide 54: Thermal conductivity of various materials
Metals: mainly free electron contribution
Resistivity from electron scatter with phonons (lattice vibrations) α T2 and with impurities α 1/T
Increasing effect of impurities in less pure metals and dominates in alloys
Maximum of conductivity reduces and shifts towards higher T with increasing impurities
Al and st.steel alloys: conductivity α T
Non cristalline materials and insulators dominated by phonon contribution

## Slide 55: Thermal conductivity in epoxy-based composite materials
Low conduction (little electron conduction), essentially phonon driven
Anisotropic structure, (fibers/matrix) with constituents-specific thermal conductivity properties
Conductivity highly depends on:
  Material (fiber) orientation
  Ratio between fibre and matrix (Vf)
Glass-fiber is the “conductive” fraction and also the “structural” constituent
Epoxy is the “isolating” fraction material and is the “less structural” constituent
Vf typically around 40-70%
Cryogenic grade G10-CR and G11 are widely used standard laminates in various forms (plates, tubes, rods). Properties data can be found in: https://trc.nist.gov/cryogenics/index.html
Chart: Conductivity (W/m.K)
| 4 | 6 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 120 | 140 | 160 | 180 | 200 | 250 | 300
l long/trasv | 0.078 | 0.093 | 0.102 | 0.111 | 0.125 | 0.14 | 0.154 | 0.174 | 0.188 | 0.22 | 0.253 | 0.285 | 0.318 | 0.356 | 0.388 | 0.446 | 0.504 | 0.543 | 0.582 | 0.616 | 0.671 | 0.702
l epaisseur | 0.076 | 0.09 | 0.099 | 0.108 | 0.121 | 0.134 | 0.146 | 0.161 | 0.173 | 0.195 | 0.219 | 0.24 | 0.261 | 0.282 | 0.302 | 0.336 | 0.368 | 0.391 | 0.413 | 0.434 | 0.474 | 0.505
Ilong/trasv | 0.086 | 0.1 | 0.107 | 0.116 | 0.129 | 0.146 | 0.163 | 0.185 | 0.201 | 0.237 | 0.275 | 0.311 | 0.348 | 0.392 | 0.429 | 0.495 | 0.561 | 0.605 | 0.649 | 0.687 | 0.748 | 0.781
l epaisseur | 0.084 | 0.097 | 0.104 | 0.113 | 0.125 | 0.141 | 0.156 | 0.174 | 0.187 | 0.215 | 0.243 | 0.269 | 0.295 | 0.322 | 0.347 | 0.39 | 0.429 | 0.458 | 0.486 | 0.511 | 0.56 | 0.592
Chart: Conductivity (W/m.K)
| 4 | 6 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 120 | 140 | 160 | 180 | 200 | 250 | 300
l matrice | 0.06 | 0.067 | 0.075 | 0.083 | 0.088 | 0.095 | 0.1 | 0.104 | 0.109 | 0.114 | 0.122 | 0.128 | 0.134 | 0.139 | 0.145 | 0.156 | 0.165 | 0.173 | 0.181 | 0.189 | 0.207 | 0.224
l fibre | 0.09 | 0.11 | 0.12 | 0.13 | 0.15 | 0.17 | 0.19 | 0.22 | 0.24 | 0.29 | 0.34 | 0.39 | 0.44 | 0.5 | 0.55 | 0.64 | 0.73 | 0.79 | 0.85 | 0.9 | 0.98 | 1.02
(Data from Schwartzberg F.H., Cryogenic Material Data Handbook,
Martin Marietta Corrp, Denvers, Co, 1970)
Longitudinal and transversal conductivity laminate studies for LHC supports, (based on data from Schwartzberg F.H.), collaboration CERN-ESIGEC (Univ.Savoie), 1997
Glass-fiber dry fabric
Standard G10-CR plate and rods

## Slide 56: LHC supports
Two intercepts, 5K & 75K
One intercept, 75K
No intercepts
Support posts measurements
  49 ± 10% mW @ 1.9K
  450 ± 10% mW @ 4.5-10K
  7.1 ± 5% W @ 50-65K
conductivity acceptance criteria at low T

## Slide 57: The conduction equation (unidirectional case)
Longitudinal
conduction
Internal
heat
generation
Thermal
inertia
(Fourier law of heat conduction)
(change of internal energy)

## Slide 58: No heat deposition and steady-state
If k constant with T:
If k = k(T):
Thermal conductivity
Integral (conductance)
(Linear solution)
For metal alloys (ex.steels) at low T:
(quadratic solution)
T
x
T0
TL
L
with boundary conditions:
x=0, T=T0
X=L, T=TL
TL
T0
L

## Slide 59: Thermal conductivity integrals (conductance) for some materials [W/m]
Highest T (Lowest T =4.2 K) | 20 K | 80 K | 290 K
OFHC Copper | 11000 | 60600 | 152000
DHP Copper | 395 | 5890 | 46100
Aluminium 1100 | 2740 | 23300 | 72100
Aluminium 2024 | 160 | 2420 | 22900
Stainless steel AISI 304 | 16,3 | 349 | 3060
Typical Glass-fiber/Epoxy Composite G-10 | 2 | 18 | 153
Tw
4.2K
Qc

## Slide 60: …more conductivity integrals

## Slide 61: Electrical network analogy
The inverse of the thermal conductance  thermal resistance:
Analogy with the electrical resistance (replace q with I, T with V):
For variable k, define an average value kAV:
We can therefore model a complex thermal conductivity problem by elementary thermal resistances Ri, and solve the network by using Kirckhoff’s laws.

## Slide 62: Thermal conduction with uniform heat deposition
x
To
Beam of length L, thickness t, width w (= 1m);
beam thermalized on one side at To
uniform heat deposition from one side, q (W/m2)
considering k constant with T
Boundary conditions: a) for x= 0 T=To (heat sink); b) q=0 for x=L (isolated tip)
Integrating and imposing the 2 boundary conditions:
x
T
T0
TL
L
practical interest:
calculate minimum thickness to limit ΔTmax
t

## Slide 63: Thermal diffusivity
Conduction equation (no heat generation)
For small variations of k with x:
Solid bar cooled at one tip
T perturbation travels towards warm end:
tc, critical time for T perturb. to reach warm tip:
Thermal diffusivity [m2/s]
T
L
x
Ti
T∞
T0=T∞
t increase

## Slide 64: Heat capacity
Specific heats:
For (incompressible) substances:
Change of internal energy U (or enthalpy):
Cooling (or heating) power:
Specific heats/unit vol.
sometimes preferred
Below 10 K enthalpy of
helium overtakes all solids
(from J. Ekin, Experimental Techniques for Low-Temperature Measurements)

## Slide 65: Measuring Thermal Performance of Cryostats

## Slide 66: Measure of Heat Loads to liquid helium tanks
Isothermal cooling (2-phase helium tanks)
He boil-off along saturation line
  Latent Heat (Lv) of vaporisation
  Isothermal cooling (T constant if P constant)
  Measure p, and
  Liquid/gas volumetric corrective factor
vapor
Latent heat of vaporization
Density of saturated liquid helium
liquid volume replaced by fraction of boil-off vapor (no filling)
measure
from diagrams
Example:
  Mass flow measurement: 7 g/s
  P measured: 1.78 bar
  (T on saturation line (phase diagram): 4.8 K)
  Lv = 14’116 (J/kg) (tables, or state equations)
  𝑣𝑔 = 3.0 10-2 (m3/kg); 𝑣l = 9.3 10-3 (m3/kg); (phase diagrams)
   = 143 W
He boil-off liquid level measurement:
  Measure level l with SC level gauge
  Boil-off mass flow measure through reservoir geometrical correlation between level and volume: f(l), and knowledge of density ρ and Lv
1.45 !
SC level gauge

## Slide 67: Measure of Heat Loads to liquid helium tanks (cont.d)
Non-isothermal cooling
Calorimetry in superfluid helium (LHeII):
  LHe II very high thermal conductivity  isothermal bath
  Measure internal energy change (closed system) in time Δt
  Needs precise measure of helium content
  and sub-mK resolution measurement of T at T< 2.17 K

## Slide 68
Non-isothermal cooling
Forced internal (conduit) convection of single-phase fluid:
  Non-isothermal cooling: enthalpy change of fluid (here cp assumed constant)
  Needs the measure of mass flow, and T at entry and exit of conduit
(Tout – Tin)
Tin
Tout
Measure of Heat Loads to Thermal Shields

## Slide 69: Forced Convection heat transfer in conduits
D
L
U
Forced flow of coolant fluid in round tube cooling lines
Considering hydro-dynamically and thermally fully developed flow
Uniform wall heat flux (linear T profiles)
Convection heat transfer from wall to fluid:
Reynolds No.:
Nusselt No.:
For laminar flow: constant
For turbulent flow, NuD=f(ReD, Pr):
(Dittus and Boelter formula)
ReD>2000  turbulent flow,
ReD<2000  laminar flow
Enthalpy balance along the line L:

## Slide 70: Frictional pressure drop in a tube
Pressure drop along tube
Experimental plots or semi-empirical formulations (ex. Colebrook formulation)
Ref. “Heat Transfer”, A.Bejan, John Wilney & Sons, Inc.

## Slide 71: Feedthroughs and heat intercepts

## Slide 72: Feedthroughs in cryostats
LHC Main dipole
Magnet Support Posts (GFRE)
Solid conduction paths:
Supporting systems
Current leads
RF main coupler
Instrumentation feed-throughs
Beam tubes Cold-to-Warm (CWT) transitions
Necks (vertical cryostats)
Current leads
Neck
LHC instrumentation capillary at assembly
CWT SPL cryomodule

## Slide 73: Heat intercepts (heat sinking) at intermediate temperatures
L
A
Tw=293K
Tc=2K
simple solid conduction
L
A
Q @ 80K
L1
L
A
Q @ 80K
Q @ 8K
L1
L2
1 heat intercept at optimal distance
2 heat intercepts at optimal distance
 L1
 L1, L2

## Slide 74: LHC supports
Two intercepts, 5K & 75K
One intercept, 75K
No intercepts
5-10 K heat intercept
GFRE composite
column
Vacuum vessel
Interface flange: 293 K
Cold mass interface flange: 2 K
50-65 K heat intercept
4-mm thickness, glass-fiber epoxy
Manufactured by Resin Transfer Moulding (RTM):
  Suited to a large-scale industrial production (4’700 units)
  High reproducibility in thermo-mechanical properties
Minimizing using efficiency factors:
C1 = 16 w/w @ 75 K
C2 = 210 w/w @ 5 K
C3 = 990 w/w @ 1.9 K

## Slide 75: Vapour cooling in solid conduction
x, T(x)
@ Tl
Tw
L
A
Large enthalpy in He vapours (1550 kJ/kg from 4.2K to 300K)  usable cooling capacity
If , which is the residual heat to the bath, is equivalent to the evaporation (i.e. self-sustained):
Vapor cooled wall
Assuming perfect exchange (Tvapor = Twall)
attenuation factor (w.r.t. solid conduction)
Reduced heat conduction in self-sustained helium cooling for selected technical materials
(ref. G. Vandoni, Heat Transfer, CAS, Erice 2003)

## Slide 76: SPL cryomodule: actively cooled RF main couplers
Vapor cooled RF coupler (double-walled tube). Effective cooling with gas regulation:
Conduction heat loads (RF off)
Conduction + RF EM resistive heating (RF on)
Double-walled tube:
~Φ 100 mm
Inner wall: st.steel, 1.5-mm thick
External wall: st.steel, 2-mm thick
260

## Slide 77: Notions of Cryogenic Safety

## Slide 78: Cryostats and safety
Cryostats include inventory of cryogenic fluids, potentially unstable stored energy, and sometimes large electro-magnetic stored energy in superconducting devices
Managing safety in cryostat covers multiple aspects: safety of personnel and equipment
  Risk assessment: cryostat as part of a cryogenic system in an environment
  Safety hazard from relief of cryogens to the environment: ODH, burns, escape paths, safety training, use of personnel protection equipment, risk to adjacent equipment, etc.
   Understanding pressure hazards and making the correct choice of the pressure relief devices to protect from overpressure of the cryostat envelopes

## Slide 79: Liquid helium (and nitrogen) stored energies
One LHC dipole:
Magnetic stored energy: ~7 MJ
Helium content: 375 liters
Total stored energy (in liquid): 74 MJ (x10 times stored magnetic energy!)
Total magnet cold surface: ~4 m2
Latent heat per unit volume: He x60 lower than N2  (roughly) venting a cryostat with ~6 liter air (~N2) boils off He in 1 dipole
Air condensation power on cold surface with MLI: 6 kW/m2 (40 kW/m2 on bare surface !)
Vacuum insulation breach: peak power deposition: 24 kW  ~40 s to boil-off liquid, ~50 min. to deplete liquid stored energy,
Property | Unit | 4He | N2
Tboil.(at 1’013 mbar) | [K] | 4.2 | 77.3
Liquid density (at Tboil.) | [kg/m3] | 125 | 804
Density ratio ρliquid/ρgas (gas at boiling temperature and 1’013 mbar) | - | 7.4 | 175
Gas/liquid volume ratio (gas at 273 K, 1’013 mbar) | - | 702 | 652
Latent heat of evaporation (at Tboil., and 1’013 mbar) | [kJ/kg] | 21 | 199
Latent heat of evaporation per unit vol. (at Tboil., 1’013 mbar) | [kJ/l] | 2.6 | 160
Enthalpy change from Tboil. to 300 K per unit vol. (at 1’013 mbar) | [kJ/l] | 194 | 189
Total liquefaction energy in 1 liter of liquid | [kJ/l] | 196.6 | 349
Expansion work (isobar at 1’013 mbar, reversible) for 1 liter of liquid | [kJ/l] | 71 | 66
Vacuum
Air
Cryogenics fluid
 Large volumetric inventories and energies released in reduced time !  pressure relief devices !

## Slide 80: Pressure hazards in cryostats
Potential pressure hazards:
  Compressors connected to cryo lines
  Heating of “trapped” volumes (typically in a circuit between valves) during warm-ups
  Helium leak to insulation vacuum, with consequent increased conduction/convection heat loads to cryogenic liquid vessels
  Cryo-condensed air leaks on cold surfaces and consequent pressure increase and increased conduct/convection heat loads during warm-ups
  …
  A) Heating/vaporization of cryogens from sudden release of stored energy in SC device (e.g. quench or arcing in a SC magnet circuit)
  B) Accidental release of cryogenic fluid to higher T surfaces (thermal shield and vacuum vessel), and consequent pressure increase and increased of conduction/convection heat loads to cold surfaces
  C) Accidental air venting of insulation vacuum with sudden condensation on cold surfaces, helium boil-off and pressure build-up
Usually the most critical
Vacuum
Air
Cryogenics fluid
C) Air venting
Vacuum
Air
Cryogenics fluid
B) Rupture of helium tank
Vacuum
Air
Cryogenics fluid
A) Quench (or el.arc)
(Snowball effect A) + B) + C)  LHC event of Sept.2008)

## Slide 81
How things can go wrong…LHC 19th sept. 2008
…5 MW arc in an interconnect
…~50 t of longitudinal;pressure force, ~50 magnets displaced …
…uprooting of jacks
~2 t of helium in first 2 minutes, estimated He peak mass flow of 20 kg/s
vacuum relief device clogged with MLI
Report of the Task Force on the incident of 19 September 2008
at the LHC, CERN EDMS 992622

## Slide 82: General approach for sizing of safety devices
Risk analysis & mitigation:
Make a thorough risk analysis and evaluate risk hazards
Identify mitigation measures (e.g. protections of exposed bellows and flanged connections)
Identify severity of consequences and appreciate probability of the event
Define the maximum credible incident(s) (MCIs) and design the safety relief system accordingly
The safety relief system must be designed to keep pressure rise within the limits of the Maximum Allowable Working Pressure (MAWP)
Design steps for the MCI:
Estimate the heat flux to helium and its conversion to mass flow rates to be discharged
Choose the type of safety device (burst disks, valves, plates) and size the safety device (orifice area and set pressure). Make use of standards and safety device manufacturers datasheet
Check the sizing of piping (generally designed for normal operation) to the relief device and increase if necessary (pressure drop limited to a few % of total discharge Δp)
 Methodology explained in a new standard EN 17527 just issued devoted to Helium cryostats - Protection against excessive pressure

Heat flux (W)  Relieving mass flow rate (kg/s)

## Slide 83: EN 17527 Helium cryostats - Protection against excessive pressure
Scope includes:
SC magnet cryostats
SRF cavity cryomodules
Ultra-low T refrigerator systems
Coldboxes of helium refrigerators and liquefiers
Helium distribution systems including valve boxes
Overall concept:
Standardization of the approach
Specification of procedures and minimum requirements
Risk assessment, definition of scenarios
Protection concepts (single-stage, multi-stage)
Dimensioning rules of pressure relief devices (HEM Model)

## Slide 84: Safety devices
Safety valves
Burst discs
Pressure relief plate
  A few suppliers
  Herose, Rembe, Ramseyer, Leser…
1.
2.
3.

## Slide 85
EN Standards for Safety Relief Devices
EN 764-7:2002 Pressure equipment. Part 7: Safety systems for unfired pressure equipment
EN ISO 4126-1:2013 / Safety devices for protection against excessive pressure - Part 1: Safety valves (ISO 4126-1:2013) 12/08/2016
EN ISO 4126-2:2013 / Safety devices for protection against excessive pressure - Part 2: Bursting disc safety devices
EN ISO 4126-3:2006 / Safety devices for protection against excessive pressure - Part 3: Safety valves and bursting disc safety devices in combination (ISO 4126-3:2006) 12/08/2016
EN ISO 4126-4:2013 / Safety devices for protection against excessive pressure - Part 4: Pilot-operated safety valves (ISO 4126-4:2013) 12/08/2016
EN ISO 4126-5:2013 / Safety devices for protection against excessive pressure - Part 5: Controlled safety pressure relief systems (CSPRS) (ISO 4126-5:2013) 12/08/2016
EN ISO 4126-7:2013 / Safety devices for protection against excessive pressure - Part 7: Common data (ISO 4126-7:2013) 12/08/2016
EN 13648-1:2008 / Cryogenic vessels - Safety devices for protection against excessive pressure - Part 1: Safety valves for cryogenic service 12/08/2016
EN 13648-2:2002 / Cryogenic vessels - Safety devices for protection against excessive pressure - Part 2: Bursting disc safety devices for cryogenic service 12/08/2016
ISO 21013-3:2016 Cryogenic vessels -- Pressure-relief accessories for cryogenic service -- Part 3: Sizing and capacity determination
EN ISO 4126-6:2013 / Safety devices for protection against excessive pressure - Part 6: Application, selection and installation of bursting disc safety devices
EN ISO 4126-10:2013 / Safety devices for protection against excessive pressure - Part 10: Sizing of safety valves for gas/liquid two-phase flow
EN 13648-3:2002 Cryogenic vessels. Safety devices for protection against excessive pressure. Determination of required discharge. Capacity and sizing
ISO 21013-1:2008 Cryogenic vessels -- Pressure-relief accessories for cryogenic service -- Part 1: Reclosable pressure-relief valves
(ISO/AWI 21013-1 Cryogenic vessels -- Pressure-relief accessories for cryogenic service -- Part 1: Reclosable pressure-relief valves [Under development])
ISO 21013-2:2007 Cryogenic vessels -- Pressure-relief accessories for cryogenic service -- Part 2: Non-reclosable pressure-relief devices + ISO 21013-2:2007/Amd 1:2018
ISO 21013-4:2012 Cryogenic vessels -- Pilot operated pressure relief devices -- Part 4: Pressure-relief accessories for cryogenic service + ISO 21013-4:2012/Amd 1:2019
EN 17527 Helium cryostats - Protection against excessive pressure NEW (May 2022)

## Slide 86: Thank you for your attention

## Slide 87: References and selected bibliography
A.Bejan, Heat Transfer, J.Wiley & Sons, Inc
CRYOGENIE, SES APPLICATIONS EN SUPRACONDUCTIVITE, IIF/IIR 1995, Techniques de l’ingenieur.
Superconducting Magnets, M.Wilson, Oxford Science Publications
R.R.Conte, Éléments de Cryogénie, Masson & Cie, Éditeurs.
Steven W.Van Sciver, Helium Cryogenics (Second Edition), The International Cryogenics Monograph Series, Springer.
K. Mendelssohn, The quest for absolute zero, McGraw Hill (1966)
R.B. Scott, Cryogenic engineering, Van Nostrand, Princeton (1959)
G.G. Haselden, Cryogenic fundamentals, Academic Press, London (1971)
R.A. Barron, Cryogenic systems, Oxford University Press, New York (1985)
B.A. Hands, Cryogenic engineering, Academic Press, London (1986)
S.W. van Sciver, Helium cryogenics, Plenum Press, New York (1986)
K.D. Timmerhaus & T.M. Flynn, Cryogenic process engineering, Plenum Press, New York (1989)
Proceedings of CAS School on Superconductivity and Cryogenics for Particle Accelerators and Detectors, Erice (2002)
  G. Vandoni, Heat transfer
Proceedings of CAS School on Superconductivity, Erice (2013)
  B. Boudoy, Heat transfer and cooling techniques at low temperature
J.Ekin, Experimental Techniques for Low-Temperature Measurements: Cryostat Design, Material Properties and Superconductor Critical-Current Testing
CRYOCOMP® is a database code of the state and thermal properties for technical materials.
NIST Cryogenic Materials database: http://www.cryogenics.nist.gov/MPropsMAY/material%20properties.htm

## Slide 88: Acknowledgments
The content of the training includes contributions from many colleagues from CERN and from external institutes. I wish to thanks them for their precious contributions.

## Slide 89: Annex: numerical calculations exercises

## Slide 90: Exercise: calculate a thermal shield thickness in an LHC-type cryostat
Symmetry: point opposite of cooling line is an isolated tip
Take ½ shield: L = = ½ 2.35 m = 1.17 m
Uniform heat deposition:
Let’s set the maximum = 5 K
Take k aluminium (Al6061) at ~80 K: 85 W/(K m)
 t = 1.7 mm
As, Ts ɛs
Ae, Te , ɛe
qshield
L,
qx(x)
q

## Slide 91: Exercise: calculate helium heat transfer in a thermal shield in an LHC-type cryostat
Non-isothermal cooling along an L=3 km long thermal shielding line
Cooling line: D=80 mm, roughness aluminium drawn tube (ɛ = 0.05 mm)
Tin = 60 K, Tout < 75 K
pin = 5 bar
Calculate mass flow, heat exchange parameters, and pressure drop
Assumptions: Tm(x) = ½(Tout+Tin) = const. ; ΔTin = ΔTout (wall with uniform heat flux )
Take Tout = 75 K
Tm = ½ (75 K + 60 K) = 67.5 K
= qshield ∙ L = 2.55 W/m ∙ 3000 m = 7’650 W (see previous calculation of thermal shield heat load)
From enthalpy balance (with Tin = 60 K, Tout = 75 K, and Cp(67.5 K) = 5300 J/kg/K)  = = 0.096 kg/s
Calculate flow properties and thermal convection heat exchange (use Cryostat Toolbox):
  Re = 1.9 10+5 , turbulent OK !
  V = 5.4 m/s
  Pr = 7.03
  Nu = 345
  h = 247 (W/m2/K)
  ΔTin = ΔTout = 0.04 K (Tw ≈ Tm)
Calculate pressure drop (use Cryostat Toolbox):
  Δp = 380 mbar (~8%)

## Slide 92
Radiation calculation exercises for
an LHC-like cryostat

## Slide 93: Exercise: Radiation in LHC-like cryostat
Thermal radiation HL for a 1-m cryostat unit length
Vacuum vessel diameter: 1m (A2= π x 1= 3.14 m2)
Cold mass diameter: 0.5 m (A1= π x 0.5 = 1.57 m2)
T1 cold mass: 2 K
T2 vac.vessel: 293 K
Ɛ1 = 0.12 (st.steel, mec.polished, 2 K)
Ɛ2 = 0.2 (low carbon.steel, mec.polished, 293 K)
q1-2 = 63.5 W (41.6 W/m2)
Sensitivity wrt Ɛ uncertainties:
  Ɛ1 = 0.18 (+50%), Ɛ2 = 0.20  q1-2 = 87 W (+37%)
  Ɛ1 = 0.12, Ɛ2 = 0.30 (+50%)  q1-2 = 69 W (+9%)
  Ɛ1 = 0.18 (+50%), Ɛ2 = 0.30 (+50%)  q1-2 = 98 W (+54%)
  Note: assuming black bodies (Ɛ1 = Ɛ2 = 1)  q1-2 = 656 W ! (x10!)
 Influence of Ɛ1 is 4 times more important than Ɛ2
A1, ɛ1
A2, ɛ2

## Slide 94: Exercise: Radiation in LHC-like cryostat (cont.d)
Di
De
Sensitivity with vac.vessel diam. (De): ~12 W/m
(Di=0.5m, De=1m)
(Di=0.5m, De=1m)
(m)
(m)
= 12.3 W/m
= 102 W/m
 Di x 8 times more sensitive than De (for similar D ratios)
Sensitivity with inner vessel diam. (Di): ~ 100 W/m

## Slide 95: Exercise: Radiation in LHC-like cryostat (cont.d)
Floating Al shield
A2, ɛ2
A1, ɛ1
As, ɛs
Floating Al shield:
Thermal shield diameter: 0.75 m (As= π x 0.75 = 2.35 m2)
F1s= 1; F2s= 1
Ts = 80 K (first guess)
Ɛs,1 = Ɛs,2 = 0.1 (Aluminum, mec.polished, 80 K)
q1-2 = q1-s = qs-2 = 28.5 W
Calculate Ts by trial and error to obtain power balance q1-s = qs-2
Ts = 260 K  increase Ɛs,1 = Ɛs,2 (0.15?) and recalculate
q1-2 = q1-s = qs-2 = 35.4 W (to be compared to 63.5 W without shield)
q1-s = 35.4 W (22.5 W/m2), qs-2 = 35.4 W (15 W/m2)
 1 floating thermal shield reduces to almost ½ the radiation to the low T (close to flat plates approximation for this geometry)

## Slide 96: Exercise: Radiation in LHC-like cryostat (cont.d)
MLI with N reflectors, floating shields
Radiation between shield and vac.vessel (flat plat approximation):
N = 30
Ts = 80 K (first guess)
Ɛe = 0.2 (low carbon.steel, mec.polished, 293 K)
Ɛs = 0.08 (Aluminium reflector, electrolytical deposition, 80 K)
Ɛav = ½ (Ɛs+ Ɛe) (average emissivity)
= ½ (+ ) (average area)
qs-e = 2.7 W (0.14 W/m2) (with Ts = 80 K )
As, Ts ɛs
Ai, Ti , ɛi
Ae, Te , ɛe
Radiation between shield (80 K) and helium vessel (formula between enclosed cylinders):
qi-s = 0.23 W (0.012 W/m2) (with Ts = 80 K )
Calculate Ts by trial and error (and tune Ɛs) to obtain power balance qs-e = qi-s (floating shield condition)
Ts = 143 K,  Ɛs= 0.1
qs-e = qi-s
qs-e = 2.6 W (1.65 W/m2)
qi-s = 2.6 W (0.96 W/m2)
 MLI (30 layers) reduce the HL to the thermal shield (and cold surface) by a factor x25 wrt no shield

## Slide 97: Exercise: Radiation in LHC-like cryostat (cont.d)
MLI with N reflectors, actively cooled shield
Radiation between shield and vac.vessel (flat plat approximation):
N = 30
Ts = 80 K (now an input)
Ɛs = 0.08 (Aluminium reflector, electrolytical deposition, 80 K)
Ɛav = ½ (Ɛs+ Ɛe) (average emissivity)
= ½ (+ ) (average area)
qs-e = 2.78 W (1.0 W/m2)
As, Ts ɛs
Ai, Ti , ɛi
Ae, Te , ɛe
Radiation between shield (80 K) and helium vessel (2 K) (formula between enclosed cylinders):
qi-s* = 0.23 W (0.14 W/m2) (x10 times less than with floating shield !)
Radiation heat extraction from thermal shield (not to be forgotten!):
qshield = qs-e - qi-s = 2.55 W
qshield
qs-e
qi-s
Actively cooled shield with MLI dramatically reduces radiation to the low T: this is a standard practice in cryostats!
(*) can be further reduced to 0.18 W (0.11 W/m2 ) with Ɛi = 0.08 (for example 1 or more MLI layers on cold mass)

## Slide 98: Exercise: Residual gas conduction in LHC-like cryostat
Typical dimension (space between surfaces): d = 0.25 m
Good insulation vacuum:
  p = 10-6 mbar (10-4 Pa) (“standard” good vacuum in cryostats)
Calculate λmolecule (between shield and inner vessel):
  λmolecule = 20 m (~ x100 d)  molecular regime
Calculate residual gas conduction with Kennard’s law:
  Qres = 0.034 W (0.02 W/m2)
Negligeable additional contribution wrt previously computed radiation:
  qi-s = 0.23 W (0.14 W/m2)
As, Ts ɛs
Ai, Ti , ɛi
Ae, Te , ɛe
qshield
qs-e
qi-s
Qres
Degraded insulation vacuum:
  p = 10-5 mbar (10-3 Pa) (degraded vacuum from helium leak in cryostats)
Calculate λmolecule (between shield and inner vessel):
  λmolecule = 2 m (~ x10 d)  still molecular regime
Calculate residual gas conduction with Kennard’s law:
  Qres = 0.34 W (0.22 W/m2)
 gas conduction contribution (Qres = 0.22 W/m2) now higher than radiation (qi-s = 0.14 W/m2). Total HL to cold surface: 0.34 W/m2

## Slide 99: Summary of radiation HL exercises
(th.shield at 260 K)
(th.shield at 50 K)
(th.shield at 143 K)

## Slide 100
Spare slides

## Slide 101: Material mechanical properties at cryogenic temperatures
Source: Techniques de l’ingenieur

## Slide 102: Material mechanical properties at cryogenic temperatures
Source: Techniques de l’ingenieur

## Slide 103: Thermal conductivity in solids (& metals)
In metals, the electron contribution dominates :
The movement of conduction electrons is impeded by scatter: interactions with phonons, and interactions with impurities/imperfections. We can introduce thermal resistivities:
The conductivity is attributed to the movement of conduction electrons (“electron gas”), ke, and the effects of phonon lattice vibrations, kl :
Therefore for metals, the resistivity can be expressed as:
And has a maximum, which shifts at higher T with increasing impurity (see coppers ) and vanishes for highly impure alloys (see steels), as impurity scatter dominates, and for which at T < RT:
For metals, analogy between electron thermal and electrical diffusion  Wiedemann-Franz law :
  Good agreement at T<< RT >> T
  Better agreement from T<< to T>> with increasing impurities
e-
N+
Lorentz constant

## Slide 104: Vacuum in cryostats
2-4.5 K surfaces have high pumping speed & capacity.
Gas species have very low vapour pressure except helium.
Without helium leaks, equilibrium pressures ≤ 10-6 mbar are obtained
Good insulation vacuum

## Slide 105: He gas conduction
Viscous regime (λmolecule << d):
  At “high” gas pressure: classical conduction (Fourier law-like)
  independent of pressure
  (natural convection should be added !)
Molecular regime (λmolecule >> d):
  At “low” gas pressure
  Kennard’s law
  Conduction is proportional to p
  A1 is the surface receiving the flow
  α(T)  accommodation coefficient depending on gas species, T1, T2 and surface geometry (applicable for flat parallel surfaces, coaxial cylinders and spheres)
d
T1
A1
T2 (>T1)
A2
λmolecule = mean free path
R = gas constant [J.K-1.mol-1] [kg/mol]
Accommodation coefficient αi
Temp. [K] | Helium
300 | 0.3
80 | 0.4
20 | 0.6
4 | 1
with λ Cv
ρ specific heat at constant pressure
/

## Slide 106: Thermal conductivity in solids (& metals)
In metals, the electron contribution dominates :
The movement of conduction electrons is impeded by scatter: interactions with phonons, and interactions with impurities/imperfections. We can introduce thermal resistivities:
The conductivity is attributed to the movement of conduction electrons (“electron gas”), ke, and the effects of phonon lattice vibrations, kl :
Therefore for metals, the resistivity can be expressed as:
And has a maximum, which shifts at higher T with increasing impurity (see coppers ) and vanishes for highly impure alloys (see steels), as impurity scatter dominates, and for which at T < RT:
For metals, analogy between electron thermal and electrical diffusion  Wiedemann-Franz law :
  Good agreement at T<< RT >> T
  Better agreement from T<< to T>> with increasing impurities
e-
N+
Lorentz constant

## Slide 107: 2 main mechanisms of interest for cryostats
Vaporisation in pool boiling (2-phase)
  Latent Heat (Lv) of vaporisation
  Isothermal cooling (T constant if P constant)
Vapor
U
Tin
Tout
Cryogen | Latent Heat (at 1tm) [kJ/kg] | [mg/s] | [l/h] (liquid) | [l/min] (gas NTP)
Helium | 21 | 48 | 1.38 | 16.4
Nitrogen | 199 | 5 | 0.02 | 0.24
Vaporisation under 1 W heat load
Forced internal (tube) convection of single-phase fluid:
  Non-isothermal cooling: enthalpy change of fluid (cp assumed constant)
  Depends on thermo-hydraulics of the flow (see next slide)
  Used in cooling of thermal shields (supercritical He)
For more cooling mechanisms see dedicated lectures

## Slide 108: Helium as a coolant for SC devices
Property | Units | 4He | N2
Boiling T (at 1 atm) | K | 4.2 | 77.3
Critical temperature | K | 5.2 | 126.1
Critical pressure | 105 Pa | 2.23 | 33.1
Latent Heat of evaporation (at 1 atm) | kJ/kg | 21 | 199
Enthalpy between T boiling and 300K | kJ/kg | 1550 | 233
Liquid density (boiling at 1 atm) | kg/m3 | 125 | 810
Saturated vapor density (at 1 atm) | kg/m3 | 17 | 4.5
Gas density (at 1 atm 273.15K) | kg/m3 | 0.18 | 1.25
Liquid viscosity (at boiling T) | μPa.s | 20 | 17
Pressurized He II, Magnets: LHC, Tore Supra
Pressurized He I, Magnets, HERA, Tevatron
Saturated He II, SRF: CEBAF, TTF, SNS, EXFEL, ESS, ILC
Pool boiling He I, SRF: HERA, LEP, KEKB
Supercritical helium: cooling of thermal shielding
(2.17 K, 49.7 mbar)
(5.2 K, 2.27 bar)
(1.8 K, 16 mbar)
(2 K, 30 mbar)
4.2 k
Density of saturated liquid helium
Latent heat of vaporization
(ref. S. Van Sciver, Helium Cryogenics)
Specific heat of saturated liquid

## Slide 109: Heat capacity
Specific heats:
For (incompressible) substances:
Change of internal energy U (or enthalpy):
Cooling (or heating) power:
Specific heats/unit vol.
sometimes preferred
Below 10 K enthalpy of
helium overtakes all solids
(from J. Ekin, Experimental Techniques for Low-Temperature Measurements)

## Slide 110: Design stresses for some materials
Material | Rp1.0 (MPa) | f (MPa) | ftest (MPa)
1.4306 (304L) | 240 | 160 | 228
1.4435/1.4404 (316L) | 260 | 173 | 247
1.4406/1.4429 (316LN) | 320 | 213 | 304
AW 5083-O/H111 |  | 83 | 
Design stresses for plates less than 12 mm thick applicable to membrane stress (safety factor 1.5 included) according to EN 13445-3
For stainless steels:
For aluminium-magnesium alloys:
Material | Rp1.0 /Rm (MPa) | f (MPa) | ftest (MPa)
AW 5083-O/H111 | 125/270 | 83 | 119

## Slide 111: Best practices
Using a coherent set of standards throughout the lifecycle of the cryostat is the simplest and safest approach. As an example when using only EN harmonised standards:
  Error margins of pressure relief devices are taken into account in the design rules
  The design rules are only applicable if the material has enough ductility
  Materials certified for pressure vessels have measured minimum fracture toughness
  Safety factors included in buckling formulae take into account shape imperfections up to the allowable tolerances layed out in the manufacturing section of the standards
  The extent of welding inspection must be compatible with the joint coefficient used in thickness calculations
  Coherence of test pressure and testing procedure with the design rules

## Slide 112: Liquid helium pumping
Pressure reduction by pumping above helium liquid at 4.2 K, with latent heat L(T) and specific heat c(T):
Neglecting external heat loads and vessel heat capacity
Integrating over T and m, with m0 initial helium mass:
(ref. R.R.Conte, Elements de Crogenie)
100%, 4.2 K
50%, 1 K

## Slide 113: LHC dipole cryostat assembly
Towing through, sliding on vacuum vessel
Assembly bench

## Slide 114: Expansion joints (bellows) for cryogenic lines
Compensation of longitudinal thermal contractions
Compensation of transversal and angular offsets (not twist!)
Extensive use in cryo-magnet interconnections
Beware of bellows stability issues in pressurised lines:
  Stiff guiding
  Limit transversal and angular off-sets
LHC String 1, CERN
RHIC, BNL