## Slide 1: MQXFB Magnet Optical feedthrough issues
Keziban KANDEMIR, Michael GUINCHARD, Sylvain MUGNIER.
EN-MME-EDM Mechanical measurement lab
Thanks to Hervé, Tavis and Nicolas (x2) and Patrick (SM18) for the support and the efficiency !

## Slide 2
Optical feedthrough : A challenge…

## Slide 3: Introduction : Coil Instrumentation based on FBG
Magnet component | Number & Directions | Bridge configuration | Type
SHELL | 20 (Θ) 20 (Z) | SG Quarter bridge | Cr – Ni / Polyimide HBM LY11-6/350
COILS | 12 (Z) 12 (Θ) | FBG | FemtoSecond® 4 arrays with 6 FBG’s
RODS | 4 (Z) | SG Full bridges | Cr – Ni / Polyimide HBM LC11-3/350
Vacuum Vessel

## Slide 4: Introduction : MQXF story
 | Optical fiber performance |  |  | 
 | Loading | Cold temp. | Powering tests | RT after TC
MQXFS6b | 100% | 100% | 100% | 100%
MQXFS6c | 100% | 100% | 100% | 100%
MQXFS6d | 100% | 100% | 100% | 100%
MQXFS7a | 100% | 50% (1) | 50% (1) | 100%
MQXFS7b | 100% | 100% | 100% | 100%
MQXFS7c | 100% | 100% | 100% | 100%
MQXFS7d | 100% | 100% | 100% | 100%
MQXFBP1 | 100% | FAILED (1) | FAILED (1) | 100%
MQXFBP2 | 100% | 100% | 75% (2) | 100%
MQXFBP3 | 100% | 12% (1) | FAILED (1) | 100%
(1) Hardware optical line issues (connectors or feedthrough); (2) SPD tracking issues;
MQXFS7a
MQXFS7b; 7c; 7d
MQXFBP1: VACOM
MQXFBP2 and P3: SEDI-ATI

## Slide 5: Introduction : Optical power budget
Loss in optical fiber performance
Effects on FBG measurements

## Slide 6: Introduction : Optical feedthrough box design
SEDI-ATI company
2 flanges:
Flange 1 : 8 FC/APC connectors welded on it,
Flange 2 : 8 optical fibers with FC/APC connectors in the extremities screwed on it.
Optical fibers with stainless steel buffer.
510 mm
250 mm
60 mm
1.9 K
RT

## Slide 7: MQXFBP3 Failure analysis
RT
Full spectra saved during cooldown for each optical fiber channels.
FBG peaks
4.5 K (QLFA1_S_TT140 Magnet inlet temperature)
85 K (QLFA1_S_TT151)
4.5 K stable
CO118 RE Z : -115 um/m
CO118 RE T : -297 um/m
Gradual optical losses during cool down were observed (only 2 FBGs remained operational at 4.5 K)

## Slide 8: MQXFBP3 Failure analysis
5m patchcords
4.8 m Fischer patchcords
6.5 m Fischer patchcords
CH3 (CO118)
Optical feedthrough box 0.82 m in total
1st flange
2nd flange
Fischer to stainless steel optical fiber connection
2 peaks at this location showing to loss from both connectorizations in the 2nd flange.
Investigations with the Luna distributed strain sensing interrogator, based on optical frequency domain reflectometer (OFDR).
Outside cryostat connections
Cold Mass
Vacuum Vessel

## Slide 9: MQXFBP3 Failure analysis
Mounting operation story:
Wrong flange first welded; it was difficult to screw the feedthrough lines;
A tooling was used to tighten the connectors.
Additional tests were performed afterward to quantify this effect :

## Slide 10: MQXFBP3 status
1
Only one bi-axial measurements valid at 4.2 K
BUT :
Peaks were still visible for two coils during the steady state conditions at 80K;
Measurements at 80K are valid thanks to the thermal compensation :
Correction factor from BP2 was used to calculate the expected strain at 4.2 K and cross-checked with CO118 RE measurement :
4.5 K stable
CO118 RE Z : -115 um/m
CO118 RE T : -297 um/m
 | Non compensated strain values at 77 K (um/m) | Compensated strain values at 77 K (um/m) | Expected strain values at 4.2 K (um/m)
CO117_RE_T_FBG | -2996 | -271 | -255
CO117_MI_T_FBG | -3088 | -363 | -341
CO117_LE_T_FBG | na | na | na
CO118_RE_T_FBG | -2974 | -249 | -234
CO118_MI_T_FBG | -3106 | -381 | -358
CO118_LE_T_FBG | -2974 | -249 | -234
CO117_RE_Z_FBG | -2922 | -197 | -181
CO117_MI_Z_FBG | -3020 | -295 | -271
CO117_LE_Z_FBG | -2940 | -215 | -198
CO118_RE_Z_FBG | -2860 | -135 | -124
CO118_MI_Z_FBG | -3055 | -330 | -304
CO118_LE_Z_FBG | -2900 | -175 | -161
CO118 RE Z (measured) : -115 um/m
CO118 RE T (measured) : -297 um/m

## Slide 11: MQXFBP3 Status
Chart: MQXFBP3 - Coils - Azimuthal Strain
| After Loading [Bdg. 180 - RT - 11/12/2020] | Welding 3rd pass [RT - 12/02/2021] | Thermal cycle n°1 [Bdg. SM18 - 4,5 K -16/08/2022] | SMI2 Before cryostat removal [RT - 17/10/2022]
CO115_RE_T_AP_FBG | -513 | -568 | 0 | 0
CO115_MI_T_AP_FBG | -689 | -697 | 0 | 0
CO115_LE_T_AP_FBG | -655 | -701 | 0 | 0
CO117_RE_T_AP_FBG | -440 | -514 | -255 | -488.4
CO117_MI_T_AP_FBG | -543 | -593 | -341 | -618.4
CO117_LE_T_AP_FBG | -515 | -580 | 0 | -572.4
CO118_RE_T_AP_FBG | -485 | -524 | -234 | -463.5
CO118_MI_T_AP_FBG | -697 | -728 | -358 | -727
CO118_LE_T_AP_FBG | -551 | -573 | -234 | -569.6
CO119_RE_T_AP_FBG | -501 | -560 | 0 | 0
CO119_MI_T_AP_FBG | -468 | -517 | 0 | 0
CO119_LE_T_AP_FBG | -488 | -577 | 0 | 0
Expected strain calculated by extrapolation from measurements at 80 K

## Slide 12: MQXFBP3 Status
Chart: MQXFBP3 - Coils - Longitudinal Strain
| After Loading [Bdg. 180 - RT - 11/12/2020] | After SS shell welding [Bdg. 180 - RT - 05/05/2022] | Thermal cycle n°1 [Bdg. SM18 - 4,5 K -16/08/2022] | SMI2 Before cryostat removal [RT - 17/10/2022]
CO115_RE_Z_AP_FBG | -3 | 136.3 | 0 | 0
CO115_MI_Z_AP_FBG | -53 | 66.5 | 0 | 0
CO115_LE_Z_AP_FBG | 24 | 170.4 | 0 | 0
CO117_RE_Z_AP_FBG | -19 | 112.8 | -181 | 66.36
CO117_MI_Z_AP_FBG | -82 | 23.04 | -271 | 121.3
CO117_LE_Z_AP_FBG | 28 | 172.8 | -198 | 202.1
CO118_RE_Z_AP_FBG | 101 | 234.5 | -115 | 189.2
CO118_MI_Z_AP_FBG | -34 | 80.42 | -304 | 204.7
CO118_LE_Z_AP_FBG | -41 | 98.77 | -161 | 97.99
CO119_RE_Z_AP_FBG | 39 | 171.8 | 0 | 0
CO119_MI_Z_AP_FBG | 76 | 200.1 | 0 | 0
CO119_LE_Z_AP_FBG | -68 | 87.6 | 0 | 0
Expected strain calculated by extrapolation from measurements at 80 K

## Slide 13: MQXFB02 solution

## Slide 14: MQXFB02 solution
Flange successfully used in MQXFBP2
8 feedthrough lines were replaced due to dismounting issues (on BP2)

## Slide 15: Testing plan prior to magnet installation
Optical feedthrough line testing down to 1.6 K
Optical feedthrough flange testing at RT
Full box testing at SM18 down to 4.2 K
Full box testing after welding at RT

## Slide 16: Optical feedthrough line testing
Distributed measurement with Rayleigh Backscattering sensing;
FBG connected at the out-patch cord: measurement with Bragg-meter.

## Slide 17: Optical feedthrough line testing
Feedthrough 4
Loss = -5.4 dB
Loss = -5.9 dB
Rayleigh backscattering distributed sensing
Fiber Brag grating measurement

## Slide 18: Optical feedthrough line testing
Feedthrough line number | Measured loss at 1.6 K with RBS (dB) | Measured loss at 1.6 K with FBG (dB)
1 | Not tested | Not tested
2 | -3.1 | -3.0
3 | -4.4 | -4.9
4 | -5.4 | -5.9
5 | -4.6 | -4.5
6 | -8.7 | -9.8
7 | Not tested | Not tested
8 | Not tested | Not tested
Same loss observed between both measurements RBS and FBG interrogator.

## Slide 19: Optical feedthrough flange testing at RT
1 m patch cord
Optical fiber with FBG
OFDR Optical interrogator
FBG Optical interrogator
1 m patch cord
Measurements in two steps (Flange and Flange + Feedthrough);
Distributed measurement with Rayleigh Backscattering sensing;
FBG connected at the out patch-cord: measurement with Bragg-meter.

## Slide 20: Optical feedthrough flange testing at RT
Flange channel 2
Rayleigh backscattering distributed sensing
Flange channel | Loss Flange (dB) | Loss Flange +feedthrough (dB)
1 | -1.5 | -1.6
2 | -0.5 | -1.5
3 | -1.4 | -1.4
4 | -0.2 | -1.4
5 | -0.4 | -1.7
6 | Broken | Broken
7 | -1.2 | -1.0
8 | -1.2 | -1.2
Study on the optical loss due to both connectors of the flange.
Moving average over 5 gage pitches (5*0.65=3.25 mm).

## Slide 21: Full box testing down to 4.5 K

## Slide 22: Summary of the testing plan
Optical feedthrough line testing down to 1.6 K
Optical feedthrough flange testing at RT
Full box testing at SM18 down to 4.2 K
Full box testing after welding at RT

## Slide 23: Summary of the testing plan
Optical feedthrough line testing down to 1.6 K
Optical feedthrough flange testing at RT
Full box testing at SM18 down to 4.2 K
Full box testing after 3 points welding at RT
Optical losses in dB at 1.6 K | Optical losses in dB at RT | Optical losses in dB at RT | Optical losses in dB at 4.2 K
Not tested | -1.5 | -1.6 | - 5
-3.1 | -0.5 | -1.5 | - 9
-4.4 | -1.4 | -1.4 | - 11
-5.4 | -0.2 | -1.4 | - 12
-4.6 | -0.4 | -1.7 | - 8
-8.7 | Broken | Broken | Broken
Not tested | -1.2 | -1.0 | - 8
Not tested | -1.2 | -1.2 | - 6
Better performance at cold comparing to MQXFBP2 ;
Green light for a complete welding and installation !

## Slide 24: After welding…
Only two channels are still working well without optical losses ;
NEXT OPERATIONS :
Extract and test at RT feedthrough lines from BP3 in progress;
The box from B02 will be dismounted to replace the 6 not working feedthrough lines (expected this week).

## Slide 25: Conclusions and Next steps
MQXFBP3 :
Gradual optical losses during cool down were observed (only 2 FBGs remained operational at 4.5 K);
Thanks to RBS electronic a deepest study was performed to understand the exact location of the optical losses. The issue was confirmed in the box at the connection between the flange and the feedthrough line : this issue is probably explained by the tightenning procedure of the connectors during mounting operation;
MQXFB02 :
A full validation test campaign was performed in short delay in order to assess the optical losses of every sub-components of the OFB down to 1.8 K;
The complete OFB assembly (only tack welded) was also tested down to 4.2 K at SM18 with optical losses slightly better that the MQXFBP2 OFB ;
An issue after complete welding was observed and need additional replacement from BP3 – Same issue was observed after dismounting operation of BP2 ;
Additional developments :

## Slide 26
Copper coated fiber brazed (250 um) on a copper insert, brazed on a SS flange.
Already tested at mbars
Two coils of MQXFB02 are equipped with high reflectivity FBG’s to improve the optical power budget at cold.