## Slide 1: MQXFB Magnet Optical feedthrough box : Status and perspectives
Keziban KANDEMIR, Michael GUINCHARD, Sylvain MUGNIER.
EN-MME-EDM Mechanical measurement lab
Thanks to Hervé, Tavis and Nicolas (x2) and Patrick (SM18) for the support and the efficiency !

## Slide 2
Optical feedthrough : A challenge…

## Slide 3: Introduction : Coil Instrumentation based on FBG
Magnet component | Number & Directions | Bridge configuration | Type
SHELL | 20 (Θ) 20 (Z) | SG Quarter bridge | Cr – Ni / Polyimide HBM LY11-6/350
COILS | 12 (Z) 12 (Θ) | FBG | FemtoSecond® 4 arrays with 6 FBG’s
RODS | 4 (Z) | SG Full bridges | Cr – Ni / Polyimide HBM LC11-3/350
Vacuum Vessel

## Slide 4: Introduction : MQXF story
 | Optical fiber performance |  |  | 
 | Loading | Cold temp. | Powering tests | RT after TC
MQXFS6b | 100% | 100% | 100% | 100%
MQXFS6c | 100% | 100% | 100% | 100%
MQXFS6d | 100% | 100% | 100% | 100%
MQXFS7a | 100% | 50% (1) | 50% (1) | 100%
MQXFS7b | 100% | 100% | 100% | 100%
MQXFS7c | 100% | 100% | 100% | 100%
MQXFS7d | 100% | 100% | 100% | 100%
MQXFBP1 | 100% | FAILED (1) | FAILED (1) | 100%
MQXFBP2 | 100% | 100% | 75% (2) | 100%
MQXFBP3 | 100% | 12% (1) | FAILED (1) | 50%
(1) Hardware optical line issues (connectors or feedthrough); (2) SPD tracking issues;
MQXFS7a
MQXFS7b; 7c; 7d
MQXFBP1: VACOM
MQXFBP2 and P3: SEDI-ATI

## Slide 5: Introduction : Optical power budget
Loss in optical fiber performance
Effects on FBG measurements (MQXFBP2)
How to improve our optical power budget :
Minimize the optical losses ;
Improve the dynamic of the optical interrogators ;
Improve the FBG design (high reflectivity gratings).

## Slide 6: Introduction : Optical feedthrough box design
SEDI-ATI company
2 flanges:
Welded flange : 8 FC/APC connectors welded on it,
Patchcord flange : 8 optical patchcords with FC/APC connectors in the extremities screwed on it.
Optical fibers with stainless steel buffer.
510 mm
250 mm
60 mm
1.9 K
RT
Welded Flange
Patchcord Flange

## Slide 7: MQXFBP3 Failure analysis
RT
Full spectra saved during cooldown for each optical fiber channels.
FBG peaks
4.5 K (QLFA1_S_TT140 Magnet inlet temperature)
85 K (QLFA1_S_TT151)
4.5 K stable
CO118 RE Z : -115 um/m
CO118 RE T : -297 um/m
Gradual optical losses during cool down were observed (only 2 FBGs remained operational at 4.5 K)

## Slide 8: MQXFBP3 Failure analysis
5m patchcords
4.8 m Fischer patchcords
6.5 m Fischer patchcords
CH3 (CO118)
Optical feedthrough box 0.82 m in total
Patchcord flange
Welded flange
Fischer to stainless steel optical fiber connection
Investigations with the Luna distributed strain sensing interrogator, based on optical frequency domain reflectometer (OFDR) done at 1.9 K
Outside cryostat connections
Cold Mass
Vacuum Vessel

## Slide 9: MQXFBP3 Failure analysis
Mounting operation story:
Wrong flange first welded; it was difficult to screw the patchcord lines;
A tooling was used to tighten the connectors.
Additional tests were performed afterward to quantify this effect :

## Slide 10: MQXFBP3 status
Only one bi-axial measurements valid at 4.2 K
BUT :
Peaks were still visible for two coils during the steady state conditions at 80K;
Measurements at 80K are valid thanks to the thermal compensation :
Correction factor from BP2 was used to calculate the expected strain at 4.2 K and cross-checked with CO118 RE measurement :
1
4.5 K stable
CO118 RE Z : -115 um/m
CO118 RE T : -297 um/m
 | Non compensated strain values at 77 K (um/m) | Compensated strain values at 77 K (um/m) | Expected strain values at 4.2 K (um/m)
CO117_RE_T_FBG | -2996 | -271 | -255
CO117_MI_T_FBG | -3088 | -363 | -341
CO117_LE_T_FBG | na | na | na
CO118_RE_T_FBG | -2974 | -249 | -234
CO118_MI_T_FBG | -3106 | -381 | -358
CO118_LE_T_FBG | -2974 | -249 | -234
CO117_RE_Z_FBG | -2922 | -197 | -181
CO117_MI_Z_FBG | -3020 | -295 | -271
CO117_LE_Z_FBG | -2940 | -215 | -198
CO118_RE_Z_FBG | -2860 | -135 | -124
CO118_MI_Z_FBG | -3055 | -330 | -304
CO118_LE_Z_FBG | -2900 | -175 | -161
CO118 RE Z (measured) : -115 um/m
CO118 RE T (measured) : -297 um/m

## Slide 11: MQXFBP3 Status
Chart: MQXFBP3 - Coils - Azimuthal Strain
| After Loading [Bdg. 180 - RT - 11/12/2020] | Welding 3rd pass [RT - 12/02/2021] | Thermal cycle n°1 [Bdg. SM18 - 4,5 K -16/08/2022] | SMI2 Before cryostat removal [RT - 17/10/2022]
CO115_RE_T_AP_FBG | -513 | -568 | 0 | 0
CO115_MI_T_AP_FBG | -689 | -697 | 0 | 0
CO115_LE_T_AP_FBG | -655 | -701 | 0 | 0
CO117_RE_T_AP_FBG | -440 | -514 | -255 | -488.4
CO117_MI_T_AP_FBG | -543 | -593 | -341 | -618.4
CO117_LE_T_AP_FBG | -515 | -580 | 0 | -572.4
CO118_RE_T_AP_FBG | -485 | -524 | -234 | -463.5
CO118_MI_T_AP_FBG | -697 | -728 | -358 | -727
CO118_LE_T_AP_FBG | -551 | -573 | -234 | -569.6
CO119_RE_T_AP_FBG | -501 | -560 | 0 | 0
CO119_MI_T_AP_FBG | -468 | -517 | 0 | 0
CO119_LE_T_AP_FBG | -488 | -577 | 0 | 0
Expected strain calculated by extrapolation from measurements at 80 K

## Slide 12: MQXFBP3 Status
Chart: MQXFBP3 - Coils - Longitudinal Strain
| After Loading [Bdg. 180 - RT - 11/12/2020] | After SS shell welding [Bdg. 180 - RT - 05/05/2022] | Thermal cycle n°1 [Bdg. SM18 - 4,5 K -16/08/2022] | SMI2 Before cryostat removal [RT - 17/10/2022]
CO115_RE_Z_AP_FBG | -3 | 136.3 | 0 | 0
CO115_MI_Z_AP_FBG | -53 | 66.5 | 0 | 0
CO115_LE_Z_AP_FBG | 24 | 170.4 | 0 | 0
CO117_RE_Z_AP_FBG | -19 | 112.8 | -181 | 66.36
CO117_MI_Z_AP_FBG | -82 | 23.04 | -271 | 121.3
CO117_LE_Z_AP_FBG | 28 | 172.8 | -198 | 202.1
CO118_RE_Z_AP_FBG | 101 | 234.5 | -115 | 189.2
CO118_MI_Z_AP_FBG | -34 | 80.42 | -304 | 204.7
CO118_LE_Z_AP_FBG | -41 | 98.77 | -161 | 97.99
CO119_RE_Z_AP_FBG | 39 | 171.8 | 0 | 0
CO119_MI_Z_AP_FBG | 76 | 200.1 | 0 | 0
CO119_LE_Z_AP_FBG | -68 | 87.6 | 0 | 0
Expected strain calculated by extrapolation from measurements at 80 K

## Slide 13: MQXFB02 solution

## Slide 14: MQXFB02 solution
Welded Flange successfully used in MQXFBP2
8 patchcord lines were replaced due to dismounting issues (on BP2)

## Slide 15: Testing plan prior to magnet installation
Patchcord lines testing down to 1.6 K
Welded flange
testing at RT
Full box testing at SM18 down to 4.2 K
Full box testing after welding at RT

## Slide 16: Patchcord lines testing
Distributed measurement with Rayleigh Backscattering sensing;
FBG connected at the out-patch cord: measurement with Bragg-meter.

## Slide 17: Patchcord lines testing
Feedthrough 4
Loss = -5.4 dB
Loss = -5.9 dB
Rayleigh backscattering distributed sensing
Fiber Brag grating measurement

## Slide 18: Patchcord lines testing
patchcord line number | Measured loss at 1.6 K with RBS (dB) | Measured loss at 1.6 K with FBG (dB)
1 | Not tested | Not tested
2 | -3.1 | -3.0
3 | -4.4 | -4.9
4 | -5.4 | -5.9
5 | -4.6 | -4.5
6 | -8.7 | -9.8
7 | Not tested | Not tested
8 | Not tested | Not tested
Same loss observed between both measurements RBS and FBG interrogator.

## Slide 19: Welded flange testing at RT
1 m patch cord
Optical fiber with FBG
OFDR Optical interrogator
FBG Optical interrogator
1 m patch cord
Measurements in two steps (Welded Flange and Welded Flange + Patchcord lines);
Distributed measurement with Rayleigh Backscattering sensing;
FBG connected at the out patch-cord: measurement with Bragg-meter.

## Slide 20: Welded flange testing at RT
Flange channel 2
Rayleigh backscattering distributed sensing
Welded Flange channel | Loss Welded Flange (dB) | Loss Welded Flange + patchcord lines (dB)
1 | -1.5 | -1.6
2 | -0.5 | -1.5
3 | -1.4 | -1.4
4 | -0.2 | -1.4
5 | -0.4 | -1.7
6 | Broken | Broken
7 | -1.2 | -1.0
8 | -1.2 | -1.2
Study on the optical loss due to both connectors of the flange.
Moving average over 5 gage pitches (5*0.65=3.25 mm).

## Slide 21: Full box testing down to 4.5 K
Channel | Optical losses in dB at 4.2 K
1 | - 5
2 | - 9
3 | - 11
4 | - 9.5
5 | - 8
6 | -27
7 | - 8
8 | - 6

## Slide 22: Summary of the testing plan
Patchcord lines testing down to 1.6 K
Welded flange
testing at RT
Full box testing at SM18 down to 4.2 K
Full box testing after welding at RT

## Slide 23: Summary of the testing plan
Patchcord lines testing down to 1.6 K
Welded flange
testing at RT
Full box testing at SM18 down to 4.2 K
Full box testing after 3 points welding at RT
Optical losses in dB at 1.6 K | Optical losses in dB at RT | Optical losses in dB at RT | Optical losses in dB at 4.2 K
Not tested | -1.5 | -1.6 | - 5
-3.1 | -0.5 | -1.5 | - 9
-4.4 | -1.4 | -1.4 | - 11
-5.4 | -0.2 | -1.4 | - 9.5
-4.6 | -0.4 | -1.7 | - 8
-8.7 | Broken | Broken | -27
Not tested | -1.2 | -1.0 | - 8
Not tested | -1.2 | -1.2 | - 6
Better performance at cold comparing to MQXFBP2 ;
Green light for a complete welding and installation !

## Slide 24: After welding…
Only two channels operational after complete welding of the box ;
Diagnostic with OFDR to confirm the position ;
Probably an issue related to the handling and not to the welding process ;
Solution implemented :
Patchcord lines from BP3 were extracted and tested at RT;
Spare patchcord lines were tested at RT;
The box from B02 was dismounted to replace the 6 not working patchcord lines with new one ;
A new batch of patchcord lines was ordered (delivery in December)

## Slide 25: Today at SM18…
OFDR traces.

## Slide 26: Our expectations at cold…
Optical lines | Optical losses Optical Feedthrough box Today @RT | Total Optical losses with FBG’s Today @RT | Optical losses Optical Feedthrough box (1) @4.2 K | Expected Total Optical losses with FBG’s (2) @1.9 K
Coil 121 In | - 6.5 dB | -12.5 dB | Not measured | -27.5 dB
Coil 121 Out | -1.3 dB | -3.1 dB |  -8 dB | -18.2 dB
Coil 123 In | -1.8 dB | -2.7 dB |  -5.5 dB | -17.7 dB
Coil 123 Out | -1.2 dB | -3.2 dB |  -7.5 dB | -18.3 dB
Coil 124 In | -2.1 dB | -2.9 dB |  -10.5 dB | -17.9 dB
Coil 124 Out | -1.6 dB | -3.5 dB |  -9.5 dB | -18.5 dB
Coil 125 In | -0.6 dB | -3.2 dB |  -8 dB | -18.2 dB
Coil 125 Out | -1.9 dB | -3.5 dB |  -9 dB | -18.5 dB
: Optical losses measured with old pigtail lines during tests at SM18
: Total optical losses with FBG’s measured today @RT – 15 dB (MQXFBP2 experience)

## Slide 27: Our expectations at cold…
Optical lines | Optical losses Optical Feedthrough box Today @RT | Total Optical losses with FBG’s Today @RT | Optical losses Optical Feedthrough box (1) @4.2 K | Expected Total Optical losses with FBG’s (2) @1.9 K
Coil 121 In | - 6.5 dB | -12.5 dB | Not measured | -27.5 dB
Coil 121 Out | -1.3 dB | -3.1 dB |  -8 dB | -18.2 dB
Coil 123 In | -1.8 dB | -2.7 dB |  -5.5 dB | -17.7 dB
Coil 123 Out | -1.2 dB | -3.2 dB |  -7.5 dB | -18.3 dB
Coil 124 In | -2.1 dB | -2.9 dB |  -10.5 dB | -17.9 dB
Coil 124 Out | -1.6 dB | -3.5 dB |  -9.5 dB | -18.5 dB
Coil 125 In | -0.6 dB | -3.2 dB |  -8 dB | -18.2 dB
Coil 125 Out | -1.9 dB | -3.5 dB |  -9 dB | -18.5 dB
: Optical losses measured with old pigtail line during tests at SM18
: Total optical losses with FBG’s measured today @RT – 15 dB (MQXFBP2 experience)

## Slide 28: Conclusions and Next steps
MQXFBP3 :
Gradual optical losses during cool down were observed (only 2 FBGs remained operational at 4.5 K);
Thanks to RBS electronic a deepest study was performed to understand the exact location of the optical losses. The issue was confirmed in the box at the connection between the flange and the feedthrough line : this issue is probably explained by the tightenning procedure of the connectors during mounting operation;
MQXFB02 :
A full validation test campaign was performed in short delay in order to assess the optical losses of every sub-components of the OFB down to 1.8 K;
The complete OFB assembly (only tack welded) was also tested down to 4.2 K at SM18 with optical losses slightly better that the MQXFBP2 OFB ;
An issue after complete welding was observed– Same issue was observed after dismounting operation of BP2 (handling operation) ;
The MQXFBP2 OFB was fully repaired and tested at RT;
Expected optical losses at cold should be fine for a successful measurement.
Additional developments :

## Slide 29
Copper coated fiber brazed (250 um) on a copper insert, brazed on a SS flange.
Already tested at mbars
Two coils of MQXFB02 are equipped with high reflectivity FBG’s to improve the optical power budget at cold.