## Slide 1: Surface contamination detection by FTIR and XRF spectroscopy
Benoit Teissandier – TE/VSC
Chemistry Lab. (10/R-015)
09/11/2022

## Slide 2
Analysis of organic compounds by infrared spectroscopy
e.g.: Silicone oil
Identification / characterisation of solid, liquid and gas
  Source of contamination (e.g. UHV applications, gas system for detector, etc.)
  Polymers, cooling fluids and some gas
  Emissivity measurement of metallic substrates
Quantification to ppm level (after calibration)
  Cleanliness assessment for UHV applications or gas system detector
  Cooling fluids purity
Lab. Spectrometer : Bruker Vertex 70
Handheld : Agilent Exoscan 4100
ZnS Window
Polychromatic IR light

## Slide 3
In situ sampling and FTIR identification
Sampling on CMS cryogenic turbine with PTFE membrane
Helium compressor oil
Breox B35
Sampling on LHC collimator tank with PTFE membrane
Oil from tungsten
machining

## Slide 4
Quantification and detection limit
Order of magnitude ≈ 1 molecular layer
Detection limit of Dimethylsiloxane on stainless steel surface
Référence Inox propre
2.3 µg/cm2
1.15 µg/cm2
0.46 µg/cm2
0.23 µg/cm2
0.115 µg/cm2
n  1235
EDMS N° : 1332721 & 347564
Detection limit of hydrocarbon contamination on stainless steel surface
  CERN standard contamination :
  Cutting oil: Blasocut BC 35 LF SW (33 % v/v)
  Machine oil: Shell Vitrea 150 (33 % v/v)
  Bearing grease : Kluber Isoflex NBU 15 (33 % v/v)
0.1 µg/cm2

## Slide 5
Others FTIR applications fields
Polyethylene terephthalate (PET)
Polyethylene (PE)
Polytetrafluoroethylene (PTFE)
Ethylene-vinyl acetate (EVA)
Identification of polymers in situ
Chart: University of Dundee - AISI 304L sample 15 (@ 27°C)
Emissivity of metallic surfaces

## Slide 6
Analysis of inorganic compounds by XRF spectroscopy
Niton XL3t GOLDD+

## Slide 7
In situ XRF analysis
Alloy discrimination
SS-304 or SS-316 ?
(presence of Mo in SS-316)
Discrimination in less than
2 seconds !
Chlorine detection in corrosion point on HL-LHC beam screen
[1] – Normal part
[2] – Corrosion part

## Slide 8
Differential Scanning Calorimetry
UV-Visible spectroscopy
Gas Chromatography
Karl Fischer Coulometry
Thermogravimetry
X-ray spectroscopy
FTIR spectroscopy
Optical Emission Spectroscopy
Atomic Absorption Spectroscopy
For more information :
CERN Chemistry Lab. (10/R-015)
Thanks for your attention !