## Slide 1: SPS-BGI design update

## Slide 2: Parts and assembly Cooling 90º adaptor
3D Model | Functionality | Material
ST1612952 | Changes the direction of the water flow | St. Steel EN 1.4404 EDMS 1380627 Tube
Diameter reduced, because of the collision with washers
Square variant as an option
Not much space for the O-ring
O-ring material to be discussed with supplier

## Slide 3: Parts and assembly Cooling 90º adaptor
3D Model | Functionality | Material
ST1612952 | Changes the direction of the water flow | St. Steel EN 1.4404 EDMS 1380627 Tube
Coupling does not fit to the assembly (stud will be used instead of the screws+nuts)
Need to find another solution

## Slide 4: Parts and assembly Rectangular flange assembly
3D Model | Functionality | Material
ST1588145 | Insures vacuum and contains cooling, power and data connectors | St. Steel EN 1.4429 EDMS 790775 Blank and EDMS 790773 Round Bar
Outgassing from inside?
Holes in arm
Pockets in the flange?
Flange updated with new electronics:
ST1637325: Ceramic Power Board
ST1637262: Flange Data Board
ST1637281: Flange Power Board

## Slide 5: Parts and assembly Support arm
3D Model | Functionality | Material
ST1572023 (L) ST1602416 (R) | Positioning of the ceramic insulators, the faraday cage, the bottom electrode, the front reinforcement, the RF fingers and the cooling plate. Is attached to the flange | St. Steel EN 1.4429 EDMS 790774 Sheet
Both arms re-designed
Inclined part removed
Back column stiffer
Feedback about radiuses expected
Feedback about RF fingers expected

## Slide 6: Parts and assembly Support arm
3D Model | Functionality | Material
ST1572023 (L) ST1602416 (R) | Positioning of the ceramic insulators, the faraday cage, the bottom electrode, the front reinforcement, the RF fingers and the cooling plate. Is attached to the flange | St. Steel EN 1.4429 EDMS 790774 Sheet
Do these holes need to cut through?
Fit to be solved (Arm, Electrode, Faraday cage)
Feedback about radiuses expected
Fit to be solved (Arm, reinforcement, Electrode)

## Slide 7: Parts and assembly Support arm
3D Model | Functionality | Material
ST1572023 (L) ST1602416 (R) | Positioning of the ceramic insulators, the faraday cage, the bottom electrode, the front reinforcement, the RF fingers and the cooling plate. Is attached to the flange | St. Steel EN 1.4429 EDMS 790774 Sheet
Which type of screw to be used?
Feedback expected

## Slide 8: Parts and assembly Support arm
3D Model | Functionality | Material
ST1572023 (L) ST1602416 (R) | Positioning of the ceramic insulators, the faraday cage, the bottom electrode, the front reinforcement, the RF fingers and the cooling plate. Is attached to the flange | St. Steel EN 1.4429 EDMS 790774 Sheet
Space is limited between the arms and the tank
Thickness of the arm is 6mm now – countersunk holes / counterbored holes?

## Slide 9: Parts and assembly Support arm
3D Model | Functionality | Material
ST1572023 (L) ST1602416 (R) | Positioning of the ceramic insulators, the faraday cage, the bottom electrode, the front reinforcement, the RF fingers and the cooling plate. Is attached to the flange | St. Steel EN 1.4429 EDMS 790774 Sheet
Thickness 20mm not ideal
Options:
Keep 20mm raw (not precise)
Reduce to 18/19mm (not ideal)
Enlarge width
  Use 30mm bulk (lot of machining)
  Use 50mm bulk (cut longitudinally)

## Slide 10: Parts and assembly Support arm
3D Model | Functionality | Material
ST1572023 (L) ST1602416 (R) | Positioning of the ceramic insulators, the faraday cage, the bottom electrode, the front reinforcement, the RF fingers and the cooling plate. Is attached to the flange | St. Steel EN 1.4429 EDMS 790774 Sheet
Threaded holes for arm dismantling
Threaded hole for Faraday cage attachment
Pins custom made or commercialized?
Which screws?

## Slide 11: Parts and assembly Cooling base plate
3D Model | Functionality | Material
ST1571284 | Positioning of the cooling pipe, the detector modules and the ceramic guide | Cu OFE C10100 (HO2)
Radiuses added
Plate was made thicker by Teresa
Added:
Contact surface with the arm
Venting hole from the bottom

## Slide 12: Parts and assembly Cooling base plate
3D Model | Functionality | Material
ST1571284 | Positioning of the cooling pipe, the detector modules and the ceramic guide | Cu OFE C10100 (HO2)
REMARKS:
The distance between the ceramic PCB (side facing the power cabling) and the power cabling NEEDS TO be 2.075 mm
The cooling base plate NEEDS TO ensure an accurate positioning of the TPX3 (pins + screws)
The detector assembly NEEDS TO be parallel to the beamline and centered
The slot for the cooling pipe NEEDS TO be precisely manufactured since it embraces the cooling pipe
Where are these Pins supposed to be?
No feature foreseen for it yet – idea shoulder screw
Careful with the assembly steps (centering and pipe welding)

## Slide 13: Parts and assembly Cooling base plate
3D Model | Functionality | Material
ST1571284 | Positioning of the cooling pipe, the detector modules and the ceramic guide | Cu OFE C10100 (HO2)
Think about positioning (pins/shoulder screws) and pipe welding
Slotted holes in cooling plate?

## Slide 14: Parts and assembly Front reinforcement
3D Model | Functionality | Material
ST1571802 | Gives rigidity to the structure | St. Steel EN 1.4429 EDMS 790774 Sheet
Side made thicker
Feedback about radiuses expected
Feedback about RF fingers expected

## Slide 15: Parts and assembly Ceramic cover and Ceramic guide
3D Model | Functionality | Material
ST1584545 | Power cabling electrical insulation | Ceramic Al2O3 96%
Inner radiuses will be added
First discussion with workshop started

## Slide 16: Parts and assembly Assembly detector
3D Model | Functionality | Material
ST1570691 | Integrates the TPX3, and power and data electronics | Module base: Alu EN AW-6082 (T6)
Is the screw connection accurate enough?
Add venting holes for the blind holes?
REMARKS:
The distance between the ceramic PCB (side facing the power cabling) and the power cabling NEEDS TO be 2.075 mm
The module base NEEDS TO ensure an accurate positioning of the TPX3
The detector assembly NEEDS TO be parallel to the beamline and centered

## Slide 17: Parts and assembly Bottom electrode
3D Model | Functionality | Material
ST1571600 | Connected to GND, is in charge of creating the electric field. Its opening also allows electrons detection | St. Steel EN 1.4429 EDMS 790774 Sheet
Which type of screws?
Countersunk?
To not influence the electrical field?

## Slide 18: Parts and assembly Faraday cage
3D Model | Functionality | Material
ST1596265 | Protects the electronic components from radiation and insulates them from the magnetic field | Cu OFE C10100 (HO2) St. Steel EN 1.4429 EDMS 790774 Sheet
Design changed to 90deg bent
Two additional screws in top part

## Slide 19: Parts and assembly Faraday cage
3D Model | Functionality | Material
ST1596265 | Protects the electronic components from radiation and insulates them from the magnetic field | Cu OFE C10100 (HO2) St. Steel EN 1.4429 EDMS 790774 Sheet
Outgassing? (pocket in bottom electrode/arm/grower)
Cut through the electrode and use longer screw with threaded hole in arm?

## Slide 20: Parts and assembly Electrical feedthrough
3D Model | Functionality | Material
ST1589217 | Conducts a maximum of 40 kV | 
Three solutions for the feedthrough
New ideas?
Feedback from Wilfried expected

## Slide 21: Parts and assembly Collet
3D Model | Functionality | Material
ST1589079 | Guarantees electrical contact between the feedthrough and the top electrode | St. Steel EN 1.4404 EDMS 1380627 Tube
New ideas?
Feedback from Wilfried expected

## Slide 22: Parts and assembly Assembly chamber
3D Model | Functionality | Material
ST1588183 | Chamber where the instrument will be installed in a vacuum environment | St. Steel EN 1.4429 and 1.4306
REMARKS:
The tank NEEDS TO fit between the magnet ribs
The open rectangular flange NEEDS TO fit between the magnet coils
The injection and pumping tubes NEEDS TO allow the magnet to slide
The feedthrough tube NEEDS TO be as long as necessary to allow the feedthrough insertion without modifying the pin
The cone chamber NEEDS TO accomplish RF standards
Tank
Conical flange
Injection tube
Open rectangular flange
Cone chamber
Pumping tube
Pumping chamber
Back tank
Feedthrough tube
Injection chamber

## Slide 23: Parts and assembly Assembly chamber
3D Model | Functionality | Material
ST1588183 | Chamber where the instrument will be installed in a vacuum environment | St. Steel EN 1.4429 and 1.4306
New connection of the back tank
All welded connections will be discussed with workshop