## Slide 1: SPS North Area SEM detectors and their absolute calibration
L. Parsons França , M.van Dijk, F. Ravotti, F. Roncarolo, C. P. Welsch, H. Zhang
BI Day 2022

## Slide 2: Outline
North Area Beamlines
SEM operating principle
Why Calibrate?
Calibration Steps
Experimental Setup m1
Activation measurement + calculations
Experimental Setup m2
Preliminary Results
Next steps
Background
Background
Calibration Steps
Preliminary Results
Next steps
North Area beamlines
SEM operating principle
Why calibrate?
Experimental setup 1
Activation measurement + calculations
Experimental setup 2
POT and spill

## Slide 3: Location
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 4: North Area Beamlines
Nearest Beam Current Transformer is in the SPS ring (before extraction)
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 5: North Area Beamlines
210
Many types of Secondary Emission Monitors (SEMs) are used in TT20:
BSM(H,V) – beam position
BSP(H,V) – beam position
BSG(H,V) – beam profile/position
BBS(H,V) – beam profile/position
BSI – beam intensity
Calibration Location
Calibration Steps
Background
Preliminary Results
Next steps
We will focus on BSI (beam intensity)

## Slide 6: Operating Principle
Beam Secondary emission Intensity (BSI)
Protons hit metal foils/bands transferring energy to the material
Energy transfer liberates electrons, generating a current, that can be measured
The number of electrons emitted per incident proton is known as the Secondary Emission Yield (SEY)
The number of ADC counts in the device is related to the number of protons via a calibration factor:
Calibration Steps
Background
Preliminary Results
Next steps
BSI (Al or Ti Plate)
e-
P+
BIAS
BIAS
Signal

## Slide 7: Protons on Target and Spill
Protons extracted from SPS ring split between different targets
7
BSIs at each target measure protons on target (POT)
Spill corresponds to one extraction period (~ 4.8s in NA)
Ideal spill measured by BSI is flat, but in practice there are fluctuations in the signal
BSI
Extraction
From SPS into NA
Current measured BCT
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 8: Why Calibrate?
Recent renewed interest in North Area
Fixed target experiments need accurate number of Protons On Target (POT) for reference
The calibration uncertainty is dominated by the SEY
Changes in the material over time could affect SEY (e.g. radiation and vacuum history)
Instrumentation in North Area has not been calibrated since the 90s
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 9: 2017 Data Highlighting Calibration Uncertainties
Slope = 0.82
Slope = 0.73
Non linearity possible causes:
• SEY drift in time
• Variation of losses with beam intensity.
Slope deviation possible causes:
• Systematic errors in the BSI calibration
• Absolute extraction losses
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 10: Damage to the BSI Foil
Upstream of Target box, minimal damage to foil
BSI Inside Target box has dip in intensity which coincides with centre of target
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 11: Challenges for Calibration
TT20 is a slow extracted beamline (beam not bunched)
Beam current transformers (BCTs ) not suitable
First order BSI calibration- compare to BCTs located in SPS ring (before de-bunching):
  Affected by knowledge of absolute losses between SPS and TT20
Use known calibration technique: “activation method”
  Performed in the past in North Area
  Used more recently in East Area
TT20
SPS
BCT
BSI
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 12: Calibration steps
Insert foil stack in beamline
Irradiate for short period
Remove foils
Cool down
Measure activity in foils
Use activation cross-section to find number of protons
Compare to current measured by BSI
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 13: Experimental Setup m1
Insert foil stack in beamline
Irradiate for short period
Remove foils
Cool down
Measure activity in foils
Use activation cross-section to find number of protons
Compare to current measured by BSI
f3
f2
f1
Beam
Gafchromic Films
Aluminium foils (100 μm thick) x 3
Gafchromic Films x2
Foils irradiated for ~ 30 mins
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 14: Activation Measurement 1
Foil 1 (most downstream) has significantly lower activity
Central foil to be used
From M. van Dijk
Thin film (beamspot only)
Thick film (full area)
Na-24 peak
Measure activity in foils
Remove foils
Cool down
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 15: POT Calculations
No. particles per second traversing foil
Activity
Surface Atomic density
Cross-section
Decay constant
Irradation time
Time elapsed since irradiation
Use activation cross-section to find number of protons
Measure activity in foils
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 16: Considerations and Corrections
Assessment of material composition (99.5% Al): impurities negligible
Interactions in material generate neutrons and other charged particles
Generates additional activation
Scales with thickness of foil material
Two values available for 24-Na cross-section in aluminium at high energy
CERN @ 400 GeV 8.79 +/- 0.18 mb (corrected for extra activation)
Fermilab @ 300 GeV 8.15+/-0.59 mb (corrected to match energy)
From M. van Dijk
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 17: BSI Measurement
Mean of 109 spills
ADC counts
Time (s)
Step 1: Add all adc counts for each spill
Step 2: Scale by the gain (400V)
Step 3: Total counts = 109 spill integrals added
Use activation cross-section to find number of protons
Compare to current measured by BSI
Calibration Steps
Background
Preliminary Results
Next steps
Spill integral

## Slide 18: Experimental Setup 2
Aluminium foils - 99.999% purity (100 μm thick) x 3
Copper foils - 99.99% purity (100 μm thick) x 3
Gafchromic Films x2
Foils irradiated for ~ 30 mins
Gafchromic Films
Al Foils
Cu Foils
Plastic spacer

## Slide 19: Measurement 2 Data Analysis
Cooldown period was much longer (~5 days)
Many more isotopes were found in gamma spectrometry analysis
Calibration Steps
Background
Preliminary Results
Next steps

## Slide 20: Preliminary Results
Current cf= 1.3e9
Cf = 99.1 +/- 1.8%
Plot from M. van Dijk
Calibration Steps
Background
Preliminary Results
Next steps
Calibraction Factor
Protons measured
With activation method
Protons measured
with BSI

## Slide 21: Conclusion and next steps
Data analysis is still ongoing
We will compare our results to each other to check for systematic errors
We aim to:
  Repeat the measurements in other target locations
  Propose new BSI monitor design optimised for this kind of calibration  ~easy (and ALARA ‘friendly’) installation and removal of activation foil close to BSI foil
We plan to standardise the procedure to make it repeatable for regular implementation in these beamlines
Background
Preliminary Results
Next steps
Calibration Steps

## Slide 22
Questions

## Slide 23: Estimation of statistical error on BSI
Deviation from mean for integrated adc counts
variation on individual measurement ~2%
But this includes variation of the beam intensity
Check for correlation between BSI and BCT measurements
Apply correlation correction to bsi data
new variation on BSI measurement: 1.2%
Error on total adc counts = error on the mean (0.11%)
Error %
2 Plots from M. van Dijk