## Slide 1: Rad-Tol COTSs Qualification for LHC BPM CONS
Selim Can Ozdogan on behalf of SY-BI-BP
16 Dec 2022

## Slide 2: Rad-Tol COTSs Qualification for LHC BPM CONS
Agenda
  Introduction & Motivation
  LHC BPM Consolidation project
  Radiation qualification of Commercial Off The Shelf (COTS) components
  Summary

## Slide 3: Rad-Tol COTSs Qualification for LHC BPM CONS
Agenda
  Introduction & Motivation
  LHC BPM Consolidation project
  Radiation qualification of Commercial Off The Shelf (COTS) components
  Summary

## Slide 4: Introduction & Motivation
Current LHC BPM system overview
~1250 BPMs
  Located at LHC ring and TI2 & TI8
  Button & Stripline BPM pick-ups
  Dual plane (H & V)
Front-End (FE) readout based on analog electronics
Back-End (BE) based on Digital Acquisition Board (DAB) from BI
Why is a consolidation required?
System operating for ~20 years
  Aging of electronics -> performance and reliability issues
  Limited spares -> long run maintenance compromised
  Complicated calibration procedures
Surface
Tunnel
Tunnel installation
BE & FE readout electronics

## Slide 5: Rad-Tol COTSs Qualification for LHC BPM CONS
Agenda
  Introduction & Motivation
  LHC BPM Consolidation project
  Radiation qualification of Commercial Off The Shelf (COTS) components
  Summary

## Slide 6: LHC BPM Consolidation Project – Radiation Environment
Increased luminosity in HL-LHC operation Higher radiation levels
Values extracted from the “Radiation level specification document for HL-LHC” (EDMS 2302154)
Expected annual (360 fb-1) HL-LHC radiation levels |  |  |  | 
 | TID (Gy) | HEH (cm-2) | Th. Neut. (cm-2) | 1MeVn-eq (cm-2)
Arcs | 1.4 | 2.4 · 109 | 1.2 · 1010 | 1.6 · 1010
RRs | 25 (max.) | 1.4 · 1010 (max.) | 1.2 · 1011(max.) | 7 · 1010 (max.)
Values extracted from “CERN-ACC-NOTE-2018-088 “ and calculated with EDMS 2302154
Current annual (360 fb-1) LHC radiation levels |  |  |  | 
 | TID (Gy) | HEH (cm-2) | Th. Neut. (cm-2) | 1MeVn-eq (cm-2)
Arcs | 0.05 | 5.0 · 107 | 5.0 · 108 | 5.0 · 108
RRs | 4.8 (max.) | 4.8 · 109 (max.) | 4.8· 1010(max.) | 4.8 · 1010 (max.)

## Slide 7
LHC BPM Consolidation Project – Architecture
Reuse of current infrastructure
  BPMs pick-ups
  Cryo-cables, Coaxial cables & Optical fibres
  Slow control fieldbus (WorldFIP)
Front-End
  Exposed to radiation
  Slow control interface: NanoFIP FMC (developed by BE-CEM)
  Crate shared with BLM (developed by SY-BI-BL)
Optical link for data read-out (in collaboration with EP-ESE)
Back-End TBD
Surface
Tunnel
BE & FE readout electronics
BLM side
BPM side
3D model courtesy of William Vigano (SY-BI-BL)
Tunnel installation
BE
FE
Cryo-cables
BPM pick-ups

## Slide 8
  BPM pick-ups as current system
  Analog Front-End TBD
  Digital Front-End:
    ADC (4 x 12-bit @ 1.28GSps)
    Clock Scheme: Crystal Oscillator & Clock Synthesizer
    Optical link: 4 x 10 Gbps per BPM
LHC BPM Consolidation Project – Front-End readout block diagram
Digital Front-End
BPM1
BPM2
DELAY
DELAY
LPF
ATT.
AMP.
ADC
ADC
ADC
ADC
Clock
Synth
Optical
MUX
Optical
MUX
Analog Front-End
DELAY
DELAY
LPF
ATT.
AMP.
LPF
ATT.
AMP.
LPF
ATT.
AMP.
Optical Transmitter
Optical Transmitter
Optical Transmitter
Optical Transmitter
Optical Transmitter
Optical Transmitter
Optical Transmitter
Optical Transmitter
Osc.
Surface
Tunnel
All active components must be radiation tolerant.

## Slide 9
LHC BPM Consolidation Project – Front-End main components
Front-End components selection | Front-End components selection |  | 
Component Type | Component Type | Number of candidates | Comments
ADC | ADC | 2 | One preferred candidate from Teledyne
Clocks scheme | Crystal osc. | 3 | Selection based on ADC requirements
 | Clock synth. | 3 | 
Optical link | CERN custom (VTTx) | 1 | Primary option RadHard by design
 | COTS Transceivers | 8 | Secondary option
Other COTS | Other COTS | - | TBD

## Slide 10
A
B
C
D
D: Other COTSs
2020
A: Clock scheme
B: Optical link
C: ADC
LHC BPM Consolidation Project – Radiation Hardness Assurance (RHA) timeline
2022
2023
E: Front-End
2024
2025
E

## Slide 11: Rad-Tol COTSs Qualification for LHC BPM CONS
Agenda
  Introduction & Motivation
  LHC BPM Consolidation project
  Radiation qualification of Commercial Off The Shelf (COTS) components
  Summary

## Slide 12
Cumulative Effects
Total Ionizing Dose (TID)
Displacement Damage (DD)
Single Event Effects (SEE)
Soft Effects
Hard Effects
Single Event Functional Interrupt
(SEFI)
Single Event Transient
(SET)
Single Event Upset
(SEU)
Single Event Gate Rupture
(SEGR)
Single Event BurnOut
(SEBO)
Single Event Latch-up
(SEL)
Lifecycle of all types of components
Bipolar
&
Opto elec.
TID unit:
Gray (Gy)
DD unit:
1MeVn-eq
Radiation qualification of COTS components – Radiation effects to electronics
SEE main figure:
https://www.famars.unibas.ch/singanim.htm
https://www.famars.unibas.ch/lunpanim.htm

## Slide 13
Crystal Oscillator
3 candidates selected
  Frequency:100MHz-200MHz
  Output Type: LVDS
  RMS Jitter: fs regime
1 candidate tested (6 samples)
  Radiation Test campaign carried out @PSI
    total ionizing dose (TID): 500 Gy
    proton fluence: 8.60x1011 p/cm-2
Monitored Parameters:
  Output Clock Frequency
  Duty Cycle
  Rise/Fall Times
  Phase Noise
  Current Consumption
Power Analyzer
Frequency Counter
Frequency Counter
Frequency Counter
Spectrum Analyzer
Radiation qualification of COTS components – Clocks Scheme
DUT1
DUT2
DUT3
DUT4
DUT5
DUT6
No significant degradation of the monitored parameters during the test
Waiting for radiation test report

## Slide 14
Radiation qualification of COTS components – Clocks Scheme
Irradiation test setup of crystal oscillator at Paul Scherrer Institute (PSI)

## Slide 15
Figure: DUT Setup
Spectrum Analyzer
Oscilloscope
Oscilloscope
Power Analyzer
Clock Synthesizer
1 sample of 3 candidates
  LMX2694-EP
  LMX2594-EP
  LMX2592-EP
Radiation Test campaign carried out @PSI.
  total ionizing dose (TID): 500 Gy
  proton fluence: 8.60x1011 p/cm-2.
Monitored Parameters:
  Output Clock Frequency
  Duty Cycle
  Rise/Fall Times
  Phase Noise
  Current Consumption
  Register Status
Radiation qualification of COTS components – Clocks Scheme
DUT | TID (Gy) |  |  |  |  |  | 
LMX2694-EP | 500 | <1.16E-12 | 0.05568 | 0.3248 | <1.40E-12 | 0.0672 | 0.392
DUT | TID (Gy) |  |  |  |  |  | 
LMX2694-EP | 500 | <1.16E-12 | 0.05568 | 0.3248 | <1.40E-12 | 0.0672 | 0.392
Values extracted from the radiation test report (EDMS 2302154)

## Slide 16
Radiation qualification of COTS components – Clocks Scheme
Irradiation test setup of clock synthesizer at Paul Scherrer Institute (PSI)

## Slide 17
COTS Optical Transceivers
4 Different type of DUTs
  Generic components
  All capable running @10Gbps
Radiation Test campaign carried out @PSI
  total ionizing dose (TID): 500 Gy
  proton fluence: 8.60x1011 p/cm-2
Monitored Parameters:
  Offline Monitoring of:
    Tx Optical Power
  Real-time Monitoring of:
    Bit Error Rate (BER)
    Loss of Lock
    Register Status
    Current Consumption
Radiation qualification of COTS components – Optical Link
17
Ethernet
DevKit
Optical Transceiver
Test Board
DUT
Register Status
Monitor
Xilinx FPGA
QSFP
QSFP
4 x SFP
4 x SFP
Slow Control
(I2C)
Tx
Tx
Tx
Tx
Rx
Rx
Rx
Rx
QSFP
Tx
Tx
Rx
Rx
QSFP
Tx
Tx
Rx
Rx
DUT
DUT RX BER
Monitor
DUT TX BER
Monitor
Type | Nb Channels | Line Rate per Channel (Gbps) | Nb candidates
SFP+ | 1 | 10 | 2
SFP28 | 1 | up to 25 | 2
QSFP+ | 4 | 10 | 1
QSFP28 | 4 | up to 25 | 3
4 out of 8 candidates showed promising results
Waiting for radiation test report

## Slide 18
LHC BPM Consolidation - Radiation Qualification of COTS– Performed Activities
DUT
Irradiation test setup of optical transceivers at Paul Scherrer Institute (PSI)
Test concept verification setup at lab

## Slide 19
ADC
Two candidates to be tested
  Candidate 1: EV12AQ600 from Teledyne
    Rad-tol by design (tested with heavy ions)
  Candidate 2: TBD
Other COTSs
Analog Front-End COTSs selection still ongoing
Front-End powering
  PSU in collaboration with SY-BI-BL
  Selection of other components in collaboration with
    BE-CEM
    RADWG
Radiation qualification of COTS components – ADC & Other COTSs
ADC & Optical Transmission
Test Board
4 x Optical
Transceiver / Transmitter

## Slide 20: Rad-Tol COTSs Qualification for LHC BPM CONS
Agenda
  Introduction & Motivation
  LHC BPM Consolidation project
  Radiation qualification of Commercial Off The Shelf (COTS) components
  Summary

## Slide 21
Summary
Current LHC BPM system requires consolidation
Consolidation project is ongoing in collaboration with different teams
Radiation test required for Front-End components
Digital Front-End
BPM1
BPM2
DELAY
DELAY
LPF
ATT.
AMP.
ADC
ADC
ADC
ADC
Clock
Synth
Optical
MUX
Optical
MUX
Analog Front-End
DELAY
DELAY
LPF
ATT.
AMP.
LPF
ATT.
AMP.
LPF
ATT.
AMP.
Optical Transmitter
Optical Transmitter
Optical Transmitter
Optical Transmitter
Optical Transmitter
Optical Transmitter
Optical Transmitter
Optical Transmitter
Osc.
Surface
Tunnel
1 candidate tested & showed promising results
2 more to be tested
3 candidates tested
1 candidate passed the test
8 candidates tested
4 candidates with promising results
2 candidates to be tested in 2023
Specifications to be finalized.
Tests foreseen in 2024

## Slide 22
Thanks to
Manoel Barros Marin from SY-BI-BP
&
RADWG:
Salvatore Danzeca
Rudy Ferraro
Antonio Scialdone