## Slide 2: CERN Vacuum acceptance tests for particle accelerator equipment: An Overview
Giuseppe Bregliozzi
TE-VSC-BVO
VELO Upgrade II Workshop
Amsterdam, 6 – 8 Feb 2023

## Slide 3
Outline
Why vacuum acceptance tests?
Acceptance tests procedures:
  Unbaked system
    Example of polymeric component
  Baked system
Conclusions & Advices
Additional slides: Some special needs and examples:
  Beam induced effects: ESD & PSD cases
  New sintered materials
  Glues vs. NEG

## Slide 4
Why Vacuum Acceptance Tests?
One of the main mandates of the CERN Vacuum Surfaces Coatings group is to provide the beam operation with a required vacuum level on all the accelerator complex.
To achieve that mandate, acceptance tests are needed to assess the compatibility of all pieces of equipment to be installed in the beam vacuum system of the accelerator complex:
    Leak tightness.
    Detection of contamination.
    Measurement of outgassing rate and its time variation.
    Measurement of virtual leaks (in leakage).

## Slide 5
CERN accelerators chain
26 GeV
450 GeV
CMS
ATLAS
LHCb
50 MeV
1-1.4 GeV
450 GeV
2 X 7 TeV
ALICE

## Slide 6
CERN accelerators chain: vacuum systems and requirements
<10-8 mbar
<10-12 mbar
LEIR
LHC Arcs
LHC straight sections <10-10 mbar
UNBAKED SYSTEMS (SEPTA..):
TMP, ION PUMPS;
SUBLIMATORS;
BAKED SYSTEMS:
ION PUMPS;
Non Evaporable Getter (TiZrV) coating
CRYO SYSTEMS:
CRYOPUMPING;
ELENA ring<4*10-12 mbar
Linac2
<2*10-6 mbar*
PSB
<5*10-8 mbar*
PS
<2*10-8 mbar*
SPS LSS <10-7 mbar*

## Slide 7
admissible outgassing rate
DRIVEN BY Beam-gas scattering: Beam losses and beam lifetime
  Estimation of average pumping speed to define a total admissible outgassing rate;
  Determine admissible molecules density;
DRIVEN BY Beam downtime: Allowed time to restart the machine in case of components exchange.
DRIVEN BY Equipment requirement: Maximum allowed pressure/molecules density to run devices like kickers or RF cavities.
How to define the acceptance criteria?
No Contamination
Anomalous presence of hydrocarbons, most probably due to error in design and/or lack of appropriate cleaning (error in cleaning procedure or post-cleaning pollution); inappropriate choice of materials (polymers, glues, lubricants …);
Higher than expected CO and CO2 outgassing indicating the presence of carbonised elements;
Any chemical element or compound usually not present in the residual gas phase, for example, F and Cl (issue with etching and cleaning), K and Na (manipulation), P and S (issue with electrolytic treatments).

## Slide 8
Some examples: Baked system
Area | Equivalent Hydrogen gas density | Effective pumping speed (indicative)
Arcs | ≤10 +15 H2 m-3 | ≥100 l.s-1
Experiments | ≤10 +13 H2 m-3 | 
LHC
ELENA ACCEPTANCE THRESHOLDS
Ensures the limitation of momentum and emittance blow up induced by the interaction of 100 keV antiprotons with a beam population of 107. No specification on gas composition
ELENA
Area | Pressure requirements | Effective pumping speed (indicative)
Ring | ≤4 x10-12 mbar | Depend upon position (NEG sticking probability)
Transfer lines | ≤10-10 mbar | 
Dynamic effects included!
LHC ACCEPTANCE THRESHOLDS
Ensure 100 h of circulating beams before intensity degradation due to residual gas interactions occurs and minimise the background for the experiments.
In both examples, the pressure requirements are driven by beam requirements.
LHC Design Report
CERN-2004-003
4 June 2004
Courtesy: C.Pasquino – OLAV - V

## Slide 9
Some examples: Unbaked system
Accelerator | Area | Operational pressure requirement (24h pumping) [mbar] | Average effective pumping speed (indicative) [l s-1] | Outgassing rate limit at 24 h [mbar.l.s-1]
PS complex | LINACS AND TRANSFER LINES CLOSE TO PSB AND PS | ≤2.10-6 | 100 | 5∙10-5 (*)
 | PSB AND TRANSFER LINES CLOSE TO THE PS RING | ≤5.10-8 | 100 | 5.10-6 (*)
 | PS ring | ≤2.10-8 | 70 | 1.5.10-6 (*)
SPS | Arcs | ≤10-6 | 10 | 10-5
 | LSS (kickers, septa, RF cavities) | ≤10-7 | 100 | 10-5 (*)
 | TI2&TI8: From SPS to TED | ≤10-5 | 1-2 | 2.10-5
 | TI2&TI8: From TED to LHC | ≤5.10-7 | 5 | 2.5.10-6
ACCEPTANCE THRESHOLDS
LHC injectors:
  Ensure beam operation after 24h pump down;
  Ensure proper functioning of High Voltage and RF devices;
  Ensure ion operation at low energy (e.g. LEIR, PS);
SPS ARCs PUMPDOWN
In this case, the acceptance thresholds are beam downtime and equipment driven.
Courtesy: C.Pasquino – OLAV - V

## Slide 10: Acceptance Criteria
VACUUM
COMPONENT
CONTAMINATION
OUTGASSING RATE
VIRTUAL LEAKS
LEAK TIGHTNESS
Chart: Time [sec]
Courtesy: C. Pasquino – OLAV -V

## Slide 11: Unbaked: Vacuum validation steps
Measurement and verification of vacuum performance
    Functionality
    Leak tightness (First: high background)
    Outgassing rate after 24h of pump down
    Residual Gas Analysis
    Leak tightness (Final)

## Slide 12: Unbaked: Acceptance Thresholds Overview
Atomic mass units from 44 to 100 (Indication of organic contamination)
Atomic mass units from 18 to 44
He leak rate:
    QAIR< 10-10 mbar∙l/s
Outgassing rate after 24h of pump down:
    QOut < Depends on the machine
Gas composition: H2O is the dominant peak
All the masses between 18 et 44 are at least 100 lower of the intensity of peak 18 (2 order of magnitude lower) except for masses 28 et 44
All the masses from 44 to 100 are at least 1000 lower of the intensity of peak 18 (3 order of magnitude lower) except for mass 44

## Slide 13: Acceptance tests for unbaked components
Pump down curve: outgassing rate value at 24h in N2 equivalent to be compared with the acceptance criteria for each specific accelerator and transfer line.
KEY FACTORS:
The chosen materials (polymers, sintered materials, high vapour pressures materials);
The thermal treatments (pre-baked in air or under vacuum (cladded cables));
The mechanical design: trapped volumes, leak tightness.
Courtesy: C. Pasquino – OLAV -V

## Slide 14: Is it 24h of pump down enough?
24h represent a good compromise that allows performing test within 1 week of time: installation, first leak detection, air venting overnight and final test. However…………..
Example of insulated cable with polyimide film of 0.2mm thickness
Exponential decay
In presence of polymeric materials, the thickness has an important role on their outgassing rate behaviour and should be addressed in advance

## Slide 15: Polymers in vacuum technology
Polymers are sometimes the only option as seal/insulator
PEEK, Kapton and Vespel -> Limited bakeout temperature
Guarantee a certain beam lifetime or certain operation conditions
Outgassing limit (maximum pressure to be reached in 24 hours) is defined for each machine AND the residual gas analysis free of contaminants
Acceptance test prior to installation:
- Pumpdown will define the outgassing rate and variation in time
- Leak detection
- Outgassing @ 24h
- RGA scan
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi

## Slide 16: Polymers in vacuum technology
Pumpdown LINAC4 Faraday Cup
LINAC 4 Faraday Cup
Vespel & Kapton wires
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi

## Slide 17: Outgassing of water Metallic surface vs Polymer
Outgassing
Molecules which have previously
been adsorbed (venting),that
desorb again
Molecules diffusing through the bulk
material, entering the surface and
desorbing from it
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi

## Slide 18: Outgassing of water Metallic surface vs Polymer
Molecules which have previously
been adsorbed (venting),that
desorb again
Chart: Elapsed Time [hours]
Ivo Wevers
Outgassing
Molecules diffusing through the bulk
material, entering the surface and
desorbing from it
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi

## Slide 19: Outgassing of water Metallic surface vs Polymer
Molecules which have previously
been adsorbed (venting),that
desorb again
Chart: Elapsed Time [hours]
Outgassing
Molecules diffusing through the bulk
material, entering the surface and
desorbing from it
Chart: Elapsed Time [hours]
| 0.11 | 0.2 | 0.4 | 0.8 | 1.1000000000000001 | 2 | 4
 | 2.2069e-05 | 1.6367e-05 | 1.1573e-05 | 8.1835e-06 | 6.9789e-06 | 5.1757e-06 | 3.6598e-06
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi

## Slide 20: Outgassing of water Metallic surface vs Polymer
Molecules which have previously
been adsorbed (venting),that
desorb again
Chart: Elapsed Time [hours]
Outgassing
Molecules diffusing through the bulk
material, entering the surface and
desorbing from it
Chart: Elapsed Time [hours]
| 0.11 | 0.2 | 0.4 | 0.8 | 1.1000000000000001 | 2 | 4 | 8 | 11 | 20 | 40 | 65 | 110 | 200 | 275 | 650 | 1100
 | 2.2069e-05 | 1.6367e-05 | 1.1573e-05 | 8.1835e-06 | 6.9789e-06 | 4.725e-06 | 2.1611e-06 | 4.521e-07 | 1.3984e-07 | 4.139e-09 | 1.6583e-12 | 9.3999e-17 | 2.1348e-24 | 1.1011e-39 | 2.0054e-52 | 4.0192e-116 | 1.4671e-192
 | 2.2069e-05 | 1.6367e-05 | 1.1573e-05 | 8.1835e-06 | 6.9789e-06 | 4.725e-06 | 2.1611e-06 | 4.521e-07 | 1.3984e-07 | 4.139e-09 | 1.6583e-12 | 9.3999e-17 | 2.1348e-24 | 1.1011e-39 | 2.0054e-52 | 4.0192e-116 | 1.4671e-192
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi

## Slide 21: Some example: Kapton
Kapton HN
Polyimide - Melting temperature: none
Expected water content: 0.96%
Samples: square sheets side 100 mm
Thicknesses: 0.0125-0.125 mm
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi

## Slide 22: Some example: PEEK
PEEK 450G
PolyEther-Ether Ketone
Semi-crystalline thermoplastic
Melting temperature: 343°C
Expected water content: ca. 0.25%
Samples shape: disks ∅ 51 mm
Thicknesses: 0.16-1.28 mm
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi

## Slide 23: Residual gas analysis Some examples
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi
Kapton
PEEK
Mylar
Vespel

## Slide 24: Outgassing measurements of PEEK, Kapton and Vespel samples
WATER
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi
Kapton
PEEK
Mylar
Vespel

## Slide 25: Outgassing measurements of PEEK, Kapton and Vespel samples
AIR
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi
Kapton
PEEK
Mylar
Vespel

## Slide 26: Cleaning methods of polymers

## Slide 27: Estimate of diffusion coefficient, moisture content and decay time
The time constant characterizes the decay of the outgassing rates
.
It depends on thickness and diffusion coefficient
The thickness has the greatest impact and can change pumping times by hundreds of hours
Low thickness
=
Much shorter !
VELO Upgrade II Workshop - Amsterdam, 6 – 8 Feb 2023 – G.Bregliozzi
Notes: 𝜏=ℎ^2/(𝜋^2 𝐷)

## Slide 28: Ongoing studies: Vacuum compatibility
Pixel detector for BGI
LCP
LCP
Kapton board
LCP Connector
Kapton cables
LISA Pathfinders Cables

## Slide 29: Acceptance tests for unbaked components
Residual Gas Analysis: after few hours of filament conditioning, the amplitude of the peaks is compared to the water content in the system. The component is considered accepted if the ratio between the water peak and the peaks up to mass 44 is higher than 100 and if the same ratio is higher then 1000 for peaks above mass 44.
KEY FACTOR
Detect the presence of hydrocarbons, CO, CO2, P, S, F, Cl or any unusual peak.
POSSIBLE SOURCES?
NONCONFORMITY IN THE DESIGN (lubricants under vacuum, glues…);
NONCONFORMITY IN THE CLEANING PROCEDURES;
NON-CONFORMITIES IN THE HANDLING PROCEDURES;
He leak rate:
    QAIR< 10-10 mbar∙l/s
x 100
x 103
Courtesy: C. Pasquino – OLAV - V

## Slide 30: Example of detected contaminants

## Slide 31: Example of contaminants Cross checking the roots cause

## Slide 32: Example of contaminants Cross checking the roots cause

## Slide 33: Baked: Vacuum validation steps
Measurement and verification of vacuum performance
    Functionality
    Leak tightness
    Outgassing rate
    Residual Gas Analysis
    Leak tightness
    Functionality
Before bake out cycle
After bake out cycle

## Slide 34: Baked: Acceptance Thresholds Overview LHC Case
Atomic mass units from 44 to 100 (Indication of organic contamination)
Acceptance criterion: RGA signals for all masses higher than 44 are at least 10000 times lower than the signal of peak H2 (mass 2).
Atomic mass units from 18 to 44 (Possible impact on NEG performance)
Different acceptance thresholds are selected as a function of the gas.
He leak rate:
    QAIR< 10-10 mbar∙l/s
Internal leak rate:
        QAIR< 5∙10-9 mbar∙l/s
Outgassing rate:
    QOut < 1∙10-7 mbar∙l/s
Gas composition: H2 is the dominant peak

## Slide 35: Baked: Acceptance Thresholds Overview LHC Case
    Maximum total outgassing ≤1∙10-7 mbar∙l/s
H2
CH4
CO
CO2
Affect the saturation level of NEG coating
    RGA Scan normalized to H2

## Slide 36: Conclusion & Advices
We need to clearly analyze and define which is your driving parameter and then set the acceptance criteria;
We should define gas density more than general pressure;
We could be flexible on the total outgassing, but we can not accept any form of contaminations;
We should be part as much as possible to the design phase to eliminates problems and nonconformities at the source.
Thank you for your attention

## Slide 37: Additional Slides: Some special needs and examples

## Slide 38: What about beam induced pressure increase?
How to deal with it?
An example in case of electron and photon flux

## Slide 39: Beam induced pressure increase
ATALS Interaction point
Courtesy: C.Y.Vallgreen

## Slide 40
ATLAS:
Dipoles & Quadrupoles
2808b , 6.5 TeV, Flat Top
100× zoom in radial direction
Q6
Q5
Q4
D2
Q7
MB
MB
D2
Q4
Q5
Q6
Q7
MB
MB
D1
IT
IP
Arc
1 Wm-2
10-7 Wm-2
10--5 Wm-2
10--3 Wm-2
SYNchrotron RADiation: SynRad
Definition of distributed photon flux
Courtesy: C.Y.Vallgreen

## Slide 41: Material characterization: hph
Courtesy: M.Ady et Al. IPAC15

## Slide 42: Definition of the SEY Threshold & Electron Flux
Without Field:
High SEY treshold
Dipole Field:
Concentrated electron flux
Quadrupole Field:
Trapping effect, area dependent on magnetic field
Example of LHC Beam Screen
SEY Threshold
Allow the definition of the minimum acceptable SEY for each materials
Electron flux distribution
Allow to determine the impact of the electron flux on the beam pipe surface
Courtesy: C.Y.Vallgreen

## Slide 43: Material characterization
SEY vs Primary energy:
The SEY scan allow determining the secondary electron evolution function of primary energy;
Accumulating electron dose bombardment it reduces the SEY and allow to create the ‘conditioning’ curve;
Conditioning Curve:
Allow determining and the SEY evolution and study its impact function of the simulated SEY threshold and electron flux.
Courtesy: R.Salemme & M.Taborelli

## Slide 44: Material characterization: hel
Courtesy: V.Baglin

## Slide 45
Synchrotron Radiation & Electron Cloud
Dynamic Pressure Profile: VAcuum Stability COde (VASCO)
ATLAS Long Straight Section
Synchrotron Radiation
Electron Cloud
Dynamic pressure profile
Courtesy: C. Y. Vallgreen