## Slide 1
By Nicolas GUILLOTIN – TE/CRG/ML
2023 JANUARY 18
2023 SM18 CRYOGENICS
SCHEDULE – CAPACITY – FLEXIBILITY – LIMITATIONS
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX
https://indico.cern.ch/event/1232235/

## Slide 2
2023 - SM18 MAIN REQUESTS
NON HL-LHC MAGNET
NON HL-LHC CAVITIES & CM
HL-LHC MAGNET
HL-LHC CAVITIES & CM
1
N. Guillotin
Maintenance,
Installation & connection of
the 1.9 K equipment for
the IT String infrastructure
FORECAST WITH SMOOTH COORDINATION
Sys. Admin days
Sys-Admin days with impact on
SM18 cryo equipment until January 9th
Restart of cryo infra equipment January 9th
Reconnection to main dewar with LHe on January 12ve
Pre-cool down to 80K with CWU can be anticipated on January 11th end afternoon
EN-EL AUG tests
EN-CV switch to the water bypass with LHC-P18 operation in parallel
Back to SFA-181 water tower on March 13
EN-EL Normal Secours tests on January 11th in the morning
EN-CV maintenance
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th

## Slide 3
2023 - SM18 MAIN REQUESTS
NON HL-LHC MAGNET
NON HL-LHC CAVITIES & CM
HL-LHC MAGNET
HL-LHC CAVITIES & CM
Dense alternation of Crab RF cavity tests in V3 and V4
FORECAST WITH SMOOTH COORDINATION
Several test benches commissioning : A2, F1, F2-SClinK
Dense alternation of HFM and Cluster_D tests  high number of quenches and high liquid helium consumption
2
N. Guillotin
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th

## Slide 4
2023 - SM18 MAIN REQUESTS
NON HL-LHC MAGNET
NON HL-LHC CAVITIES & CM
HL-LHC MAGNET
HL-LHC CAVITIES & CM
Dense alternation of tests between HFM and Cluster_D :
Cool down can not be performed simultaneously and
Quench recoveries can not be done in parallel
Regular crab RF cavity tests
Several test benches commissioning : M7 (SY-RF), B2, C2 (TE-MSC)
3
N. Guillotin
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th

## Slide 5
Chart: Liquid helium consumoption in g/s
g/s
LHE CONSUMPTION 2023
27g/s of pure liquefaction
35g/s of pure liquefaction
MAINTENANCE
4
N. Guillotin
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th

## Slide 6
Chart: 1.9K pumping consumption in g/s
2K PUMPING CONSUMPTION 2023
g/s
24 g/s with 2 WPU
@20 mBar
MAINTENANCE
FORECAST WITH SMOOTH COORDINATION – NO MARGIN
5
N. Guillotin
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th

## Slide 7
10 horizontal magnet
test benches
6 vertical magnet test benches
4 vertical RF-cavity cryostats
2 bunkers for RF-cryomodules
Typical performances :
6 g/s @ 10 mbar
12 g/s @ 20 mbar
Cold Box_B :
Liquefier in operation
Cold Box_A :
Back up liquefier
and for IT STRING
2 CWU :
Cold/warm up units,
with 3 Kaesers compressors (CWS)
2 WPU
Warm pumping units,
with very low pressure compressors
80g/s for each CWU
120kW @80K for each CWU
SM18 CRYOGENICS INFRASTRUCTURE
6
N. Guillotin
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th
CWU are LN2 cooled units

## Slide 8
Dewar
25’000L
Existing 6 kW
Refrigerator
[CB-A]
New 35 g/s
Liquefier
[CB-B]
RF Cavities &
Cryomodules
Horizontal Test Benches,
Vertical Test Benches,
F2-SClink
Cryo Distribution
USERS Supply & Return
Cold BoxES
LHe Supply & Return
Cold Gas Return
LHe Delivery
User Supply
LHe Delivery
New TLs
Auxiliary Valve Box
+ Purge
+ Balloon
CompressorS Station
Existing Sulzer process
Compressor Station
New Helium process
Compressor Station
HL-LHC IT STRING
Distribution Valve Box
Main
Valve Box
SM18 LIQUID HELIUM DISTRIBUTION infrastructure LAYOUT
2K pumping
units
WPU #2
WPU #1
Warm Pumping Units
10’000L
CCU
7
N. Guillotin
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th
25’000L Dewar reconnection foreseen after the next shutdown

## Slide 9
Chart: 
2015
2016
2017
2018
2019
2020
2021
MONTHLY CRYO AVAILABILITY (%) OVER THE LAST 8 YEARS
MAINTENANCE
MAINTENANCE
MAINTENANCE
MAINTENANCE
MAINTENANCE
COVID’19
MAINTENANCE
25kL-Dewar
Inlet clogging
TU1 filter
clogging
(25 g/s CB)
CP2 blocking & 3.3kV
motor alteration
during Open Days
-18kV Electrical cut-
25kL-Dewar pit’s flooding
Valve box insulation vacuum
2022
99.9%
81.2%
96.8%
95.8%
99.4%
99.3%
99.4%
97.3%
Yrly
availability
8
N. Guillotin
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th

## Slide 10
 | 2020 With 6 kW CB CB-A + 25kL Dewar | 2021 With new liquefier CB-B + 10kL DewAR | 2022 With new liquefier CB-B + 10kL DEWAR
Number of days of CRYO operation at SM18 | 274 days COVID’19 | 262 days 25kL-Dewar pit’s flooding | 310 days
Yearly Average Liquid Helium consumption by the test benches in g/s | 18.0 g/s | 19.8 g/s | 14.1 g/s CryoOK but no tests from wk 11 to wk 15 included
Peak Liquid Helium consumption by the test benches in g/s | 27 g/s | 36 g/s | 37 g/s
Total number of tests | 50 | 52 | 62
HL-LHC project tests | 34 | 27 | 26
KEY figures for the SM18 CRYO operation with liquid Helium
9
N. Guillotin
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th

## Slide 11
SM18 CRYOPLANT AVAILABILITIES COMPARISON
 | 2020 With 6 kW CB CB-A + 25kL Dewar | 2021 With new liquefier CB-B + 25 or 10kL DewAR | 2022 With new liquefier CB-B + 10kL DEWAR
SM18 CRYOPLANT AVAILABILITY | 96.8 % Lock down period for COVID’19 excluded, Equivalent to 55 days of stop for the tests | 81.2% 99.9 % IF Shutdown period for repair of the main helium liquid storage excluded, Equivalent to 57 days of stop | 99.9%
ROOT CAUSES | major unavailability related to : . FILTER clogging on 1st CB-A turbine . oil coalescer default on compressor side | minor defaults reported : . temporary water cut at the SM18 . under cooling of the CB-B FOLLOWING maladjusted cold gas returns from test benches | MINOR defaults reported : . CWU oil issues . CWS ISSUES . 80K-ADSORBER LEAKS
10
N. Guillotin
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th

## Slide 12
SM18 MAGNET TESTS PARTICULARITIES
About 20 Quenches performed in 2022 on A1 test bench, Nov. 14th to Dec. 15th
Average duration for these quenches recovery 2h30
With extrema from  50 min up to 19h for only one of the quenches in 2022
Recovery duration strongly depends on :
Number of tests performed in parallel and pumping availability
Type of tests and cryostat volumes concerned
Number of quenches in parallel
Peak current value reached
10kL Dewar liquid helium level
Average duration for quenches recovery on Cluster_D and HFM is 4h, With extrema from  1h up to 10h
High liquid helium volumes for specific cryostats like Cluster_D and HFM (as well as V3 for RF tests)
Significant heat load for Cryomodule tests on M7
High heat loads will be experimented on new shuffling modules with electrical heaters to ease electrical current ramp up
High 1.9K pumping consumption requested for HL-LHC RF cavity tests in 2023 with necessity of very stable values
RF
MAGNETS
Pressure
Magnet current
Quench
CrOK
11
N. Guillotin
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th

## Slide 13
KEY POINTS FOR SM18 OPERATION AND TESTS IN 2023
Installation of all or part of the new 1.9K pumping equipment for IT STRING during maintenance period.
Dense alternation between HFM and Cluster_D tests in 2nd semester
 Strong LHe and 1.9K pumping consumption with punctual high-level quench recovery.
Average LHe consumption for the test benches should stand theoretically at 75% of the max capacity [slide 4],
Leading to a sufficient margin in terms of coordination and operation
1.9K pumping consumption for the test benches should stand between 80 - 100% in the 2nd semester [slide 5]
 No operational margin
Conclusion :
No showstopper, but new constraint would impact the ideal tests planning as few margin on pumping capacities and LHe after quenches
High LHe consumption expected with new shuffling modules when electrical heaters are activated to ease magnets current ramp-up.
Specific conditions of operation not tested on long periods formerly  LHe and pumping consumption uncertainty for A2, B2, C2 and F1
CCU commissioning will require a period of 1.9K pumping availability reduction, not taken into account in the current simulations for 2023
Periods with magnets maintained at 80K could be optimized
12
N. Guillotin
TM REVIEW SESSION BEFORE LHC WORKSHOP CHAMONIX - Jan. 18th
Magnets cool down to 80K and warm up to 300K with CWU have to be anticipated due to high LN2 consumption, especially during WE