## Slide 1
MAJOR REFERENCES
John Kadyk: Proceedings of the Workshop on Radiation Damage to Wire Chambers. LBL-21170 (1986).
John Kadyk: Wire Chamber Aging. Nucl. Instr. Meth. A300 (1991) 436
Marcus Hohmann: Aging phenomena in gaseous detectors-perspectives from the 2001 workshop. Nucl. Instr. Meth. A494 (2002)494.
Mar Capeáns: Aging and materials: lessons for detectors and gas systems. Nucl. Instr. Meth. A515(2003)73.
Maxym Titov: Radiation damage and long-term aging in gas detectors. Proc. 42nd Workshop of the INFN ELOISATRON Project (2003)
Fabio Sauli: Gaseous Radiation Detectors: Fundaments and Applications. Cambridge University Press (2014).
RD10: ﻿A study to improve the radiation hardness of gaseous detectors for use at very high luminosity
RD28: Development of microstrip gas chambers for high-rate applications
RD51: Development of Micro-Pattern Gas Detectors Technomogies
GDD: Gas Detectors Development Group (CERN)
☛ Special Thanks to Eraldo Oliveri
Limited to MWPCs & MPGDs

## Slide 2
G. Charpak et al,
Nucl. Instr. Meth. 62(1968)282
1968: Multiwire Proportional Chamber
Drift Chambers
Time Projection
Jet Chambers
….

## Slide 3
G. Charpak et al, Nucl. Instr. Meth. 99(1972)269
Deterioration of Performances under Irradiation:
Rate Dependence of Gain
Higher Noise
After
~ 107 cm-2 𝞫 90Sr
AGING OF MWPCs
The Night Sky
Magic Gas
Ar-iC4H10-CF3Br

## Slide 4
R. Kotthaus,
Nucl. Instr. Meth. A252 (1986) 531
Examples of the Effect of Trace Pollutants in the Gas
Adding a Short PVC Tube
M. Capeáns,
ICFA Instr. Bull. (2002)
ATLAS/TRT Straws
Change of a Gas Flowmeter
PLASTICIZERS
SILICON COMPOUNDS

## Slide 5
Deposit on Anodes and Cathodes
Polymers of organic quencher
Pollutants
Materials Outgassing
SILICONE Deposits
(see Later)

## Slide 6
1988: Micro-Strip Gas Counter (MSGC)
A. Oed,
Nucl. Instr. Meth. 263(1988)351
Extreme Dependence on Gas Purity
R. Bouclier et al, Nucl. Instr. Meth. A348(1994)109

## Slide 7
MSGC: Aging and Discharges
DISCHARGES
AGING
☛ Discharges: See Piotr Gasik Talk
Discolorations (Deposits?)
Pitting (Micro-Discharges)
Full Discharge

## Slide 8
Mechanisms of Polymer Formation
Plasma Chemistry
J. Va'vra,
Nucl. Instr. Meth. A252(1986)547
﻿e+CH4  CH2:+H2+e
Production of CH2: radicals
(precursors of Polyethylene)
Polyethylene Compounds
☛ Plasma Chemistry See Katerina Kudnetzova

## Slide 9
W. Haberichter and H Spinka, AHL-HEP-CP-86-10
Silicone Compouds Deposits
Anode Wire Surface Analysis with Ion Microprobe Mass Analysis
NEW WIRE
IRRADIATED WIRE
Silicone from Materials J. Kadyk, Nucl. Instr. Meth. 300A(1991)436

## Slide 10
Polysilicon
Silicate
AVOID:
Plasticizers (Phtalates)
Sealant: RTV
Vacuum Grease
Oil Bubblers
Silicone oil (Dow 704)
Duo-Seal
Epoxies: G10
…. and many others
Very Effective Formation in the Avalanche of Silicone Polymers
G. Gavrilov et al. Phys. Part. Nuclei 49(2018)33

## Slide 11
RD10-RD28 Setup at CERN:
Study of Aging of Detectors Under Irradiation
M. Capeáns: Nucl. Instr. Meth. A515(2003)73.

## Slide 12
Units Used for Long-Term Aging Studies:
Irradiation Time
Radiation Flux : Counts/cm2
Charge per unit Surface: Coulomb/cm2
Charge per unit Wire Length: Coulomb/mm
LHC Years
Accelerated lifetime measurements:
High source intensity
High charge gain
Both ➙ Local Charge Density 10 to 100 times than normal

## Slide 13
Effect on Ageing Rate of Irradiation Current Density:
CAVEAT: HIGH-RATE MEASUREMENTS ARE OPTIMISTIC !
M. Capeáns Nucl. Instr. Meth. A515(2003)73

## Slide 14
M. Capeáns: Nucl. Instr. Meth. A515(2003)73.
Materials Outgassing and Effects on Detectors
☛ Maxym Titov: Gaseous Detectors and Radiation Hardness

## Slide 15
K. Kurvinen, RD-51 WG2 meeting 10 Dec. 2008
More Studies on Wire Chamber Chemistry

## Slide 16
K. Kurvinen, RD-51 WG2 meeting 10th Dec. 2008
Outgassing of Nylon Tubing

## Slide 17
The CF4 Case: Benefits and Harms
Wire Chamber Aging
Ar-C2H6
CF4-iC4H10
Recovery With CF4
R. Openshaw et al, Nucl. Instr. Meth A307(1991)298
CF4 + e --> CF3* + F*+ e
CF4 + e --> CF2* + 2F*+ e
CF4 + e --> CF3- + F*
Very Fast Drift Velocity and Low Diffusion
Dissociates to Reactive Species uner Avalanche Conditions
F + H2O  HF
Efficient Etching of Deposits on Wires
10 mC 90Sr
CAVEAT:
CF4 dissociation products react also with the detector building materials !

## Slide 18
Anode Wire Aging Recovery in CF4
Aged Wire Irradiated with 55Fe source in ﻿Ar-CO2 + 10% CF4
G. Gavrilov et al, ﻿Physics of Particles and Nuclei49(2018)33
"Training": Cyclic Exposure to 90Sr Source
in 80% CF4-20% CO2

## Slide 19
+
+
+
+
+
+
+
-
-
-
Positive Ions Deposit on the Thin Insukating Layer and Create an Electric dipole Field Extracting Electrons
Self-Sustained Effect Leading to Current Increase and Discharge
Cathode Field Emission – Malter Effect
Rate Dependent
Enhanced Sensitivity to Light
The Current Increase due to Malter Effect continues after removal of source
Source Removed
LHCb Muon MWPC:
G. Gavrilov et al,
Phys. Atomic Nuclei 82(2019)1273

## Slide 20
LHCb Muon Chambers : ~ 1400 MWPC with Ar-CO2-CF4
Standard Mixture
+2% O2 Training
Operated for 9 years without gain reduction
~ 20% MWPC developed high current and HV trips, attributed to Malter effect on Cathodes
F. P. Albicocco et al, JINST 14(2019)P11031
Malter Current
TRAINING : Reduction of Mather CurrentAdding O2 to Ar-CO2-CF4

## Slide 21
M. A, Buzoveria et al, Phys. Atomic Nuclei 82(2019) 1252
Catode Plane Damage in CMS CSC Chambers
The gas mixture used, (Argon-CO2–CF4) prevents the formation of anode and cathode deposits; damages to electrodes under sustained irradiation are attributed to CO2 dissociation products (O* , CO* O2 , O3).
CAVEAT: CF4 is Expensive, Difficult to Procure and Greenhouse Gas
Search for Eco-Friendly Replacements:
☛ Beatrice Mandelli, Maxym Titov + Many Presentations

## Slide 22
GOLDEN RULES OF AGING PREVENTION
Ultra-Pure gases
Non-Organic Quenchers (CO2)
Choice of non-Outgassing Building Materials
Non-Polymerizing Additives: Methylal, Propilic Alcohol
Improved Cleaning Protocols
Avoid Silicon-Containing Materials: Tubing, Sealings, ….
Zapping: Burning Deposits with High Current on Anodes
Addition of Compounds with Etching Properties (O2 , CF4)
The Creation of Reactive Species is Enhanced at High Values of the electric field

## Slide 23
Maximum Electric Field in MWPC & MSGC:
MWPC
MSGC
Computed by Djunes Janssens (CERN-GDD)
E ~ 400 kV/cm
E ~ 500 kV/cm

## Slide 24
The Solution: Get Rid of Thin Wires and Strips
➽ Micro-Pattern Gas Detectors (MPGD)
MICROMEGAS
Y. Giomataris et al,
Nucl. Instr. Meth. A376(1996)29
F. Sauli,
Nucl. Instr. Meth. A386(1997)531
Gas Electron Multiplier (GEM)
MiroGroove, MicroGap,MicroPixel
Resistive Plate Well ………..
+

## Slide 25
Electric Field Strength: MICROMEGAS & GEM
GEM
E~20 kV/cm
E~30 kV/cm
Computed by Djunes Janssens (CERN-GDD)
MICROMEGAS
E~ 140 kV/cm
E~ 100 kV/cm
Interwoven
Micromesh
Lower GEM Hole Edge

## Slide 26
Long-Term Aging Measurements: MICROMEGAS
HL-LHC ATLAS Upgrade
X-Ray Generator 11 kV
Ar-CO2 (90-10)
918 mC on 4 cm2 (230 mC/cm2)
J. Galan et al, JINST7(2012) ﻿C01041

## Slide 27
Long-Term Aging Mesurements
CMS Forward Muon Triple-GEM Upgrade
Ar-CO2(70-30)
F. Fallavollita, PhD Theses Pavia Univ. (2019)
X-Ray Source

## Slide 28
CAVEAT:
MPGD ARE ALSO SENSITIVE TO POLLUTANTS!
COMPASS Pixel GEM– After Four Years of Operation (165 mC/cm2)
Silicon Coating: (R2SiO) - Sulphur: Sulphuric Acid (H2SO4)
C. Dreisbach, MPGD Conf. Zaragoza (2013)
The Culptit: External Sealant DowCorning 1-2577

## Slide 29
﻿M. Jung, Ch. Garabatos and R. Negrão, RD51 Coll. Meeting Munich 19/06/2018
ALICE TPC AGING in Ar-CH4
Hydrocarbon Deposits on Copper

## Slide 30
﻿M. Jung, Master Thesis, Univ. Frankfurt (2019)
ALICE TPC AGING in Ar-CH4
Recovery after Replacement of CH4 with CO2
Kapton Ablation? Outgassing from Epoxy?
For Updates: ☛ See Other Presentations at this Conference
Circular Halo Around Holes

## Slide 31
J.A.Merlin
RD51 Miniweek CERN 4-7 Dic 2018
GEM: Silicon Deposits on Rim
SEM Analysis after long-Term Irradiation

## Slide 32
CONCLUSIONS (IF ANY!)
Despite many problems, wire-based detector systems have been successfully operated for over two decades
The development of novel MPGD devices and a better understanding of the aging processes have extended the lifetime of existing systems, as amply documented by the work discussed at this conference
A continuing development effort is needed to ensure operation at the ever increasing luminosity of colliders

## Slide 33
Everybody's Dream