## Slide 1: Magnetic Measurements on MCBRDP4
EDMSID: 2779805 01/02/2023
Piotr Rogacki
Lucio Fiscarelli
Carlo Petrone
Matthias Bonora
Agnieszka Chmielinska
and other colleagues in the section

## Slide 2: Outline
  Measurement setup
  Measurement Cycles Overview
  Standard Cycles
  Comparison with other MCBRD magnets

## Slide 3: Setup
Vertical field in AP1:
- main field B1 (normal)
- first allowed multipole b3
Horizontal field in AP2:
- main field A1 (skew)
- first allowed multipole a3
3
  5 segments (500 mm each, 2.5 m total)
    Centered around Segment 3
Rotating-coil in the helium bath of vertical cryostat ClusterD in SM18
Reference radius 35 mm
AP2
AP1

## Slide 4: Measurement Cycles Overview
Standard cycles

## Slide 5: Standard Cycles Integrated Field Strength
  AP1 at 393 A
    Individual Powering: 4.991 Tm
    Combined Powering: 4.982 Tm
  AP2 at 393 A
    Individual Powering: 4.986 Tm
    Combined Powering: 4.980 Tm
5
Difference AP1-AP2 at warm:
-10 units (AP2>AP1)
Difference individual-combined:
AP1: 18 units
AP2: 12 units
Difference AP1-AP2:
+10 units (individual)
+4 units (combined)

## Slide 6: Standard Cycles Integrated Field Strength Comparison
6
 | AP1 |  |  | AP2 |  | 
 | Individual | Combined | Diff from average | Individual | Combined | Diff from average
 | Tm | Tm | unit | Tm | Tm | unit
MCBRDP1 | 5.003 | 4.993 | 18 | 4.996 | 4.989 | 6
MCBRDP2 | 4.997 | 4.988 | 6 | 4.996 | 4.989 | 6
MCBRDP3 | 4.987 | 4.977 | -14 |  |  | 
MCBRDP4 | 4.991 | 4.982 | -6 | 4.986 | 4.980 | -14
MCBRD01 | 4.992 | 4.983 | -4 | 4.994 | 4.987 | 2
Difference individual-combined:
AP1: 19±1 units
AP2: 13±1 units

## Slide 7: Transfer function

## Slide 8: AP1 - Multipoles b2/a2

## Slide 9: AP1 - Multipoles b3/a3

## Slide 10: AP1 - Multipoles b5/a5

## Slide 11: AP2 - Multipoles b2/a2

## Slide 12: AP2 - Multipoles b3/a3

## Slide 13: AP2 - Multipoles b5/a5

## Slide 14: Field-quality tables - Integral
APERTURE 1 |  |  |  |  | 
Individual powering (positive-negative mean) |  |  | Difference from individual to combined powering |  | 
AP1 at ±393A AP2 at 0A |  |  | AP1 at ±393A AP2 at ±393A |  | 
 | Unit | Value |  | Unit | Value
Norm. TF | Tm/kA | 12.699 | Norm. TF Δ* | units | -18
Int. B1 | Tm | 4.991 | Int. B1 Δ* | units | -18
n | bn | an | n | Δbn** | Δan**
2 | 4.98 | 3.16 | 2 | 3.97 | ±5.60
3 | -4.29 | 0.14 | 3 | -3.41 | ±1.90
4 | -0.36 | -0.15 | 4 | 0.34 | ±0.81
5 | -1.74 | 0.23 | 5 | -0.32 | ±0.34
6 | 0.14 | 0.05 | 6 | -0.02 | ±0.12
7 | 0.36 | -0.09 | 7 | -0.02 | ±0.04
8 | 0.07 | -0.03 | 8 | -0.01 | ±0.00
9 | 0.02 | -0.07 | 9 | -0.01 | ±0.01
10 | 0.01 | 0.01 | 10 | 0.00 | ±0.00
11 | 0.01 | -0.01 | 11 | 0.00 | ±0.00
12 | -0.01 | 0.00 | 12 | -0.01 | ±0.00
13 | 0.00 | 0.00 | 13 | 0.00 | ±0.00
14 | 0.00 | 0.00 | 14 | 0.00 | ±0.00
15 | 0.00 | 0.00 | 15 | 0.00 | ±0.00
* Relative difference in units to nominal value in individual powering
** Difference in units between individual and combined powering - the ‘±’ indicates that the sign of the difference is dependent on the powering quadrant.
APERTURE 2 |  |  |  |  | 
Individual powering (positive-negative mean) |  |  | Difference from individual to combined powering |  | 
AP2 at ±393A AP1 at 0A |  |  | AP2 at ±393A AP1 at ±393A |  | 
 | Unit | Value |  | Unit | Value
Skew TF | Tm/kA | 12.687 | Skew TF Δ* | units | -12
Int. A1 | Tm | 4.986 | Int. A1 Δ* | units | -12
n | bn | an | n | Δbn** | Δan**
2 | -4.66 | -7.46 | 2 | ±7.46 | -5.48
3 | 0.05 | 1.84 | 3 | ±2.22 | -0.82
4 | 0.11 | -0.18 | 4 | ±0.10 | -0.17
5 | 0.02 | -1.50 | 5 | ±0.14 | -0.16
6 | 0.07 | 0.15 | 6 | ±0.05 | -0.07
7 | 0.04 | -0.31 | 7 | ±0.02 | -0.01
8 | -0.10 | 0.04 | 8 | ±0.01 | 0.00
9 | 0.01 | -0.01 | 9 | ±0.01 | 0.01
10 | 0.01 | 0.01 | 10 | ±0.00 | 0.00
11 | 0.00 | -0.01 | 11 | ±0.00 | 0.00
12 | 0.00 | 0.00 | 12 | ±0.00 | 0.00
13 | 0.00 | 0.00 | 13 | ±0.00 | 0.00
14 | 0.00 | 0.00 | 14 | ±0.00 | 0.00
15 | 0.00 | 0.00 | 15 | ±0.00 | 0.00

## Slide 15: Field-quality Multipole decay
Chart: Harmonic decay MCBRDP4 Individual powering
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
AP1 bn | 4.5404 | 3.5979 | 0.28441 | 1.7005 | 0.14526 | 0.35674 | 0.069133 | 0.026517 | 0.015519 | 0.006457 | 0.002655 | 0.002542 | 0.001495 | 0.001323
AP1 an | 3.6322 | 0.14061 | 0.010475 | 0.25226 | 0.08548 | 0.063193 | 0.018078 | 0.056383 | 0.012827 | 0.004064 | 0.002635 | 0.001054 | 0.000892 | 0.000476
AP2 bn | 4.2848 | 0.043084 | 0.061079 | 0.003879 | 0.057473 | 0.040505 | 0.095611 | 0.008815 | 0.005699 | 0.00114 | 0.00161 | 0.000755 | 7.3e-05 | 0.000177
AP2 an | 7.4641 | 1.8316 | 0.17885 | 1.5028 | 0.1519 | 0.30433 | 0.040115 | 0.006939 | 0.011747 | 0.011893 | 0.0031 | 0.002903 | 0.001807 | 0.000135
Decay law | 4.5918 | 2.1429 | 1 | 0.46667 | 0.21778 | 0.10163 | 0.047427 | 0.022133 | 0.010329 | 0.00482 | 0.0022493 | 0.0010497 | 0.00048986 | 0.0002286
Chart: Harmonic decay MCBRDP4 Combined powering
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
AP1 bn | 7.9666 | 6.6217 | 0.073351 | 2.0222 | 0.1442 | 0.34037 | 0.063042 | 0.038542 | 0.012268 | 0.008232 | 0.005649 | 0.003584 | 0.000196 | 0.001441
AP1 an | 8.4288 | 1.5771 | 0.68333 | 0.019379 | 0.18837 | 0.090136 | 0.033859 | 0.050921 | 0.011473 | 0.003645 | 0.001063 | 0.000461 | 0.000124 | 0.000132
AP2 bn | 2.841 | 2.2779 | 0.2061 | 0.091545 | 0.017998 | 0.02223 | 0.099469 | 0.00245 | 0.004648 | 0.001253 | 0.000945 | 0.000777 | 0.000224 | 0.000123
AP2 an | 12.6079 | 1.1798 | 0.34859 | 1.673 | 0.082434 | 0.31564 | 0.032715 | 0.010791 | 0.0134 | 0.010753 | 0.002118 | 0.002896 | 0.000873 | 0.000462
Decay law | 4.5918 | 2.1429 | 1 | 0.46667 | 0.21778 | 0.10163 | 0.047427 | 0.022133 | 0.010329 | 0.00482 | 0.0022493 | 0.0010497 | 0.00048986 | 0.0002286

## Slide 16: Standard Cycles Integrated Allowed Multipoles Comparison
16
 | AP1 |  | AP2 | 
 | b3 | b5 | a3 | a5
MCBRDP1 | -11 (-14) | 0.5 (0.4) | 7.5 (7) | 1 (0.9)
MCBRDP2 | -13.3 (-17.1) | 0.7 (0.3) | 10.3 (9.4) | 0.6 (0.5)
MCBRDP3 | -5 (-9) |  |  | 
MCBRDP4 | -4.3 (-7.7) | -1.7 (-2.0) | 1.8 (1.0) | -1.5 (-1.7)
MCBRD01 | -5.1 (-8.9) | -1.7 (-2.1) | 2.3 (1.6) | -2.1 (-2.3)
multipoles in units at nominal current
values in parenthesis - combined powering

## Slide 17: Summary
17
TF value within ±15 units from the average
Field quality: all multiples within ±10 units
Cross talk as expected from previous measurements on other magnets

## Slide 18: Measurements at ambient temperature

## Slide 19: Ambient temperature measurements
Chart: Angle [mrad]
| -1.0399999999999994 | -0.97499999999999942 | -0.90999999999999948 | -0.84499999999999942 | -0.77999999999999936 | -0.71499999999999941 | -0.64999999999999947 | -0.58499999999999952 | -0.51999999999999957 | -0.45499999999999963 | -0.38999999999999968 | -0.32499999999999973 | -0.25999999999999979 | -0.19499999999999984 | -0.12999999999999989 | -6.4999999999999947E-2 | 0 | 6.4999999999999947E-2 | 0.12999999999999989 | 0.19499999999999984 | 0.25999999999999979 | 0.32499999999999973 | 0.38999999999999968 | 0.45499999999999963 | 0.51999999999999957 | 0.58499999999999952 | 0.64999999999999947 | 0.71499999999999941 | 0.77999999999999958 | 0.84499999999999953 | 0.90999999999999948 | 0.97499999999999942 | 1.0399999999999994
AP1 (Pink) | 10.9948 | 7.872 | 3.0651 | 0.39874 | 0.077945 | 0.12331 | 0.16245 | 0.18526 | 0.21762 | 0.25821 | 0.2975 | 0.31221 | 0.31708 | 0.32177 | 0.30569 | 0.2861 | 0.27708 | 0.22942 | 0.17865 | 0.11652 | 0.062794 | 0.050795 | 0.0024985 | -0.15945 | -0.36913 | -0.51165 | -0.72175 | -0.90639 | -1.1145 | -1.7519 | -4.548 | -10.1641 | -16.0701
AP2 (Black) | 8.589 | 7.4832 | 3.0749 | 0.21253 | -0.11435 | 0.088065 | -0.15778 | 0.17705 | 0.15027 | 0.19128 | 0.1361 | 0.19593 | 0.22425 | 0.18737 | 0.23055 | 0.22817 | 0.25831 | 0.25954 | 0.1536 | 0.066714 | -0.10203 | -0.063341 | -0.11079 | -0.13108 | -0.19783 | -0.29722 | -0.34616 | -0.41817 | -0.60844 | -0.97256 | -3.9033 | -9.6986 | -13.407
Chart: Main field TF - str section [mT / A]
| -0.77999999999999936 | -0.71499999999999941 | -0.64999999999999947 | -0.58499999999999952 | -0.51999999999999957 | -0.45499999999999963 | -0.38999999999999968 | -0.32499999999999973 | -0.25999999999999979 | -0.19499999999999984 | -0.12999999999999989 | -6.4999999999999947E-2 | 0 | 6.4999999999999947E-2 | 0.12999999999999989 | 0.19499999999999984 | 0.25999999999999979 | 0.32499999999999973 | 0.38999999999999968 | 0.45499999999999963 | 0.51999999999999957 | 0.58499999999999952 | 0.64999999999999947 | 0.71499999999999941 | 0.77999999999999958
AP1 (Pink) | 6.593 | 6.5872 | 6.5867 | 6.5849 | 6.5837 | 6.5831 | 6.5821 | 6.5816 | 6.5821 | 6.5818 | 6.5807 | 6.5797 | 6.5805 | 6.5812 | 6.5808 | 6.5809 | 6.58 | 6.5791 | 6.5788 | 6.578 | 6.5775 | 6.5781 | 6.5806 | 6.5852 | 6.594
AP2 (Black) | 6.5964 | 6.5924 | 6.5909 | 6.5895 | 6.5894 | 6.5889 | 6.5877 | 6.5874 | 6.5874 | 6.5872 | 6.5865 | 6.5859 | 6.5865 | 6.5871 | 6.5867 | 6.5868 | 6.5859 | 6.585 | 6.5845 | 6.584 | 6.5834 | 6.5837 | 6.5859 | 6.59 | 6.5995

## Slide 20: Ambient temperature summary
AP1 (Pink) | Str. section |  |  | Integral | 
Current | 0.2475 | [A] |  | 0.2475 | [A]
B1 | 1.6290 | [mT] |  | 3.1130 | [mT m]
TF | 6.5825 | [mT/A] |  | 12.5794 | [mT m/A]
n | bn | an | n | bn | an
2 | 8.70 | -3.02 | 2 | 8.16 | -4.42
3 | 1.66 | 0.31 | 3 | 5.00 | -0.02
4 | 0.28 | -0.02 | 4 | 0.47 | -0.22
5 | -1.30 | 0.08 | 5 | -2.41 | -0.14
6 | 0.12 | 0.00 | 6 | 0.12 | 0.00
7 | 0.19 | 0.00 | 7 | 0.31 | -0.01
8 | 0.01 | -0.02 | 8 | 0.02 | -0.01
9 | -0.06 | 0.01 | 9 | -0.04 | 0.01
10 | 0.00 | 0.00 | 10 | 0.00 | -0.01
11 | 0.01 | 0.00 | 11 | 0.00 | 0.00
12 | 0.00 | 0.00 | 12 | 0.00 | 0.00
13 | 0.00 | 0.00 | 13 | 0.00 | 0.00
14 | 0.00 | 0.00 | 14 | 0.00 | 0.00
15 | 0.00 | 0.00 | 15 | 0.00 | 0.00
 | Units at 35 mm reference radius |  |  |  | 
AP2 (Black) | Str. section |  |  | Integral | 
Current | 0.2476 | [A] |  | 0.2476 | [A]
A1 | 1.6311 | [mT] |  | 3.1167 | [mT m]
TF | 6.5880 | [mT/A] |  | 12.5885 | [mT m/A]
n | bn | an | n | bn | an
2 | 4.97 | 4.87 | 2 | -4.88 | -6.08
3 | 0.19 | -0.67 | 3 | 8.36 | -8.33
4 | 0.07 | 0.27 | 4 | 0.27 | -0.27
5 | 0.02 | 0.09 | 5 | -2.26 | -2.26
6 | 0.00 | 0.04 | 6 | -0.04 | 0.03
7 | -0.01 | -0.01 | 7 | 0.27 | -0.27
8 | 0.01 | 0.01 | 8 | 0.01 | -0.02
9 | -0.01 | -0.01 | 9 | -0.04 | -0.04
10 | 0.00 | 0.00 | 10 | 0.00 | 0.01
11 | 0.00 | 0.00 | 11 | 0.01 | -0.01
12 | 0.00 | 0.00 | 12 | 0.00 | 0.00
13 | 0.00 | 0.00 | 13 | 0.00 | 0.00
14 | 0.00 | 0.00 | 14 | 0.00 | 0.00
15 | 0.00 | 0.00 | 15 | 0.00 | 0.00
 | Units at 35 mm reference radius |  |  |  | 

## Slide 21: Ambient temperature field multipoles
AP1 (Pink)
AP2 (Black)
Chart: bn [units]
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
 | 7.5508 | 2.0122 | 0.3523 | -1.211 | 0.12958 | 0.1885 | -0.0058315 | -0.049168 | -0.00013737 | 0.012535 | 0.0013215 | 0.0025271 | 0.0011524 | -0.0011175
 | 8.0796 | 1.912 | 0.29786 | -1.2198 | 0.14705 | 0.19571 | 0.015414 | -0.048374 | 0.0056207 | 0.015201 | 0.0046931 | 0.00037441 | 0.0015611 | -0.0013417
 | 8.6853 | 1.9133 | 0.24265 | -1.1853 | 0.1351 | 0.21357 | 0.012002 | -0.044828 | 0.0096024 | 0.012442 | 0.0068526 | 0.0017387 | 0.0029319 | -0.0013616
 | 8.7195 | 1.7272 | 0.19443 | -1.2357 | 0.078115 | 0.21947 | -0.023366 | -0.035906 | 0.0033596 | 0.011978 | 0.0028866 | -0.0058831 | 0.0021068 | -0.00055067
 | 8.9395 | 1.5602 | 0.16054 | -1.3572 | 0.05811 | 0.1702 | -0.013549 | -0.046125 | 0.0050496 | 0.015891 | -0.0039379 | -0.0015831 | 1.8889e-05 | -0.0015186
 | 9.5755 | 1.6574 | 0.25766 | -1.3592 | 0.10514 | 0.18138 | 0.0095575 | -0.04587 | 0.01678 | 0.01402 | 0.00017507 | -0.0023087 | 0.00088281 | -0.0011205
 | 9.2663 | 1.8698 | 0.20987 | -1.2935 | 0.11816 | 0.18411 | 0.0019279 | -0.045623 | 0.014546 | 0.011694 | 0.00041439 | -0.00074655 | 0.0023751 | -0.00099809
 | 8.3867 | 1.7079 | 0.045843 | -1.2819 | 0.072889 | 0.19396 | -0.014947 | -0.042778 | 0.0034812 | 0.011284 | -0.002581 | -0.001082 | 0.00059879 | 0.00082244
 | 8.7314 | 1.4946 | 0.097458 | -1.3248 | 0.066646 | 0.22165 | -0.017742 | -0.04525 | 0.00063104 | 0.0059723 | -0.0051928 | -0.0031579 | 8.4831e-05 | -0.00052877
 | 8.9751 | 1.4636 | 0.18533 | -1.3398 | 0.132 | 0.19317 | -0.0043953 | -0.060424 | -0.0045265 | 0.0079861 | -0.004169 | -0.0036842 | -0.0005223 | -0.0008198
 | 9.0228 | 1.5467 | 0.21735 | -1.2968 | 0.14028 | 0.17822 | 0.015587 | -0.067503 | -0.007807 | 0.013867 | -0.0088766 | -0.0010302 | -0.00064056 | -0.0012254
 | 9.0632 | 1.7106 | 0.31526 | -1.2933 | 0.12601 | 0.18618 | 0.014911 | -0.058825 | -0.0014438 | 0.010108 | -0.0078568 | 0.0018982 | -0.00028575 | -0.00093643
Chart: an [units]
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
 | -3.6748 | 0.34855 | -0.10457 | 0.029339 | -0.03384 | -0.036148 | -0.02519 | 0.0041287 | -0.0068701 | -0.0022059 | -0.0049901 | -0.0030941 | -0.0016896 | 0.0012642
 | -2.9678 | 0.38118 | -0.098417 | 0.087416 | -0.035384 | -0.026312 | -0.020361 | -0.0027235 | -0.0041695 | -0.0036874 | -0.0047045 | 5.5857e-05 | -0.0023708 | 0.0013017
 | -2.6996 | 0.41212 | 0.10534 | 0.10802 | 0.024426 | -0.012743 | -0.0165 | -0.0024626 | -0.0030946 | -0.0013717 | -0.00179 | -0.0034765 | -0.00076974 | 0.00061434
 | -3.5044 | 0.53356 | 0.034579 | 0.028276 | 0.0046512 | -0.036626 | -0.024503 | 0.0095582 | -0.0078468 | 0.0030359 | -0.0040664 | 0.0013021 | -0.0021016 | -0.00019798
 | -3.3378 | 0.55066 | -0.090904 | 0.051667 | -0.044778 | -0.0083226 | -0.0042419 | 0.027615 | 0.0020997 | 0.0048104 | -0.0031052 | 0.00050971 | -0.0022681 | -0.00017578
 | -2.9328 | 0.53597 | -0.010577 | 0.11129 | 0.015472 | 0.025294 | 0.0073409 | 0.03293 | -0.0020416 | 0.0026172 | -0.0022142 | 0.0010998 | -0.0010039 | 0.00036378
 | -3.1514 | 0.43212 | 0.02011 | 0.1024 | 0.016155 | 0.0091842 | 0.0093572 | 0.014889 | -0.0069469 | 0.00066524 | -0.0044069 | -0.0020382 | -0.0019641 | 8.4667e-05
 | -2.9101 | 0.25177 | 0.044569 | 0.057595 | 0.0014788 | 0.016982 | 0.011351 | 0.017506 | 0.0012492 | 0.0019422 | 0.00040165 | 0.00072552 | -5.8624e-06 | -0.0011355
 | -2.7387 | 0.17131 | 0.023425 | 0.040467 | -0.0054728 | 0.030517 | -0.035925 | 0.023861 | -0.001898 | 0.00151 | 0.0043305 | -0.00059558 | 0.0012059 | -0.00080751
 | -3.0218 | 0.27382 | 0.050969 | 0.093326 | 0.010251 | 0.0017449 | -0.05492 | 0.0082557 | -0.011096 | 0.0020721 | -0.0052991 | 0.0026525 | -0.00060707 | -0.0014045
 | -3.3522 | 0.26367 | 0.072955 | 0.033477 | 0.025055 | -0.022861 | -0.050339 | 0.0085342 | -0.010361 | 0.0047188 | -0.0016951 | -0.0010772 | -0.0016159 | -0.00065064
 | -2.9073 | 0.13462 | 0.029895 | 0.024625 | 0.01211 | -0.025827 | -0.030441 | 0.0079325 | -0.010278 | 0.0049407 | 0.0048785 | -0.003859 | -0.0026422 | -0.0014069
Chart: bn [units]
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
 | 2.947 | 0.19012 | 0.15641 | -0.014334 | -0.016605 | 0.013647 | 0.0023471 | -0.015219 | 0.0060127 | 0.0058472 | -0.0025705 | -0.0031281 | -0.00046991 | -0.0001214
 | 2.653 | 0.19649 | 0.065856 | -0.01054 | -0.0086892 | -0.0018236 | 0.012683 | -0.021849 | 4.4389e-05 | 0.0033897 | -0.0019516 | -0.0028375 | 0.00033524 | 0.0010106
 | 2.66 | 0.28616 | 0.040474 | -0.013302 | -0.022521 | 0.019177 | 0.022894 | -0.015392 | 0.00018361 | -0.00078811 | -0.00082652 | -0.00076681 | -0.00025869 | 0.0011055
 | 2.989 | 0.38024 | 0.093296 | 0.053782 | -0.018656 | 0.017177 | 0.017541 | -0.01545 | 0.0033332 | 0.00039547 | -0.0013387 | 0.00093429 | -0.00067629 | 0.0012005
 | 3.199 | 0.42973 | 0.12519 | 0.028812 | -0.029076 | 0.0059295 | 0.0089766 | -0.021503 | 0.0026068 | 0.0050563 | -0.0043699 | -0.0021592 | 0.00022575 | 0.00056835
 | 3.143 | 0.36309 | 0.1962 | 0.0062906 | -0.059665 | -0.018589 | 0.011612 | -0.015277 | 0.00023365 | 0.005819 | -0.00013624 | -0.0005841 | -0.00014824 | 0.00092917
 | 3.647 | 0.34545 | 0.17527 | -0.012629 | -0.046555 | -0.020942 | 0.023965 | -0.020753 | -0.00075461 | 0.0088132 | 0.00082652 | -0.0031973 | -0.0014533 | 0.00046051
 | 4.186 | 0.18571 | 0.12039 | 0.02305 | 0.012488 | -0.022118 | 0.025365 | -0.017813 | -0.0031597 | 0.0051128 | 0.00012892 | 0.0022561 | -0.00079837 | 0.00033301
 | 5.25 | 0.022393 | 0.12691 | 0.070589 | -0.021681 | -0.018 | 0.013177 | -0.0065719 | 0.0048021 | -0.0051693 | -0.0040733 | -0.0015364 | 0.00041953 | -0.00026993
 | 5.677 | 0.16905 | 0.033134 | -0.014454 | -0.0079665 | 0.028589 | -0.011447 | -0.021042 | 0.0092006 | -0.00044066 | -0.0030846 | 0.00025606 | -0.00086522 | 0.00054733
 | 5.579 | 0.28175 | 0.012965 | 0.018584 | -0.0041849 | -0.010612 | -0.011035 | -0.01251 | 0.01481 | 0.00093499 | -0.001568 | 0.002353 | -0.00023544 | 0.00069179
 | 6.076 | 0.30821 | 0.069972 | 0.065547 | -0.038824 | -0.060472 | -0.00077001 | 0.0070907 | 0.0205 | -0.0051128 | -0.0054772 | 0.001744 | 0.00023447 | 0.00024077
Chart: an [units]
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
 | -5.138 | -4.1601 | -0.37387 | -1.0348 | 0.024202 | -0.19765 | -0.020424 | -0.046925 | 0.0019531 | -0.013954 | 0.0028671 | -0.0045815 | 6.0265e-05 | -0.00052901
 | -5.411 | -4.2532 | -0.40474 | -1.0348 | 0.048572 | -0.18236 | -0.023224 | -0.038624 | 0.0034744 | -0.017881 | 0.0030055 | -0.0045815 | -0.00071407 | -0.00023331
 | -5.418 | -4.1944 | -0.40131 | -1.0636 | 0.040001 | -0.16236 | -0.033601 | -0.03603 | 0.0048828 | -0.017259 | 0.005754 | -0.00045538 | -0.00010948 | 0.00011191
 | -5.754 | -4.0376 | -0.40817 | -1.0588 | 0.0090254 | -0.16941 | -0.026436 | -0.056495 | -0.00060127 | -0.011751 | 0.0076325 | -0.0011613 | -0.0019281 | 0.00022178
 | -6.251 | -4.0866 | -0.33374 | -0.9652 | 0.0031093 | -0.1553 | -0.024789 | -0.059377 | 0.0052863 | -0.018869 | 0.0027485 | 0.00025606 | -0.001434 | -0.00019804
 | -6.517 | -4.1258 | -0.27646 | -0.98681 | 0.015597 | -0.16589 | -0.036565 | -0.057648 | 0.0089989 | -0.019689 | 0.0030649 | -0.0025191 | 0.00074508 | -0.00082065
 | -7.07 | -4.0621 | -0.24765 | -1.0276 | 0.0072774 | -0.18471 | -0.039942 | -0.073789 | 0.015375 | -0.0040959 | 0.0049631 | -0.0023115 | 0.0010755 | -1.9465e-05
 | -8.19 | -4.067 | -0.17493 | -0.9556 | -0.024538 | -0.20706 | -0.030142 | -0.062836 | 0.016706 | -0.0088132 | 0.0028869 | -0.0010492 | 0.0012014 | -3.0724e-05
 | -7.98 | -4.1405 | -0.26171 | -0.94599 | 0.023194 | -0.21294 | -0.057566 | -0.042083 | 0.014729 | -0.017429 | -0.0020762 | 0.00026714 | 0.003333 | 0.0005514
 | -7.98 | -4.2091 | -0.22123 | -1.0252 | 0.068573 | -0.18941 | -0.064319 | -0.040411 | 0.018361 | -0.011017 | -0.0023332 | -0.0023945 | 0.0028388 | 5.1206e-05
 | -7.91 | -3.9396 | -0.067914 | -1.0588 | 0.041513 | -0.17412 | -0.044636 | -0.044562 | 0.021307 | -0.013644 | -0.0005754 | -0.00089415 | -0.00031392 | 7.5961e-05
 | -7.42 | -3.8857 | -0.14234 | -1.0804 | 0.024202 | -0.15765 | -0.034506 | -0.059954 | 0.019612 | -0.0094064 | 0.0031044 | -0.0014257 | -0.0017537 | -0.0002679

## Slide 22: Other Integral Multipoles

## Slide 23: AP1 - Multipoles b4/a4

## Slide 24: AP1 - Multipoles b7/a7

## Slide 25: AP1 - Multipoles b9/a9

## Slide 26: AP1 - Multipoles b11/a11

## Slide 27: AP1 - Multipoles b13/a13

## Slide 28: AP2 - Multipoles b4/a4

## Slide 29: AP2 - Multipoles b7/a7

## Slide 30: AP2 - Multipoles b9/a9

## Slide 31: AP2 - Multipoles b11/a11

## Slide 32: AP2 - Multipoles b13/a13

## Slide 33: Integral Field

## Slide 34: AP1 Field-quality table - Integral Field
AP1 at 393A AP2 at 0A |  |  | AP1 at -393A AP2 at 0A |  |  | AP1 at 393A AP2 at 393A |  |  | AP1 at -393A AP2 at 393A |  |  | AP1 at 393A AP2 at -393A |  |  | AP1 at -393A AP2 at -393A |  | 
 | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value
TF | Tm/kA | 10.003 | TF | Tm/kA | 10.002 | TF | Tm/kA | 9.984 | TF | Tm/kA | 9.983 | TF | Tm/kA | 9.983 | TF | Tm/kA | 9.984
B1 int | Tm | 3.931 | B1 int | Tm | -3.931 | B1 int | Tm | 3.924 | B1 int | Tm | -3.923 | B1 int | Tm | 3.923 | B1 int | Tm | -3.924
n | bn | an | n | bn | an | n | bn | an | n | bn | an | n | bn | an | n | bn | an
2 | 4.54 | 3.63 | 2 | 4.33 | 3.58 | 2 | 7.97 | 8.43 | 2 | 8.15 | -1.46 | 2 | 8.12 | -1.52 | 2 | 7.99 | 8.38
3 | -3.60 | -0.14 | 3 | -3.52 | -0.14 | 3 | -6.62 | -1.58 | 3 | -6.53 | 1.79 | 3 | -6.47 | 1.83 | 3 | -6.66 | -1.54
4 | -0.28 | -0.01 | 4 | -0.32 | -0.01 | 4 | 0.07 | 0.68 | 4 | -0.02 | -0.75 | 4 | -0.02 | -0.75 | 4 | 0.08 | 0.69
5 | -1.70 | 0.25 | 5 | -1.69 | 0.25 | 5 | -2.02 | -0.02 | 5 | -1.96 | 0.59 | 5 | -1.95 | 0.60 | 5 | -2.02 | -0.00
6 | 0.15 | 0.09 | 6 | 0.14 | 0.08 | 6 | 0.14 | 0.19 | 6 | 0.11 | -0.02 | 6 | 0.12 | -0.02 | 6 | 0.15 | 0.19
7 | 0.36 | -0.06 | 7 | 0.36 | -0.06 | 7 | 0.34 | -0.09 | 7 | 0.35 | -0.03 | 7 | 0.35 | -0.03 | 7 | 0.33 | -0.10
8 | 0.07 | -0.02 | 8 | 0.07 | -0.02 | 8 | 0.06 | -0.03 | 8 | 0.07 | -0.04 | 8 | 0.07 | -0.04 | 8 | 0.06 | -0.03
9 | 0.03 | -0.06 | 9 | 0.03 | -0.05 | 9 | 0.04 | -0.05 | 9 | 0.01 | -0.05 | 9 | 0.01 | -0.05 | 9 | 0.04 | -0.05
10 | 0.02 | 0.01 | 10 | 0.02 | 0.01 | 10 | 0.01 | 0.01 | 10 | 0.01 | 0.01 | 10 | 0.01 | 0.01 | 10 | 0.01 | 0.01
11 | 0.01 | -0.00 | 11 | 0.01 | -0.01 | 11 | 0.01 | -0.00 | 11 | 0.01 | -0.01 | 11 | 0.01 | -0.01 | 11 | 0.01 | -0.01
12 | -0.00 | 0.00 | 12 | -0.00 | 0.00 | 12 | -0.01 | 0.00 | 12 | -0.01 | 0.00 | 12 | -0.00 | 0.00 | 12 | -0.01 | 0.00
13 | -0.00 | 0.00 | 13 | -0.00 | 0.00 | 13 | -0.00 | 0.00 | 13 | -0.00 | -0.00 | 13 | -0.00 | -0.00 | 13 | -0.00 | 0.00
14 | 0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | -0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | 0.00 | 0.00
15 | 0.00 | -0.00 | 15 | 0.00 | 0.00 | 15 | 0.00 | 0.00 | 15 | 0.00 | -0.00 | 15 | 0.00 | -0.00 | 15 | 0.00 | -0.00

## Slide 35: AP2 Field-quality table - Integral Field
AP2 at 393A AP1 at 0A |  |  | AP2 at -393A AP1 at 0A |  |  | AP2 at 393A AP1 at 393A |  |  | AP2 at -393A AP1 at 393A |  |  | AP2 at 393A AP1 at -393A |  |  | AP2 at -393A AP1 at -393A |  | 
 | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value
TF | Tm/kA | 12.687 | TF | Tm/kA | 12.686 | TF | Tm/kA | 12.671 | TF | Tm/kA | 12.669 | TF | Tm/kA | 12.670 | TF | Tm/kA | 12.670
A1 int | Tm | 4.986 | A1 int | Tm | -4.986 | A1 int | Tm | 4.980 | A1 int | Tm | -4.979 | A1 int | Tm | 4.979 | A1 int | Tm | -4.979
n | bn | an | n | bn | an | n | bn | an | n | bn | an | n | bn | an | n | bn | an
2 | -4.28 | -7.46 | 2 | -5.04 | -7.45 | 2 | 2.84 | -12.61 | 2 | -12.08 | -13.26 | 2 | -12.24 | -13.24 | 2 | 2.68 | -12.62
3 | 0.04 | 1.83 | 3 | 0.05 | 1.85 | 3 | 2.28 | 1.18 | 3 | -2.16 | 0.87 | 3 | -2.17 | 0.85 | 3 | 2.28 | 1.18
4 | 0.06 | -0.18 | 4 | 0.16 | -0.18 | 4 | 0.21 | -0.35 | 4 | 0.02 | -0.35 | 4 | 0.02 | -0.36 | 4 | 0.21 | -0.34
5 | 0.00 | -1.50 | 5 | 0.04 | -1.50 | 5 | -0.09 | -1.67 | 5 | 0.18 | -1.64 | 5 | 0.17 | -1.64 | 5 | -0.11 | -1.67
6 | 0.06 | 0.15 | 6 | 0.07 | 0.15 | 6 | 0.02 | 0.08 | 6 | 0.12 | 0.09 | 6 | 0.13 | 0.09 | 6 | 0.02 | 0.08
7 | 0.04 | -0.30 | 7 | 0.04 | -0.31 | 7 | 0.02 | -0.32 | 7 | 0.07 | -0.31 | 7 | 0.06 | -0.31 | 7 | 0.02 | -0.32
8 | -0.10 | 0.04 | 8 | -0.09 | 0.04 | 8 | -0.10 | 0.03 | 8 | -0.09 | 0.05 | 8 | -0.09 | 0.05 | 8 | -0.10 | 0.03
9 | 0.01 | -0.01 | 9 | 0.01 | -0.01 | 9 | 0.00 | -0.01 | 9 | 0.01 | -0.00 | 9 | 0.01 | -0.00 | 9 | 0.00 | -0.01
10 | 0.01 | 0.01 | 10 | 0.00 | 0.01 | 10 | 0.00 | 0.01 | 10 | 0.00 | 0.01 | 10 | 0.01 | 0.01 | 10 | 0.00 | 0.01
11 | -0.00 | -0.01 | 11 | -0.00 | -0.01 | 11 | -0.00 | -0.01 | 11 | -0.00 | -0.01 | 11 | -0.00 | -0.01 | 11 | -0.00 | -0.01
12 | 0.00 | 0.00 | 12 | 0.00 | 0.00 | 12 | 0.00 | 0.00 | 12 | 0.00 | 0.00 | 12 | 0.00 | 0.00 | 12 | 0.00 | 0.00
13 | 0.00 | -0.00 | 13 | 0.00 | -0.00 | 13 | 0.00 | -0.00 | 13 | 0.00 | -0.00 | 13 | 0.00 | -0.00 | 13 | 0.00 | -0.00
14 | 0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | 0.00 | 0.00
15 | 0.00 | -0.00 | 15 | 0.00 | -0.00 | 15 | -0.00 | -0.00 | 15 | 0.00 | 0.00 | 15 | 0.00 | 0.00 | 15 | 0.00 | -0.00

## Slide 36: Central Field on Straight section

## Slide 37: Field-quality tables - Straight section
APERTURE 1 |  |  |  |  | 
Individual powering (positive-negative mean) |  |  | Difference from individual to combined powering |  | 
AP1 at ±393A AP2 at 0A |  |  | AP1 at ±393A AP2 at ±393A |  | 
 | Unit | Value |  | Unit | Value
Norm. TF | T/kA | 6.676 | Norm. TF Δ* | units | -19
B1 | T | 2.624 | B1 Δ* | units | -19
n | bn | an | n | Δbn** | Δan**
2 | 4.44 | 3.61 | 2 | 3.62 | ±4.95
3 | -3.56 | -0.14 | 3 | -3.01 | ±1.69
4 | -0.30 | -0.01 | 4 | 0.33 | ±0.72
5 | -1.70 | 0.25 | 5 | -0.29 | ±0.30
6 | 0.15 | 0.09 | 6 | -0.02 | ±0.11
7 | 0.36 | -0.06 | 7 | -0.02 | ±0.03
8 | 0.07 | -0.02 | 8 | -0.01 | ±0.01
9 | 0.03 | -0.06 | 9 | -0.01 | ±0.00
10 | 0.02 | 0.01 | 10 | -0.01 | ±0.00
11 | 0.01 | -0.01 | 11 | 0.00 | ±0.00
12 | 0.00 | 0.00 | 12 | -0.01 | ±0.00
13 | 0.00 | 0.00 | 13 | 0.00 | ±0.00
14 | 0.00 | 0.00 | 14 | 0.00 | ±0.00
15 | 0.00 | 0.00 | 15 | 0.00 | ±0.00
* Relative difference in units to nominal value in individual powering
** Difference in units between individual and combined powering - the ‘±’ indicates that the sign of the difference is dependent on the powering quadrant.
APERTURE 2 |  |  |  |  | 
Individual powering (positive-negative mean) |  |  | Difference from individual to combined powering |  | 
AP2 at ±393A AP1 at 0A |  |  | AP2 at ±393A AP1 at ±393A |  | 
 | Unit | Value |  | Unit | Value
Skew TF | T/kA | 6.671 | Skew TF Δ* | units | -12
A1 | T | 2.622 | A1 Δ* | units | -12
n | bn | an | n | Δbn** | Δan**
2 | -5.26 | -8.91 | 2 | ±8.37 | -6.09
3 | -0.13 | 2.38 | 3 | ±2.46 | -0.85
4 | 0.09 | -0.11 | 4 | ±0.06 | -0.17
5 | 0.07 | -1.51 | 5 | ±0.18 | -0.18
6 | 0.07 | 0.16 | 6 | ±0.07 | -0.08
7 | 0.03 | -0.31 | 7 | ±0.03 | -0.01
8 | -0.09 | 0.05 | 8 | ±0.01 | 0.00
9 | 0.02 | 0.00 | 9 | ±0.01 | 0.00
10 | 0.00 | 0.01 | 10 | ±0.00 | 0.00
11 | 0.00 | -0.01 | 11 | ±0.00 | 0.00
12 | 0.00 | 0.00 | 12 | ±0.00 | 0.00
13 | 0.00 | 0.00 | 13 | ±0.00 | 0.00
14 | 0.00 | 0.00 | 14 | ±0.00 | 0.00
15 | 0.00 | 0.00 | 15 | ±0.00 | 0.00

## Slide 38: AP1 Field-quality table - Straight section
AP1 at 393A AP2 at 0A |  |  | AP1 at -393A AP2 at 0A |  |  | AP1 at 393A AP2 at 393A |  |  | AP1 at -393A AP2 at 393A |  |  | AP1 at 393A AP2 at -393A |  |  | AP1 at -393A AP2 at -393A |  | 
 | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value
TF | T/kA | 6.676 | TF | T/kA | 6.676 | TF | T/kA | 6.664 | TF | T/kA | 6.664 | TF | T/kA | 6.663 | TF | T/kA | 6.664
B1 | T | 2.624 | B1 | T | -2.624 | B1 | T | 2.619 | B1 | T | -2.619 | B1 | T | 2.619 | B1 | T | -2.619
n | bn | an | n | bn | an | n | bn | an | n | bn | an | n | bn | an | n | bn | an
2 | 5.08 | 3.18 | 2 | 4.88 | 3.13 | 2 | 8.85 | 8.63 | 2 | 9.06 | -2.57 | 2 | 9.04 | -2.63 | 2 | 8.86 | 8.57
3 | -4.32 | 0.14 | 3 | -4.25 | 0.14 | 3 | -7.77 | -1.47 | 3 | -7.63 | 2.33 | 3 | -7.58 | 2.37 | 3 | -7.80 | -1.42
4 | -0.34 | -0.15 | 4 | -0.38 | -0.14 | 4 | 0.03 | 0.64 | 4 | -0.08 | -0.98 | 4 | -0.09 | -0.98 | 4 | 0.04 | 0.65
5 | -1.74 | 0.22 | 5 | -1.73 | 0.23 | 5 | -2.10 | -0.08 | 5 | -2.01 | 0.61 | 5 | -2.01 | 0.62 | 5 | -2.10 | -0.06
6 | 0.14 | 0.05 | 6 | 0.13 | 0.04 | 6 | 0.13 | 0.17 | 6 | 0.09 | -0.08 | 6 | 0.10 | -0.07 | 6 | 0.14 | 0.17
7 | 0.35 | -0.09 | 7 | 0.36 | -0.09 | 7 | 0.33 | -0.12 | 7 | 0.35 | -0.05 | 7 | 0.35 | -0.06 | 7 | 0.33 | -0.13
8 | 0.07 | -0.03 | 8 | 0.07 | -0.03 | 8 | 0.06 | -0.05 | 8 | 0.07 | -0.05 | 8 | 0.06 | -0.05 | 8 | 0.06 | -0.05
9 | 0.02 | -0.06 | 9 | 0.02 | -0.07 | 9 | 0.03 | -0.06 | 9 | -0.00 | -0.07 | 9 | 0.00 | -0.07 | 9 | 0.03 | -0.06
10 | 0.01 | 0.01 | 10 | 0.01 | 0.01 | 10 | 0.01 | 0.01 | 10 | 0.01 | 0.01 | 10 | 0.01 | 0.01 | 10 | 0.01 | 0.01
11 | 0.01 | -0.01 | 11 | 0.01 | -0.00 | 11 | 0.01 | -0.00 | 11 | 0.01 | -0.00 | 11 | 0.01 | -0.01 | 11 | 0.01 | -0.01
12 | -0.00 | 0.00 | 12 | -0.01 | 0.00 | 12 | -0.01 | 0.00 | 12 | -0.01 | 0.00 | 12 | -0.01 | 0.00 | 12 | -0.01 | 0.00
13 | -0.00 | -0.00 | 13 | -0.00 | 0.00 | 13 | -0.00 | 0.00 | 13 | -0.00 | -0.00 | 13 | -0.00 | -0.00 | 13 | -0.00 | -0.00
14 | 0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | 0.00 | -0.00 | 14 | -0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | 0.00 | 0.00
15 | 0.00 | 0.00 | 15 | 0.00 | -0.00 | 15 | 0.00 | -0.00 | 15 | 0.00 | -0.00 | 15 | 0.00 | 0.00 | 15 | 0.00 | -0.00

## Slide 39: AP2 Field-quality table - Straight section
AP2 at 393A AP1 at 0A |  |  | AP2 at -393A AP1 at 0A |  |  | AP2 at 393A AP1 at 393A |  |  | AP2 at -393A AP1 at 393A |  |  | AP2 at 393A AP1 at -393A |  |  | AP2 at -393A AP1 at -393A |  | 
 | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value |  | Unit | Value
TF | T/kA | 6.671 | TF | T/kA | 6.670 | TF | T/kA | 6.661 | TF | T/kA | 6.660 | TF | T/kA | 6.661 | TF | T/kA | 6.661
A1 | T | 2.622 | A1 | T | -2.621 | A1 | T | 2.618 | A1 | T | -2.617 | A1 | T | 2.618 | A1 | T | -2.618
n | bn | an | n | bn | an | n | bn | an | n | bn | an | n | bn | an | n | bn | an
2 | -4.88 | -8.92 | 2 | -5.64 | -8.90 | 2 | 3.15 | -14.65 | 2 | -13.57 | -15.36 | 2 | -13.74 | -15.33 | 2 | 3.00 | -14.66
3 | -0.14 | 2.37 | 3 | -0.12 | 2.39 | 3 | 2.34 | 1.70 | 3 | -2.57 | 1.37 | 3 | -2.57 | 1.35 | 3 | 2.35 | 1.70
4 | 0.03 | -0.11 | 4 | 0.14 | -0.11 | 4 | 0.15 | -0.28 | 4 | 0.03 | -0.28 | 4 | 0.03 | -0.29 | 4 | 0.15 | -0.28
5 | 0.05 | -1.51 | 5 | 0.09 | -1.51 | 5 | -0.08 | -1.71 | 5 | 0.27 | -1.66 | 5 | 0.26 | -1.67 | 5 | -0.10 | -1.70
6 | 0.06 | 0.16 | 6 | 0.07 | 0.16 | 6 | 0.01 | 0.08 | 6 | 0.14 | 0.09 | 6 | 0.15 | 0.09 | 6 | 0.01 | 0.07
7 | 0.03 | -0.31 | 7 | 0.03 | -0.31 | 7 | 0.01 | -0.32 | 7 | 0.06 | -0.32 | 7 | 0.06 | -0.32 | 7 | 0.00 | -0.32
8 | -0.09 | 0.05 | 8 | -0.09 | 0.05 | 8 | -0.10 | 0.04 | 8 | -0.09 | 0.06 | 8 | -0.09 | 0.07 | 8 | -0.10 | 0.04
9 | 0.02 | -0.00 | 9 | 0.02 | -0.00 | 9 | 0.01 | -0.01 | 9 | 0.02 | 0.01 | 9 | 0.02 | 0.01 | 9 | 0.01 | -0.01
10 | -0.00 | 0.01 | 10 | -0.00 | 0.01 | 10 | -0.00 | 0.01 | 10 | -0.00 | 0.01 | 10 | -0.00 | 0.01 | 10 | -0.00 | 0.01
11 | -0.00 | -0.01 | 11 | -0.00 | -0.01 | 11 | -0.00 | -0.01 | 11 | -0.00 | -0.01 | 11 | -0.00 | -0.01 | 11 | -0.00 | -0.01
12 | 0.00 | 0.00 | 12 | 0.00 | 0.00 | 12 | 0.00 | 0.00 | 12 | 0.00 | 0.00 | 12 | 0.00 | 0.00 | 12 | 0.00 | 0.00
13 | 0.00 | -0.00 | 13 | 0.00 | -0.00 | 13 | 0.00 | -0.00 | 13 | 0.00 | -0.00 | 13 | 0.00 | -0.00 | 13 | 0.00 | -0.00
14 | 0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | 0.00 | 0.00 | 14 | -0.00 | 0.00
15 | -0.00 | -0.00 | 15 | -0.00 | -0.00 | 15 | -0.00 | -0.00 | 15 | 0.00 | -0.00 | 15 | 0.00 | -0.00 | 15 | -0.00 | -0.00

## Slide 40: Operation Levels

## Slide 41: AP1 - Operation Levels TF

## Slide 42: AP2 - Operation Levels TF

## Slide 43: AP1 Nominal Op. Level 1 - Multipoles

## Slide 44: AP1 Nominal Op. Level 2 - Multipoles

## Slide 45: AP2 Nominal Op. Level 1 - Multipoles

## Slide 46: AP2 Nominal Op. Level 2 - Multipoles

## Slide 47: Operation Levels Field-quality table - Integral
AP1 |  |  |  |  | 
AP1 at 393A AP2 at 314A |  |  | AP1 at 80A AP2 at 40A |  | 
 | Unit | Value |  | Unit | Value
TF | Tm/kA | 12.690 | TF | Tm/kA | 12.683
B1 int | Tm | 4.987 | B1 int | Tm | 1.015
n | bn | an | n | bn | an
2 | 5.30 | 4.53 | 2 | 4.77 | 4.12
3 | -4.46 | -0.36 | 3 | 0.84 | -0.82
4 | -0.24 | 0.14 | 4 | -0.14 | -0.11
5 | -1.80 | 0.19 | 5 | -1.45 | 0.19
6 | 0.14 | 0.10 | 6 | 0.12 | 0.09
7 | 0.35 | -0.07 | 7 | 0.14 | -0.02
8 | 0.07 | -0.02 | 8 | 0.06 | -0.01
9 | 0.03 | -0.05 | 9 | 0.02 | -0.04
10 | 0.02 | 0.01 | 10 | 0.01 | 0.01
11 | 0.01 | -0.00 | 11 | 0.00 | -0.00
12 | -0.00 | 0.00 | 12 | -0.00 | 0.00
13 | -0.00 | 0.00 | 13 | -0.00 | 0.00
14 | 0.00 | 0.00 | 14 | 0.00 | 0.00
15 | 0.00 | 0.00 | 15 | 0.00 | 0.00
AP2 |  |  |  |  | 
AP2 at 314A AP1 at 393A |  |  | AP2 at 40A AP1 at 80A |  | 
 | Unit | Value |  | Unit | Value
TF | Tm/kA | 12.688 | TF | Tm/kA | 12.678
A1 int | Tm | 3.984 | A1 int | Tm | 0.507
n | bn | an | n | bn | an
2 | 0.11 | -5.92 | 2 | -1.61 | -2.93
3 | 1.48 | 1.50 | 3 | 1.37 | 0.54
4 | 0.31 | -0.09 | 4 | 0.70 | 0.05
5 | 0.01 | -1.19 | 5 | 0.25 | -1.76
6 | 0.07 | 0.06 | 6 | 0.10 | 0.06
7 | 0.05 | -0.28 | 7 | 0.06 | 0.06
8 | -0.07 | 0.03 | 8 | -0.10 | 0.06
9 | -0.00 | -0.00 | 9 | 0.01 | 0.00
10 | 0.01 | 0.01 | 10 | 0.01 | 0.01
11 | -0.00 | -0.01 | 11 | -0.00 | 0.00
12 | 0.00 | 0.00 | 12 | 0.00 | 0.00
13 | 0.00 | -0.00 | 13 | 0.00 | -0.00
14 | 0.00 | 0.00 | 14 | -0.00 | 0.00
15 | -0.00 | -0.00 | 15 | -0.00 | 0.00